## Item 1 Hemodialysis (HD) machin
Form D: Technical Quotation Form
RFQ reference no: RFQ/2025/56702
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 1: Hemodialysis (HD) machine including Hemofiltration unit, mobile, for emergency therapy
Model (full name) |  | Insert Details
Manufacturer (full name) |  | Insert Details
Country of Origin |  | Insert Details
| Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Description: HD machines clean blood by passing it through a filter called a dialyzer. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Conformity Certificate IEC 60601-1-2 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Conformity Certificate IEC 60601-1-8 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Conformity Certificate IEC 60601-1-11 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Conformity Certificate IEC 60601-2-16 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Conformity Certificate IEC 62304 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical specifications:
8 | Weight of the unit: V -/+ 10 % 90 kg (including options) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Dimensions: HxWxD ± 10% : 140 x 48 x 48 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Inlet Pressure: 1 - 6.0 bar | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Water Inlet Temperature: 5 - 30 degrees Celsius | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Internal Battery autonomy at least 30 minutes of operation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | LAN (RJ 45) port for data exchange with Therapy Data Management System Technical Remote access | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Operational Features
15 | Device mounted on a mobile base equipped with a braking system and a transportation handle | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Supplied with IV pole | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Equipped with 4 castor wheels: At least two wheels with brakes | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Minimum 10-inch colour touchscreen monitor | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Display of pictograms and messages for intuitive and quick use | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Visual, Sound messages, Display and labelling: Russian | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Control of volumetric ultrafiltration and sequential ultrafiltration | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Fully automatic system readiness tests before dialysis start | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Short time between treatment less than 40 minutes including Disinfection + functions check | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Heparin pump for syringes of at least 20 ml | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Bolus Function: 1.0 to 5.0 mL (minimum range) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Blood leak detector and air bubble detector Sensitivity of ≤ 0.5 mL blood/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Patient disconnection available during treatment | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Self clearing Arterial/Venous alarms | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Possibility to pre-set common clinical settings to limit the number of parameters that need to be programmed for each patient | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Treatment, Alarm and disinfection history | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Dialysis fluid flow rate setting of at least 100 to 1000 ml/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Pyrogen and endotoxin filter to produce ultra-pure dialysates | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Treatment Time adjustable up to 9 hr 59 min. in 1 min increment. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | TMP Monitoring should be displayed | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Integrated patient card | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | 1- A dialysate circuit equipped with components enabling the following functions to be performed: Adjustment and control of the dialysate flow rate by a pump characterised by:
37 | Variable flow 0 - 800 ml/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Flow setting accuracy is less than or equal to ±10%. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | Dialysate temperature setting and control | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | Dialysate conductivity control 12 - 15 mS/cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | Bicarbonate & Concentrate Pump | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | Blood leakage detector | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | 2- A system of audible and visual alarms with adjustable values ​​concerning the conductivity, the temperature and the flow rate of the dialysate
44 | Audible and visual alarm system for venous, arterial, transmembrane pressure and retrofiltration. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | Venous pressure range: -60 to 500mmHg minimum Accuracy: ± 20 mmHg, Resolution: 10 mmHg or better | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | Arterial Blood pressure range: -280 to 300mmHg minimum Accuracy: ± 20 mmHg, Resolution: 10 mmHg or better | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
47 | 3-A degassing system
48 | System that ensures the degassing of the hydraulic circuit | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
49 | 4- A blood circuit equipped with:
50 | Blood pump with adjustable flow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
51 | Blood pump maximum flow greater than or equal to 500 ml/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
52 | Blood pump Accuracy less than or equal to ± 10% , Resolution: 10 mL/min or better and adaptable to all standard lines. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
53 | Integrated heparin pump with adjustable flow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
54 | Flow rate, double needle: 0 to 500 mL/min Flow rate, single needle: 0 to 500 mL/min, pressure-controlled | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
55 | Air detector with safety venous clamp | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
56 | 5- Ultrafiltration controller allowing the adjustment and permanent control of weight loss
57 | Maximum ultrafiltration rate higher than 4L/hour. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
58 | 6- Kt /V Module | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
59 | An integrated module for instantaneous Kt/V measurement | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
60 | 7- A thermochemical disinfection system for the device.
61 | Disinfection temperature greater than or equal to 84°C. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
62 | Programmable hot cleaning and disinfection modes | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
63 | 8- Included Automatic temperature , body composition monitor and blood pressure measurement.
64 | Integrated blood pressure monitoring system with cuff, BPM & BTM included | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
65 | systolic, diastolic, average blood pressure and heart rate (PAS, PAD, PAM, FC) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
66 | Temperature monitoring range 35.0 degree C to 40.0 degree C if applicable | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
67 | BCM body composition monitor | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
68 | 9- Preventive Maintenance Plan
69 | Provide the composition of these 2500 hrs and 5000 hrs kits or equivalent in accordance with the manufacturer's requirements | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
70 | 10- Accessories & consumables & reagents/solutions Price List
71 | Provide Spare parts, Preventive Maintenance kit and accessories price list | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
72 | Provide Consumables price list | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
73 | Provide reagents , solutions price list | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
74 | Included Ancillary Services
75 | Delivery on site, installation , See Associated services | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
76 | Testing & Commissioning, See Associated services | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
77 | Preventive Maintenance, Corrective/Repair and technical assistance contract for users 1 year | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
78 | 230 V AC ± 10%, 50 - 63 Hz | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
79 | Current Consumption: max 9 A | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
80 | Power consumption: Max 2100 Watt | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
81 | Built-in electrical protection device: Fuses | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
82 | The supply of complete revision, preventive maintenance kit for the warranty year | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
83 | The supply of complete revision kits for each machine: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
84 | •03 revision kit - Approx 2500 hrs of operation or Equivalent | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
85 | •03 revision kit - Approx 5000 hrs of operation or Equivalent | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
86 | The supply of Qty 04, dialysis fluid ultrafilters per machine. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
87 | The supply of 1 kit of consumables (Patients cards, Qty 1 dialysis fluid filter, Qty 1 Citrate, Qty 1 dry bicarbonate, Qty 2 Dialyzers authorised in Uzbekistan, Qty 2 tubing set , Qty 4 Syringes for heparin, etc.) for machine start-up operations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
88 | The supply of reagents and solutions including cleaning solutions (Qty 1 Citric Acid, Qty 1 disinfectant, Qty 1 Cleaning Sporotal, Qty 1 Diasteril/Citrosteril, Qty 1 Puristeril 340/plus, pH test strips, Potassium-iodide-starch paper, etc.) necessary for machines start-up operations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
89 | Various accessories (Qty 3 Bp Cuff adult size 17-25, 23 - 33, 30 - 40 cm; Mains cable; hoses (for water inlet and drain); dialyzer holder; dialyzer connector set; Top tray; BPM cuff holder; infusion stand; CD (Operator’s manual included ...) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
90 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
91 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification) | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.3.2 of the Schedule of Requirements of draft contract | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section III, 3.3 of the Schedule of Requirements of draft Contract | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section III. 3.4 of the Schedule of Requirements of draft contract | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section III.2 of the Schedule of Requirements are met of the draft contract | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
92 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
93 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
94 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements and Schedule IV: Draft Contract for Goods.
☐ Yes ☐ No
ANY DEVIATION MUST BE LISTED BELOW:
Name :
Title :
Date :
Signature : _____________________________________________________________
[Stamp this form with official stamp of the bidder]

## Item 2 Water treatment system f
Form D: Technical Quotation Form
RFQ reference no: RFQ/2025/56702
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 2: Water treatment system for dialysis (for one patient)
Model (full name) |  | Insert Details
Manufacturer (full name) |  | Insert Details
Country of Origin |  | Insert Details
| Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Description: Water treatment through reverse osmosis (RO) systems helps ensure patients undergoing hemodialysis receive safe treatment. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical specifications:
3 | Certificate of conformity / declaration of conformation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Function: Device for the treatment of water by osmosis for mobile individual dialysis. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Mini station capable of serving two dialysis machines. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Production flow from 80 to 100 (L / Hour). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Filtration system. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | DE chlorination system. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Membrane system to remove mineral salts. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Conductivity meter with digital display. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Automatic sterilization. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Conductivity alarm. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Components, spare parts, consumables:
13 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
14 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
D | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification) | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
E | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
F | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
G | Confirm that Assembly and Calibration shall be completed as per Section II.3.2 of the Schedule of Requirements of draft contract | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
H | Confirm that Training services shall be provided as per Section III, 3.3 of the Schedule of Requirements of draft Contract | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
I | Confirm that Warranty and After-sales services shall be provided as per Section III. 3. of the Schedule of Requirements of draft contract | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
J | Confirm that General Requirements as per Section III.2 of the Schedule of Requirements are met of the draft contract | ☐ Yes ☐ No | Insert details of how these requirements are met
K | Documentation and Licences
15 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
16 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
17 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements and Schedule IV: Draft Contract for Goods.
☐ Yes ☐ No
ANY DEVIATION MUST BE LISTED BELOW:
Name :
Title :
Date :
Signature : _____________________________________________________________
[Stamp this form with official stamp of the bidder]

## Services and delivery requireme
Form D: Technical Quotation Form
RFQ reference no: RFQ/2025/56702
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Service requirements – Comparative Data Table
S/N | Description of Services | Is quotation compliant? Bidder to complete | Details Bidder to complete
1 | The Supplier shall provide all the necessary documentation in a timely manner for facilitating the Customs Clearance process by the Consignee. If applicable, This includes the Certificate of Registration / Authorization / Waiver for the import and use of the devices by the health regulatory authority of the Republic of Uzbekistan that shall be obtained by the Supplier prior to the goods’ arrival to the country (or application for registration shall be made and letter from relevant regulatory authority for the Customs shall be obtained). Any demurrage caused due to incomplete documentation shall be borne by the Supplier. | ☐ Yes ☐ No | Insert details
2 | Delivery, ( If applicable installation or assembly, calibration, testing, commissioning), users’ training services as well as maintenance during warranty period at Republican Specialized Scientific and Practical Medical Center of Oncology and Radiology. Address: Shifokorlar Street, Almazar District, Tashkent city. | ☐ Yes ☐ No | Insert details
Delivery requirements –– Comparative Data Table
UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete
Delivery time for the goods | Bidder shall make the goods available at FCA (Incoterms2020) point within 30 days after Contract signature and/or deliver them to DPU (Incoterms 2020) Tashkent, Uzbekistan within 60 days after Contract signature. | ☐ Yes ☐ No | Insert details
Delivery time for the services | Bidder shall deliver the goods from Tashkent to all the final destination and perform the related services within 30 days after goods’ arrival to Tashkent (or its vicinity). | ☐ Yes ☐ No | Insert details
Delivery place and Incoterms rules | Bidder’s quoted FCA point – FCA (Incoterms2020) Tashkent, Uzbekistan – DPU (Incoterms2020) Door-to-door within Tashkent, Uzbekistan, as provided in the Section IV: Draft Contract for Goods. | ☐ Yes ☐ No | Insert details
Consignee details | United Nations Office for Project Services (UNOPS) AUMCO Uzbekistan Office 4A Afrosiyob Street, Tashkent, Uzbekistan Phone: +998 71 205 1255 | ☐ Yes ☐ No | Insert details
UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantities of the equipment components and consumables specified above, provided this does not exceed +/- 20%, without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details
The offered goods and related services (if applicable) are in accordance with the required specifications and requirements specified in Section II: Schedule of Requirements and Schedule IV: Draft Contract for Goods.
☐ Yes ☐ No
ANY DEVIATION MUST BE LISTED BELOW:
Name :
Title :
Date :
Signature : _____________________________________________________________
[Stamp this form with official stamp of the bidder]