Section II: Schedule of Requirements

eSourcing reference: RFQ/2020/16759

Technical specifications for Goods and Comparative Data Table:

| Item No | UNOPS minimum technical requirements | Qty. | Is bid compliant? Bidder to complete | Details of goods offered. Bidder to complete | Brochure Page No. |
| --- | --- | --- | --- | --- | --- |
| 1 | Water Treatment Plant for Hemodialysis Equipment | 1 |  |  |  |
| 2 | Portable Water Treatment Plant for Hemodialysis Equipment | 1 |  |  |  |
|  | Name of Manufacturer |  |  | Insert details | Insert |
|  | Model/ catalogue number |  |  | Insert details | Insert |
|  | Country of Origin for the offered model |  |  | Insert details | Insert |
|  | Country where the manufacturer is based |  |  | Insert details | Insert |
|  | Delivery time |  |  | Insert details | Insert |
|  | Full warranty period |  |  | Insert details | Insert |
|  | FDA clearance OR CE Mark |  |  | Insert details | Insert |
| Technical specifications for items 1 and 2: |  |  |  |  |  |
| 1 | Sand Filter Design: Suspended solids & turbidity. Filter Media: a. Graged silica sand & supported gravel. b. High quality, resisting attrition and possessing durability and long life. Tank: a. Working Pressure: 6.9 bar (100 psi). b. Hydrostatically test at 150% of working pressure. c. Opening on top to replace filter media. a. Provide single tank. Controls: b. Piston type, motor driven control valve. c. Self-adjusting flow control to control the backwash and rinse rates. d. Time clock initiates backwash on a calendar clock basis. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 2 | Carbon Filter Design: - Dechlorinates and absorbs dissolved organics. Filter Media: c. Activated carbon mesh & supported gravel. d. High quality, resisting attrition and possessing durability and long life. e. Shall not impart taste, odor, or color to the water. Tank: d. Working Pressure: 6.9 bar (100 psi). e. Hydrostatically test at 150% of working pressure. f. Opening on top to replace filter media. g. Provide single tank. Controls: a. Piston type, motor driven control valve. b. Self-adjusting flow control to control the backwash and rinse rates. c. Time clock initiates backwash on a calendar clock basis. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 3 | Reverse Osmosis Unit Design Requirements: Provides water that is free of dissolved solids, bacteria, virus and pyrogenic contaminants. Design Parameters: a. The RO system shall produce minimum water flow as indicated in the drawings when operating at 60% recovery. b. The system shall use Thin-Film-Composite (TFC) membranes with salt rejection of 96-98% and > 99% rejection of organics > 150 molecular weight (mw). c. Membrane elements shall be provided with internal interconnectors and double O-Ring seals. Pump: a. Construction: Multi-stage centrifugal pump with stainless steel type 316 and Noryl impellers. d. Motor: 3 Hp, 208-230/ 460VAC, 3 phase. Member Housing: Side entry, stainless steel with internal load transfer device to prevent telescoping of membrane elements. Housing shall be mounted horizontally for ease of servicing. Accessories: a. Sampling valves and flow meters that permit monitoring of product water and reject water without line disconnection. b. Adjustable waste flow control with minimum flow preset at factory. c. Sanitation ports with valves on feed, permeate and concentrate lines. d. Automatic waste water flushing system based on timed cycles (field adjustable). e. Adjustable waste flow control. 5 micron pre-filter to be built on to the RO frame. Instrumentation: a. On/ off switch. b. Low feed water pressure alarm post 5 micron filter, with time delay. c. Built-in water quality monitor on permeate. d. Pressure gages for pre-filter, post-filter, pump discharge, and final pressure. e. Motor overload alarm with immediate shutdown. f. High temperature shutdown. g. Temperature monitor (panel mounted). |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 4 | Storage Tank and Accessories Storage Tank: a. Level control with low-level alarm. b. Dished bottom, flat top with sealed cover.PE. Tank shall be FDA approved. c. Capacity: 2000 liters (500 gallons) minimum. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 5 | Ultraviolet Light Construction: a. Housing: Type 304 stainless steel with No.4 finish. b. Chamber: Type 316 stainless steel, passivated and electro polished. c. Provide with flanges on inlet and outlet. Design: a. Operating pressures: 830 kPa (120 psi) operating, 1240 kPa (180 psi) test pressure. b. Water temperature range: 2 to 38°C (35-100°F). c. Ambient temperature range: 7 to 55°C (45-130°F). d. Designed for continuous operation with 99% bacterial reduction. Controls: a. Interlocked with loop distribution pumps turn OFF whenever the pump goes OFF. b. LED UV lamp display providing continuous visual indication of lamp status. c. Running time meter (non-resettable) for recording operating hours. e. Tie UV sensor to pump for shutoff. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 6 | Cation Bed resin / water softener Provide duplex sets, two in parallel. Operating Data: a. Capacity: 125,000 grains. b. Flow rate: as indicated in the drawings. c. Pressure Drop: 1. Minimum: 70kPa (10psi). 2. Nominal: 140kPa (20psi). Primary Separate Bed Cation/softening Unit Bed Cation/softener Unit: Provide cation unit consisting of a strongly acidic cation resin for the removal of hardness. Resin Trap Cartridge Filter Housing and Filters Resin Trap Cartridge Filter Housing: Similar in construction to pre-filter housing. Provide cartridge filter housing with liquid filled stainless steel pressure gages on inlet and outlet. Resin Trap Filter Cartridges: 5 micron absolute pre-rinsed, high quality, nonfiber releasing, with polypropylene core, designed for use with RO water. Provide O-ring connectors for fitting into the housing. Fiberglass-Reinforced Polyester Tank. Design: a. Base wall and head thickness calculations using a 10 to 1 safety factor. Provide vertical tanks with convex dished top heads and continuously supported flat bottoms. b. Tank Tie-Down Lugs: Provide four or more hold-down lugs for vertical tanks. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 7 | Pumps Construction: a. Stainless steel, type 316; suction and discharge chambers, shaft and shaft seals. b. Ceramic; bearings. c. EPDM; O-rings. d. Noryl impellers. Design: a. Design working pressure: 16 bar (230 psi) (maximum) b. Pump: Horizontal with multi-stage impeller design. Centrifugal Pumps: Designed to meet the requirements of ANSI B-73.1 with passivated 316 stainless steel wetted parts, Viton, elastomers, and bellows- type mechanical seals. Provide 316 stainless steel, ceramic, or glass filled Teflon wetted seal components. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 8 | Filters: Housings: a. Constructed from stainless steel type 304 or polypropylene. b. Designed for minimum 100 psi operating pressure. c. Holds multiple cartridges. d. Provide bypass valves around each housing to allow for servicing without shutdown. Biological Cartridges: a. Designed to filter micron particles. e. Designed to filter 0.2-micron particles with maximum of 34kPa (5-psi) pressure drop. Polycarbonate media. 0.2 micron. Sediment Cartridges: Construct of polypropylene. Provide primary pre-filters with string wound on a polypropylene center core and 5 micron absolute polypropylene membrane secondary prefilters. Carbon cartridges: Consist of specially processed, granular, highly activated carbon to provide removal of a broad spectrum of low molecular weight organics, specific non-polar high molecular weight organics, free chlorine, and partially dissolved gases. a. Provide 12 by 40 granular carbon mesh with a total surface area of 1000m²/ g. Iodine number shall be 1000 minimum, molasses number shall be 230 minimum, and abrasion number shall be 75 minimum. b. Precondition activated carbon by hydration and backwashing to remove fines. Do not use dry, unconditioned activated carbon in tank. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |
| 9 | System Instrumentation Instrumentation and monitoring shall include but not be limited to the following: a. Feed water flow meter. b. Pressure gage on inlet and outlet of pre-filters, resin traps, and final filters. c. Flow monitor on outlet. d. Pressure gage on distribution/ recirculation pumps discharge. e. Pressure gages before and after the back pressure control valve. Provide ESP dedicated bacteria sample ports at the following locations: a. Outlet of water softener. b. Inlet to RO unit. c. Outlet 0.2 micron final filters. d. Return loop outlet. |  | ☐ Yes ☐ No | Insert details of goods offered, including specifications and brand/model offered if applicable | Insert |

Delivery requirements and Comparative Data Table:

| UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete |
| --- | --- | --- | --- |
| Delivery time | Bidder shall deliver the goods [4 weeks] after Contract signature. | ☐ Yes ☐ No | Insert details |
| Delivery place and Incoterms rules | DAP to Royal Medical Services (Medical City in Amman) Incoterms rules as per Incoterms 2010. | ☐ Yes ☐ No | Insert details |
| Consignee details | Royal Jordanian Army with notifying UNOPS | ☐ Yes ☐ No | Insert details |
| UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- [20%] , without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details |

00ax: +45 45 33 75 01