## Annex A
INVITATION TO BID: ITB 2021 008 FOR THE ESTABLISHMENT OF FA FOR MEDICAL ITEMS Annex A – Technical specification
Item number | Item Description | Specification | Unit of Measure | Estimated quantity
SECTION 1: HOSPITAL SUPPLIES
1 | Endotracheal tube, without cuff, disposable, sterile, PVC (different sizes) | • Open distal end and Magill-type point with oral angle of 37.5º. • Standard connector (ext. Ø 15mm) at the proximal end enabling the tube to be connected to the ventilation system. • Radio opaque mark. • With Murphy's eye. • Graduations.• Endotracheal tube without cuff. • Size: Ø internal 2 / 2.5 / 3 / 3.5 / 4 mm • Material: Polyvinyl chloride (PVC). • Disposable. • Sterile. • Initial sterilisation method: Ethylene oxide gas or Gamma radiation. | EA | Size 2 --> 6 Size2.5 --> 7 Size 3 --> 7 Size 3.5 --> 6 Size 4 --> 4
2 | Endotracheal tube, with cuff, disposable, sterile, PVC (different sizes) | • Open distal end and Magill-type point with oral angle of 37.5º. • Standard connector (ext. Ø 15mm) at the proximal end enabling the tube to be connected to the ventilation system. • Radio opaque mark. • With Murphy's eye. • Graduations. • Endotracheal tube without cuff. • Size: Ø internal 3 mm, 3.5 mm, 4 mm, 4.5 mm, 5mm, 5.5 mm, 6 mm, 6.5mm, 7mm, 7.5mm, 8mm or 8.5 mm • Material: Polyvinyl chloride (PVC). • Disposable. • Sterile. • Initial sterilisation method: Ethylene oxide gas or Gamma radiation. | EA | Size 3 --> 1 Size 3.5 --> 1 Size 4 --> 2 Size 4.5 --> 2 Size 5 --> 4 Size 5.5 --> 5 Size 6 --> 5 Size 6.5 --> 20 Size 7 --> 40 Size 7.5 --> 60 Size 8 --> 60 Size 8.5 --> 40
3 | Ultrasound gel | Hypoallergenic, water soluble gel suitable for ultrasonography, doppler and defibrillator applications. It ensures the ultrasound waves come to the equipment display in a clear and uninterrupted way. Salt, formaldehyde, fragrance free. Appearance and Odor: Colourless, viscous gel | EA | 1 unit of 6 Liters 42 units of 260 ml
4 | Nasal cannula (prong) replacement for use with Oxygen concentrator | Nasal prongs (nasal cannula) is a device designed for easy administration of oxygen and comfort of patient. The device consists of a plastic tube which fits behind the ears, and a set of two prongs which are placed in the nostrils. Soft twin prongs nasal tips to ensure equal oxygen flow to both.Star lumen main tube to avoid accidental blockage. Adjustable, smoothly finished, nasal tips for maximum patient comfort. Soft funnel shaped connector to facilitate easy connection to oxygen source. Oxygen tube length: approximately 2m. | EA | 800
5 | Nebulizer Mask Adult | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, adult, nonsterile, single use | Pcs | 100
6 | Nebulizer Mask Pediatric | Mask, oxygen, with connection tube, reservoir bag and valve, high-concentration, pediatric, non-sterile, single use | Pcs | 50
7 | Needle, sterile 23G (0.6 x 25 mm), s.u. | General Description: Needle with base (Luer type fitting and complies with ISO 594-1:1986, ISO 594-2:1998 and BS EN 20594-1:1994), and protective cap. Blister packed. Technical specifications: Material: Needle: Stainless steel conform to ISO 9626-1:1991. Base and protective cap: medical grade plastic. Measurements: External diameter expressed in Gauge and millimetres. Length expressed in inch and millimetres. Size selected: 23Gx1” (0.6x25mm). Colour code: visible at the base of the needle and complies with ISO 6009-3:1992. Disposable. Sterile, (ethylene oxide sterilisation and compiles to BS EN 550:1994-ISO 11135-1:2007) Standards: Product and packaging conform to ISO 7864:1993, ISO 9626-1:1991, BS EN 550:1994, ISO 11135-1:2007. Quality management system (QMS): Products and manufacturer must be certified and comply with ISO 13485: 2003 and 2008 requirements (or equivalent). Marketing approval certificate: Product and manufacturer must conform to European Directive MDD 93/42/EEC on Medical Devices (or equivalent). Packaging and labelling: Primary packaging: Each needle packed in an individual sterilised peel-off pack made of paper and/or plastic. Name of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date: month and year. The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilisation protection before use" (if space allows). The words "destroy after use" (if space allows). CE marking. Secondary packaging: Protective packaging - 1 carton of 100 needles. Name and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date: month and year. The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the primary container individual sterile peel-pack before use" (if space allows). The words "destroy after use" (if space allows). Number of units per secondary package. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. (Alternatively, the instruction for use can be indicated on a separate insert). CE marking. | EA | 1000
8 | Syringe , 10ml, sterile, s.u. | General Description: Syringe, 10ml, sterile, with bi-packed needle, 21G (0.80 x 38-40mm), disposable, box of 100 Technical Specifications: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Capacity: 10ml. Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Concentric or eccentric luer nozzle. Material: Polyethylene (PEF) or polypropylene (PP). Needle with base (conical Luer type fitting) and protecting cap. Material: Needle: stainless steel; base and protecting cap: plastic. Size selected: 21G (0.80 x 38-40mm). Disposable. Sterile. Sterilisation method: Ethylene oxide. Syringe with bi-packed needle conforms to EN-ISO 7886-1, EN-ISO 7864, EN-ISO 20594, EN-ISO 9626-1991, EN 550-ISO 11135. Packaging and labelling: Primary packaging: Each syringe and needle bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Name of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilisation protection before use" (if space allows). The words "destroy after use" (if space allows). CE marking. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Name and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date: month and year. The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the primary container individual sterile peel pack before use" (if space allows). The words "destroy after use" (if space allows). Number of units per secondary package. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. (Alternatively, the instruction for use can be indicated on a separate insert). CE marking. Weight/Volume/Dimensions: Estimated weight: 0.866kg Estimated volume: 7.917cdm (The above weight and volume is for box of 100 syringes w/needles). Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Safety boxes are available as a UNICEF standard warehouse item reference: S0782208 - Safety box f.used syrgs/ndls 5lt/BOX-25. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. For more information on waste management, please refer to WHO publication "Safe Management of Wastes from Health Care Activities". | EA | 2100
9 | Infusion giving set with air inlet and needle, s.u. | Infusion giving set, with airinlet and needle, sterile, single-use | EA | 750
10 | Infusion giving set pediatric, with burette, s.u. | Infusion giving set pediatric, with airinlet and needle, sterile, single-use | EA | 190
11 | Tongue depresseor, wooden, disp. | General Description Tongue depressor, wooden, disp/BOX-100 with rounded extremities. Size: approx. Size: 17 x 149mm Thickness: Approximately 1.6mm Single use. Non sterile. Packaging and labelling: Primary packaging: One (1) Box of 100 pcs wooden tongue depressors. Secondary packaging: Carton of 50 boxes of 100 pcs Over packaging: Carton of 50 boxes of 100 pcs | EA | 4066
12 | BAG BODY, 8 handles, U-shaped zip, white, 400 microns, adult, 230x100cm | General Description: Body bag including absorbent pad used for packing and transporting contagious human corpses. Adult size. Technical Specifications: Body bag: With zipper, includes an absorbent pad prefixed inside of the body bag and padded carry handles during handling and transportation of the corpse ·Single use, disposable ·Tear proof and puncture resistant ·Seams entirely heat-sealed, width minimum of 5 mm ·100% leak proof seams ·Integrated transparent label pouch 10 x 15 cm (l x w) for placement of identification tag ·Thickness: 300 to 400 microns (um) ·Color: white ·Carry capacity: up to 270 kg ·Shelf life: minimum 5 years ·Certificate of origin ·Minimum: 2 ply Body bag dimensions (with prefixed absorbent pad inside): ·Unfolded: 230 x 110 cm (l x w) (± 10%) ·Individually packed, supplied folded in a transparent plastic bag ·Label indicates product short description with size in cm and inch Body bag material: ·PEVA ·Special linear reinforced material ·Chloride and carbon free ·Non-porous and non-degradable ·Buried in soil, the body bag will remain solid over a period of minimum 5 years Carry handles: ·Minimum of 6 padded carry handles ·Non splintering handles ·Reinforced and integrated into the bag material Zipper: ·High quality black or white zipper, smooth running with teeth made of nylon ·Zipper is sewed into the body bag fabric with double stitch ·Full length U shaped ·With 25 cm return allowing to show the face of the decease person ·Double runner ·Large plastic or sturdy cloth loops, on each of the zipper runners ·Stop for the zipper is reinforced into the body bag and heat-sealed Absorbent pad: for body bags to absorb fluids. Integrated and prefixed inside of the body bag, used in the context of infection control ·Preapplied inside of the body bag with water resistant double sided tape ·Two layers of padding ·Absorption capacity: minimum 5 liters ·Resistant to 0.5% chlorine solution ·Single use, disposable ·Size: 160 x 70 cm (± 15%) ·Shelf life: minimum 5 years Absorbent pad material: ·Thick cotton lining on polyethylene base ·Chloride free Conforms to: (a)OSHA regulation 3130 for Universal Precautions involving containment of body fluids and protection from blood borne pathogens or (b)Equivalent applicable product or other performance standard(s). Intended use and considerations: Bag intended for epidemics with contagious corpses (e.g. cholera and haemorrhagic fever) and for emergencies, used for enclosing human corpses, before burying. Wherever possible, follow local customs regarding funeral rites. Information concerning the corpses should be written on the body bags with permanent markers. The head of the deceased should be positioned on the side of the zipper's return in order to show the face without having to open the whole body-bag. Shelf life: Shelf life in normal storage conditions (protected from humidity, sunlight, high temperatures and dust): minimum 5 years. Storage conditions: As indicated by manufacturer Supplied with: Manufacturer's instruction for use | EA | 100
13 | Electric Sprayer with Anti-Microbal, Disinfecting solution for Fogging machine, 5 Liter | Fumigation Machine for Hospitals - Supplier to provide the detailed specifications | EA | 24
14 | Thermometer, clinical, IR, handheld set | Infrared (IR) clinical thermometer, digital, handheld, minimmal range 34– 42 °C TECHNICAL SPECIFICATIONS: Infrared (IR) non-contact clinical thermometer, Measurement of skin temperature trough thermal radiation detection, Factory calibrated and pre-set emissivity for all skin types Measuring distance between 3 – 12 cm / 1.22 – 6 inch, depending on the model Handheld design with ergonomic pistol type grip or similar design. IR sensor located at the distal end (patient side) Digital display located at the proximal end (user side) Handgrip with index finger or thumb switch to trigger measurements Preferably backlit LCD display, suitable for reading in bright sunlight Display visualisation of measured temperature, battery status ,operational status, errors including incorrect reading Digital read-out, user-selectable, in Celsius or Fahrenheit scale Low and high temperature visual or audiable alarms Temperature measurement range 34 – 42 °C / 93 – 108 °F Graduation 0.1 °C / 0.3 °F Accuracy 0.3 °C (minimum guaranteed) Automatic switch-off when not in use Hold function, to keep the latest measurement visualised on the display Internal memory (the presence of this feature depends on the available model), to retain several readings Visual or audio alert on switch-on, ready-to-use and measurement completed. Water spill proof and suitable to be cleaned and disinfected Battery powered, 2 x 1.5 V AA or AAA commercial batteries depending on the model (included in the supply). ITEMS SUPPLIED WITH: 1 x set of batteries (separately packed) Instruction for use and maintenance in English | EA | 34389
15 | Paper roll, for examination table | White non transparent, Paper material, roll Width: 50 cm Length: 100 m Fluid absorbent, Perforated | EA | 184120
16 | Viral transport medium with swab. Medium 3ml | Viral Transport Medium with Swab., Medium 3 ml and Nylon Flocked swab | EA | 760
17 | Ethanol (Alcohol )95% (in Gallon 20 Liter) | Ethanol, Alcohol 95% or 5 Liter X4 | EA | 8
SECTION 2: HYGIENE, INFECTION, PREVENTION AND CONTROL - IPC ITEMS
18 | Alcoholic, Hand rub – (Sterilium) 70% ethanol, 250 ML | Bottle of 250 ML at least 80% ethanol or 75% isopropyl alcohol (v/v) Quote for each size sperately | EA | 300
19 | Alcoholic, Hand rub – (Sterilium) 70% ethanol, 500ML | Bottle of 500ml, at least 80% ethanol or 75% isopropyl alcohol (v/v) Quote for each size sperately | EA | 161000
20 | Alcoholic, Hand rub – (Sterilium) 70% ethanol, 1L | Bottle of 1000ml , at least 80% ethanol or 75% isopropyl alcohol (v/v) Quote for each size sperately | EA | 3000
21 | Hand sanitizer, gel, 60ml | Fluid absorbent, Perforated , at least 80% ethanol or 75% isopropyl alcohol (v/v) | Bottle | 29050
22 | Disinfecting liquid hand wash (broad spectrum, added emollients) 1 Liter | Liquid soap in bottle with automatic hand soap dispenser,good quality, skin friendly,recommended unscented, antibacterial with added emollients , free of harmful or dangerous substances. Remaining shelf life of 80% Note for external use only, keep away from eyes, mouth and open wounds. Weight: 5 Ltr, Self-life | EA | 27500
23 | Disinfecting liquid hand wash (broad spectrum, added emollients) 5 Liter. Bottle of 5000ml, 1000ml & 500ml Liquid soap, Antibacterial Unscented Quote for each size sperately |  | EA | 125094
24 | Bleach, 1L | 5% chlorine | Bottle | 14020
25 | Eau de javel, 5L | 5% chlorine | EA | 176485
26 | Multipurpose cleaning solution, 1L | Liquid with anionic and non ionic surfactants, other cleaning agents, water. Can be diluted or ready to use. Produces foam | EA | 2801
27 | Antiseptic, Multi action similar to dettol, septol- | Antiseptic 1L or 500 ml X2 | EA | 17692
28 | Household floor detergent (1L) | Size: 500ml Hypoallergenic. Multi-Purpose Cleaner. With its proven anti-bacterial formula, it helps eliminate 99.9% bacteria, leaving your household surfaces clean and shiny | Bottle | 14020
29 | Bacillol Medica, Surface Spry, Alcohol, 500ML or 1L | Surface Spry, Alcohol 75% or sodium hypochlorite 0.1%, 500ML or 1L | EA | 7200
30 | Disinfecting cleaner for surfaces and equipment (broad spectrum, residual activity, fragranced) 5 Liter | Multi-Surface antibacterial cleaning solution, unscented, hypoallergenic Free of harmful or dangerous substances. Alcohol 75% Remaining shelf life of 80% For surface cleaning only Indicating “For external use only, avoid contact with eyes”. Additional Warning/Danger signal sticker stating product content is “flammable". in a bottle with spray dispenser, good quality | EA | 81
SECTION 3: PERSONAL PROTECTIVE EQUIPMENT
31 | Mask, surgic, typeIIR, earloop, disp. | Medical mask, good breathability, internal and external faces should be clearly identified box of 50 masks | Pack of 50 | 538053
32 | Face Mask Disposable 3 Ply with tie (box=50pcs) | Medical mask, good breathability, internal and external faces should be clearly identified | Box | 1080
33 | Mask, surgic, typeI, earloop, disp. | . Bacterial filtering efficiency (BFE): equal to or greater than 95%. . Breathing resistance . Splash resistance . Fabric, non-woven with outer layer . Completely covering nose, mouth and chin. . Clearly identifiable inner and outer surfaces. . Malleable nose strip, made of aluminum, allowing a snug fit. . With attached 2 x 2 tie-straps, allowing correct fixation and securing at the ear loops. | Box of 50 | 33000
34 | Mask, surgic, typeII, earloop, disp. |  | Box of 50 | 15000
35 | Surgical respirator, high-filtration, FFP2/N95, mask, no valve, non-sterile | Material: non-woven filter layer. Meets the requirements of FFP2 or N95 Meets the requirements of type IIR: Bacterial filtration efficiency (BFE) > or = 98%. Shape of the mask: duckbill (folded horizontal width-wise), or cup-shaped. Good breathability with design that does not collapse against the mouth. Strong elastic straps, fitting Non-sterile. | EA | 2279235
36 | Cloth Mask, 3 layers, re-washable | A minimum of three layers is required for non-medical masks. 1) an innermost layer of a hydrophilic material (e.g. cotton or cotton blends); 2), an outermost layer made of hydrophobic material (e.g., polypropylene, polyester, or their blends) 3) a middle hydrophobic layer of synthetic non-woven material such as polyproplylene It is preferable not to select elastic material for making masks; during wear, the mask material may be stretched over the face, resulting in increased pore size and lower filtration efficiency throughout use. Also, elastic materials may degrade over time and are sensitive to washing at high temperatures. Coating the fabric with compounds like wax is not recommended Select washable fabrics that can be washed and that can support high temperatures (60° or more) are preferable. Non-woven polypropylene (PP) spunbond may be washed at high temperatures, up to 125°C.(72) | EA | 1000
37 | Cloth Face mask, ear loop or tie-on | 3 layers filtration: - (minimum 70% solid particle filtration or droplet filtration) breathability: - (maximum pressure difference of 0.6 mbar/cm2 or maximum inhalation resistance of 2.4 mbar and maximum exhalation resistance of 3 mbar). Type of material: - cotton woven Diameter 22 cm minimum Height 12 cm minimum to cove top of the nose till chin Elastic rubber band adjustable Recommended with logo both ear loop and tie-on, Logo in one color to be added upon organization request. For kids and adult | EA | 4033000
38 | Coverall, protection, CatIII, type 6b, L | General description: Spray/aerosol-penetration resistant, biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. Technical specifications Elasticated hood around face. Elasticated cuffs and ankles. Elasticated sleeves. Preferably bound seams. Zipper with re-sealable flap protecting leakage through seams. Stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa (class 2) (or equivalent international standard) Non-sterile Single Use, disposable Size: L Conforms to: European Directive 89/686/EEC on personal protective equipment Category III: Chemical protective coverall, Type 6 comply with EN 13034:2005+A1 (or equivalent) marketing approval certificate. Barrier to infective agent standards: EN 14126:2003 certified passing infectious agent test according to ISO 16604:2004 (Resistance to penetration by blood‐borne pathogens using bacteriophage Phi‐X 174) standard at minimum exposure pressure of 1.75kPa (class 2) (or equivalent international standard) Performance requirement ISO standards: EN 340 – General requirements EN ISO 17491-4: 2008 (supersedes: EN 468 mod. – Spray aerosol test) EN 14325 – Test methods & performance classification EN ISO 3758 (ISO 3758) – Textile care symbols EN ISO 13935-2 – Seam strength Intended use and considerations: Disposable spray/arosol-tight biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. After use, the coverall must be disposed in a biohazard waste container, collected and destroyed. This is also applicable if the coverall is damaged (perforation, etc.). Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Supplier's instruction for use Shelf life: As indicated by supplier Storage conditions: As indicated by supplier | EA | 22784
39 | Coverall, protection, CatIII, type 6b, M | General description: Spray/aerosol-penetration resistant, biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. Technical specifications Elasticated hood around face. Elasticated cuffs and ankles. Elasticated sleeves. Preferably bound seams. Zipper with re-sealable flap protecting leakage through seams. Stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa (class 2) (or equivalent international standard) Non-sterile Single Use, disposable Size: M Conforms to: European Directive 89/686/EEC on personal protective equipment Category III: Chemical protective coverall, Type 6 comply with EN 13034:2005+A1 (or equivalent) marketing approval certificate. Barrier to infective agent standards: EN 14126:2003 certified passing infectious agent test according to ISO 16604:2004 (Resistance to penetration by blood‐borne pathogens using bacteriophage Phi‐X 174) standard at minimum exposure pressure of 1.75kPa (class 2) (or equivalent international standard) Performance requirement ISO standards: EN 340 – General requirements EN ISO 17491-4: 2008 (supersedes: EN 468 mod. – Spray aerosol test) EN 14325 – Test methods & performance classification EN ISO 3758 (ISO 3758) – Textile care symbols EN ISO 13935-2 – Seam strength Intended use and considerations: Disposable spray/arosol-tight biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. After use, the coverall must be disposed in a biohazard waste container, collected and destroyed. This is also applicable if the coverall is damaged (perforation, etc.). Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Supplier's instruction for use Shelf life: As indicated by supplier Storage conditions: As indicated by supplier | EA | 6000
40 | Coverall, protection, CatIII, type 6b, XL | General description: Spray/aerosol-penetration resistant, biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. Technical specifications Elasticated hood around face. Elasticated cuffs and ankles. Elasticated sleeves. Preferably bound seams. Zipper with re-sealable flap protecting leakage through seams. Stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa (class 2) (or equivalent international standard) Non-sterile Single Use, disposable Size: XL Conforms to: European Directive 89/686/EEC on personal protective equipment Category III: Chemical protective coverall, Type 6 comply with EN 13034:2005+A1 (or equivalent) marketing approval certificate. Barrier to infective agent standards: EN 14126:2003 certified passing infectious agent test according to ISO 16604:2004 (Resistance to penetration by blood‐borne pathogens using bacteriophage Phi‐X 174) standard at minimum exposure pressure of 1.75kPa (class 2) (or equivalent international standard) Performance requirement ISO standards: EN 340 – General requirements EN ISO 17491-4: 2008 (supersedes: EN 468 mod. – Spray aerosol test) EN 14325 – Test methods & performance classification EN ISO 3758 (ISO 3758) – Textile care symbols EN ISO 13935-2 – Seam strength Intended use and considerations: Disposable spray/arosol-tight biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. After use, the coverall must be disposed in a biohazard waste container, collected and destroyed. This is also applicable if the coverall is damaged (perforation, etc.). Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Supplier's instruction for use Shelf life: As indicated by supplier Storage conditions: As indicated by supplier | EA | 22784
41 | Coverall, TYVEK or equivalent | Breathable Fabric providing coolness, water resistant/repellent and quick drying, stretch knit cuffs safely cover the wrist area. | EA | 648500
42 | Goggles, protective, indirect-side-venti | Good seal with the skin of the face, Flexible PVC frame to easily fit with all face contours with even pressure, Enclose eyes and the surrounding areas, Accomodate wearers with prescription glasses, Clear plastic lens with fog and scratch resistant treatments, Adjustable band to secure firmly so as not to become loose during clinical activity, Indirect venting to avoid fogging, May be re-usable (provided appropriate arrangements for decontamination are in place) or disposable. | EA | 452363
43 | Faceshield, fog-resistant, fullface, disp | Made of clear plastic and provides good visibility to both the wearer and the patient, Adjustable band to attach firmly around the head and fit snuggly against the forehead, Fog resistant (preferable), Completely cover the sides and length of the face, May be re-usable (made of robust material which can be cleaned and disinfected) or disposable. For kids and adult | EA | 382985
44 | Gloves,surgical, Sterile, length to forearm large (longer than examination gloves), latex, disp. - size 7.5 | Gloves, surgical, latex, sterile, powder-free, single use. Gloves should have long cuffs, reaching well above the wrist, ideally to mid-forearm. minimum thickness 0.10mm. Sizes | EA | 20 Pairs
45 | Gloves, w/o powder, nitrile, L, disp,box/100 | Gloves, examination, nitrile, powder-free, non-sterile. (eg. minimum 230mm total length. Sizes, minimum thickness 0.05mm | Box of 100 | 136485
46 | Gloves, w/o powder, nitrile, M, disp, box/100 |  | Box of 100 | 146564
47 | Gloves, w/o powder, nitrile, S, disp, box/100 |  | Box of 100 | 21200
48 | Gloves, w/o powder, Latex, XL, disp, box/100 | Gloves, examination, LATEX, powder-free, non-sterile. Cuff length preferably reaching above the wrist, minimum thickness 0.05mm. | Box of 100 | 168
49 | Gloves, w/o powder, Latex, L, disp, box/100 |  | Box of 100 | 60445
50 | Gloves, w/o powder, Latex, M, disp, box/100 |  | Box of 100 | 48510
51 | Gloves, w/o powder, Latex, S, disp, box/100 |  | Box of 100 | 300
52 | HE*Gloves, heavy-duty, rubber/nitrile, L | Outer glove should have long cuffs, reaching well above the wrist, ideally to mid-forearm. Cuff length preferably reach mid-forearm (eg. minimum 280mm total length).Reusable | EA | 150542
53 | HE*Gloves, heavy-duty, rubber/nitrile, M |  | EA | 59315
54 | HE*Gloves, heavy-duty, rubber/nitrile, S |  | EA | 15000
55 | Gown, surgic, nonsterile, nonwoven, disp, L | Single use, long sleeves, fiber made non woven, thumb loop, tape tab for neck closure, water and liquids proof, compliant with the EN 13795 high performance level, or AAMI level 3 performance or equivalent | EA | 623947
56 | Gown, surgic, nonsterile, nonwoven, disp, XL |  | EA | 619947
57 | *HE Apron, protect, plastic,disp/PAC-100 | *HE Apron,protect,plastic,disp/PAC-100 Technical Specifications: Single-use straight sleeveless protective apron, for use in healthcare settings Seamless liquid proof and stain resistant Comfortable to wear, apron has back- and neck-band strips attached (4 in total) Both back- and neck-band can be adjusted/fastened Color: white Material: polyethylene (PE) or biodegrdable or compostable material Size: 85 x 145 cm (w x l) (+/- 15%) Thickness, at not less than: 50 um Can resist water and disinfectant (ethanol 70% and chlorine solution 0.5%) Product performance testing if biodegradable: EN 13432, ASTM D6400 or alternative equivalent set of standards Packaging and labelling: Primary packaging: One (1) pack r roll of 100 aprons. Intended use and considerations Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. The size has been chosen as the most commonly used. The protection apron is single use only. After use, dispose used aprons in waste container, collect and destroy them. Please refer to WHO publication "Safe Management of Waste for Health Care Activities". Product attributes | Pack of 100 | 60
58 | *HE Apron, protection, plastic, reusable | Straight apron with bib, Fabric: 100% polyester with PVC coating, or 100% PVC, or 100% rubber, or other fluid resistant coated material, Waterproof, Sewn strap for neck and back fastening Minimum basis weight: 300g/m2 covering size: 70-90 cm (width) X 120-150cm (height) Reusable (provided appropriate arrangements for decontamination are in place) | EA | 300
59 | Cover Shoes(box=100pcs) | Material: PVC, non-conductive, disposable, non-woven, elastic in hem at ankle, non-permeable to liquid, waterproof, one-size-fits-all. Single use, disposable. | Box | 54489
60 | Disposable covershoes | Length should cover the foot till the knee, with tie or rubber | 100 m | 288500
61 | Head Cap Disposable (Box 100pcs) | A non-sterile head covering designed as a cap to completely cover the hair. Elasticated cap made of non-woven materials. Non-woven (polypropylene, viscose, etc.) Non-permeable to liquid. Waterproof. For medical use. Elastic opening permitting complete coverage of all hairstyles Latex-free. One-size-fits-all. Non sterile, single use. | Box of 100 | 891500