## Instructions
Annex 9. Technical Information and Price Bid Form
1. The structure of Annex 9: Tab 1. Product List and Price Bid Form Tab 2. Technical Information Form for the Pharmaceuticals Tab 3. Technical Information Form for the Medical Devices Tab 4. Technical Information Form for In Vitro Diagnostic Tests Tab 5. Technical Specification Form for the Medical Devices and IVDs Tab 6. Technical Specification Form for Other Items Tab 7. Prices for the IARH kits Tab 8. Prices for the Other Medical Kits Tab 9. Prices for Data Loggers and Cold Boxes Bidders should fill out the tabs in this file and submit this form in Excel with unlocked cells. Tab 1, 7, 8 should be submitted in both signed pdf and excel versions. 2. Bidders provide clear and complete information and documents for each lot the company is bidding for. For Lot 1: Pharmaceuticals - Tabs #1, 2, 9 should be filled in and submitted. Bidders must quote for as many products as possible but not less than 80% of the total list. For Lot 3: Medical Devices and IVDs - Tabs #1, 3, 4, 5, 9 should be filled in and submitted. Bidders must quote for as many products as possible but not less than 800% of the total list. For Lot 4: IARH Kits - Tabs #1-7, 9 should be filled in and submitted. In order to be considered a prospective supplier of IARH Kits, ALL the products in each kit should be offered, and ALL the IARH Kits should be quoted for. The information for the components of IARH kits should be provided in the respective tabs. For Lot 5: Other Kits - Tabs #1-6, 8, 9 should be filled in and submitted. In order to be considered a prospective supplier of Other Medical Kits, ALL the products in each kit should be offered, and ALL the Other Medical Kits Kits should be quoted for. The information for the components of kits should be provided in the respective tab 3. In instances where a Bidder submits products whose attributes differ slightly from the specifications, Bidders should highlight these deviations in bold/red in the text. Products that meet all UNFPA’s specifications will be given priority.

## 1. Product list-price form
Tab 1. Product List and Price Bid Form
| Name of Company (Please enter in the cell on the right): | to be filled-in
| FCA Bidders Warehouse Location (Please enter in the cell on the right): | to be filled-in
| FCA Bidders closest port and airport (Please enter in the cell on the right): | to be filled-in
| Please indicate lots for which your company is bididng | to be filled-in
Other comments (if any):
Important guidance to Bidders:
1. Please fill in the blanks under yellow and green columns.
2. Product List and Price Form should be submitted in both excel and signed PDF version.
3. Clarification on packaging sizes: Unit of measurement (UOM) refers to the tablet, capsule, vial, ampoule, etc.; primary packaging is the one that is in direct contact with the product: vial, amp, blister, jar. Secondary packaging is the sale pack, usually cardboard boxes.
4. Color coding: Items included in Lot 4: Interagency Reproductive Health Kits (IARH kits) and complementary kits, Emergency PEP Kit for HIV Treatment are marked and terms included in Lot 5 Other Kits are marked as following:
| Lot 4: IARH Kits | Lot 5: Other Kits | Items in both Lot 4: IARH Kits & Lot 5: Other Kits
Please note that the proposed products should align with quantities in the respective kit compositions if the bidder bids for kits.
5. Weighted scoring will be used during the commercial and financial evaluation per each lot separately, where the weight score is given per each item based on past procurement levels. The total score for each Bidder will be the sum of weighted scores for all quoted items, which will identify the lowest-priced bids that will be passed to further Technical Evaluation. Please note that for pharmaceutical items where several strengths or dosages are allowed, the price comparison will be done for IU or mg/mcg of API.
6. Please list the secondary sources, where applicable, at the end under the sub-section "secondary sources".
7. Columns highlighted in blue are linked automatically to tabs 7 and 8.
PRODUCT LIST | [5 empty] | Product Information |  |  |  |  | FINANCIAL INFORMATION
Bid Item No. | Category | Product sub-category | UNFPA Item Description | Supplier's sales pack description including details of the UOM, primary and secondary packaging MAY NOT BE LEFT BLANK AND SHOULD MATCH ONE OF THE OPTIONS IN COLUMN D | Supplier's sales pack size IN NUMBERS to ensure proper sales pack calculation MAY NOT BE LEFT BLANK AND SHOULD MATCH ONE OF THE OPTIONS IN COLUMN D | Brand name | Manufacturer(s) name | Manufacturing country (Country of origin) | Name and address(es) of ALL manufacturing site(s) of the finished product(s), including address of site of final product release. | Storage conditions | Minimum Order Quantity (based on Supplier's Sales Pack Size) | Multiple Order Quantity (for products ordered in pre-set multiple quantities) (based on Supplier's Sales Pack Size) | Maximum Order Quantity for the offered Lead Time (based on Supplier's Sales Pack Size) | Bidder's Price per UOM for delivery FCA closest port/airport PALLETIZED (USD) | FCA Bidders Warehouse Price per Sales Pack (USD) | FCA Bidders Warehouse Price per UOM (tablet, caps, vial, amp, syringe, etc.) (USD) | Weighting score | Calculated automatically weighted price per item | Shelf life in months | Delivery Lead Time FCA (Days) | Please provide information about discounts if any | Bidder's comments
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 12 | 14 | 15 | 16=6*17 | 17 | 18 | 19=17*18 | 20 | 21 | 22 | 23
1 | Pharmaceuticals | Orals | Amoxicillin, powder oral suspension, 125mg/5ml, 100 ml, pack of 10 or 40 bottles | [11 empty] | 0 |  | 0.2209% | 0
2 | Pharmaceuticals | Orals | Amoxicillin trihydrate, 250mg, packs of 15, 20, 21, 50, 100, 500 or 1000 capsules | [11 empty] | 0 |  | 0.2209% | 0
3 | Pharmaceuticals | Orals | Amoxicillin trihydrate, 500mg, packs of 15, 20, 21, 100, 500 or 1000 bottles or 15 or 21 blister strips | [11 empty] | 0 |  | 0.2209% | 0
4 | Pharmaceuticals | Orals | Amoxicillin trihydrate suspension, 250mg/5ml, in 100ml, pack of 10 or 40 bottles | [11 empty] | 0 |  | 0.2209% | 0
5 | Pharmaceuticals | Parenterals | Ampicillin sodium, powder for injection, 1000mg (1g), pack of 1, 2, 5, 10, 20, 25 or 50 vials | [11 empty] | 0 |  | 0.2209% | 0
6 | Pharmaceuticals | Parenterals | Ampicillin sodium, powder for injection, 500mg, pack of 50 vials | [11 empty] | 0 |  | 0.2209% | 0
7 | Pharmaceuticals | Parenterals | Anti-D (RhO) immunoglobulin, 300mcg (winrho SDF), pack of 1 or 5 vials | [11 empty] | 0 |  | 2.963% | 0
8 | Pharmaceuticals | Orals | Atazanavir (ATV) 300mg + Ritonavir (r) 100mg, pack of 30 tablets | [11 empty] | 0 |  | 0.2209% | 0
9 | Pharmaceuticals | Parenterals | Atropine sulphate 1mg/ml injection in 1ml ampoule, pack or 10, 20, 50 or 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
10 | Pharmaceuticals | Orals | Abacavir (as sulphate) 120mg + Lamivudine 60mg, pack of 60 dispersible tablets | [11 empty] | 0 |  | 0.2209% | 0
11 | Pharmaceuticals | Orals | Azithromycin Anhydrous, 250mg, pack of 1, 6, 10 or 12 capsules | [11 empty] | 0 |  | 0.2209% | 0
12 | Pharmaceuticals | Orals | Azithromycin, 250mg, pack of 4 or 6 tablets | [11 empty] | 0 |  | 0.2209% | 0
13 | Pharmaceuticals | Orals | Azithromycin, 500mg, pack of 3 or 500 tablets | [11 empty] | 0 |  | 0.2209% | 0
14 | Pharmaceuticals | Orals | Azithromycin dihydrate, 200mg base/ 5ml suspension, 15ml bottle | [11 empty] | 0 |  | 0.2209% | 0
15 | Pharmaceuticals | Parenterals | Benzathine benzylpenicillin, 1.44G (2.4 M.I.U) powder for injection, 5ml or 10ml, pack of 1, 10, 50 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
16 | Pharmaceuticals | Parenterals | Benzathine benzylpenicillin, 900mg (1.2 M.I.U) powder for injection, 5ml or 10ml, pack of 1, 10, 50 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
17 | Pharmaceuticals | Parenterals | Betamethasone sodium phosphate 3mg base/ml + betamethasone acetate 3mg base/ml in an aqueous vehicle, pack of 20 vials | [11 empty] | 0 |  | 0.2209% | 0
18 | Pharmaceuticals | Parenterals | Bupivacaine hydrochloride 0.5%, 10 ml, pack of 10, 20, 50 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
19 | Pharmaceuticals | Parenterals | Bupivacaine hydrochloride 0.5%, 20 ml, pack of 10, 20, 50 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
20 | Pharmaceuticals | Parenterals | Bupivacaine hydrochloride 0.5% 4ml ampoule + 7.5% glucose solution, pack of 5, 10 or 20 ampoules | [11 empty] | 0 |  | 0.2209% | 0
21 | Pharmaceuticals | Parenterals | Calcium gluconate, 100mg base/ml injection, 10ml, pack of 10, 20 or 50 ampoules | [11 empty] | 0 |  | 1.843% | 0
22 | Pharmaceuticals | Parenterals | Carbetocin, 100mcg/ml in 1 ml ampoule, pack of 10, 20 or 50 ampoules | [11 empty] | 0 |  | 0.2209% | 0
23 | Pharmaceuticals | Orals | Cefixime, 200mg, pack of 8 or 10 tablets | [11 empty] | 0 |  | 0.2209% | 0
24 | Pharmaceuticals | Orals | Cefixime, 400mg, pack of 8, 10 or 100 tablets | [11 empty] | 0 |  | 2% | 0
25 | Pharmaceuticals | Orals | Cefixime trihydrate, 100mg/5ml, powder oral suspension, bottle of 30ml, 40ml or 50ml | [11 empty] | 0 |  | 0.2209% | 0
26 | Pharmaceuticals | Orals | Cephalexin, 500mg, pack of 21, 100 or 500 capsules | [11 empty] | 0 |  | 0.2209% | 0
27 | Pharmaceuticals | External Use | Chlorhexidine digluconate 7.1% w/v solution, bottle of 10ml with a dripper | [11 empty] | 0 |  | 0.2209% | 0
28 | Pharmaceuticals | External Use | Chlorhexidine digluconate solution 5%, 1 litre bottle | [11 empty] | 0 |  | 0.2209% | 0
29 | Pharmaceuticals | External Use | Chlorhexidine digluconate 7.1% (equivalent to 4% chlorhexidine) gel, 10g or 20g, aluminium tube | [11 empty] | 0 |  | 0.2209% | 0
30 | Pharmaceuticals | External Use | Chlorhexidine gluconate 7.1% (equivalent to 4% chlorhexidine) gel, 3g, pack of 7 plastic individual sachets | [11 empty] | 0 |  | 0.2209% | 0
31 | Pharmaceuticals | External Use | Chlorhexidine gluconate solution 4.0% (detergent solution), 500ml bottle | [11 empty] | 0 |  | 0.2209% | 0
32 | Pharmaceuticals | Orals | Ciprofloxacin, 500mg, pack of 10 or 100 tablets | [11 empty] | 0 |  | 0.2209% | 0
33 | Pharmaceuticals | Parenterals | Clindamycin injection, 125mg/ml in 4 ml ampoule/vial (600mg/4ml), pack of 1, 5, 10, 20 or 50 vials/ampoules | [11 empty] | 0 |  | 0.2209% | 0
34 | Pharmaceuticals | Vaginal | Clotrimazole, 100mg, vaginal tablet | [11 empty] | 0 |  | 0.2209% | 0
35 | Pharmaceuticals | Vaginal | Clotrimazole, 500mg, vaginal tablet | [11 empty] | 0 |  | 0.2209% | 0
36 | Pharmaceuticals | Parenterals | Cloxacillin, as sodium salt, 500mg, pack of 10 or 50 vials | [11 empty] | 0 |  | 0.2209% | 0
37 | Pharmaceuticals | Parenterals | Dexamethasone sodium phosphate, 4mg base/ml, 1ml, pack of 10, 50 and 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
39 | Pharmaceuticals | External Use | Disinfectant tablets for water, 1.67g NaDCC, pack of 10, 100 or 200 tablets | [11 empty] | 0 |  | 0.2209% | 0
40 | Pharmaceuticals | Orals | Doxycycline, 100mg, pack of 8, 10 or 12 tablets | [11 empty] | 0 |  | 0.2209% | 0
41 | Pharmaceuticals | Orals | Dolutegravir, 50mg, pack of 30 film coated tablets | [11 empty] | 0 |  | 0.2209% | 0
42 | Pharmaceuticals | Orals | Dolutegravir, 10mg, pack of 90 dispersible tablets | [11 empty] | 0 |  | 0.2209% | 0
43 | Pharmaceuticals | Orals | Efavirenz 600mg+Emtricitabine 200mg+Tenofovir 300mg, pack of 28, 30 or 500 tablets | [11 empty] | 0 |  | 0.2209% | 0
44 | Pharmaceuticals | Orals | Emtricitabine 200mg + Tenofovir 300mg, pack of 30 tablets | [11 empty] | 0 |  | 0.2209% | 0
45 | Pharmaceuticals | Parenterals | Ephedrine hydrochloride injection, 30 mg/ml, pack of 5 or 10 ampoules | [11 empty] | 0 |  | 0.2209% | 0
46 | Pharmaceuticals | Parenterals | Epinephrine (Adrenaline) as hydrochloride or hydrogen tartrate injection, 1mg/ml in 1 ml ampoule, pack of 10 ampoules | [11 empty] | 0 |  | 0.2209% | 0
47 | Pharmaceuticals | Parenterals | Ergometrine maleate, 0.2mg base/ml injection in 1ml ampoule, pack of 6 or 10 ampoules | [11 empty] | 0 |  | 0.2209% | 0
48 | Pharmaceuticals | Orals | Ferrous salts/Folic acid coated (60mg/0.4mg), pack of 100 or 1000 tablets | [11 empty] | 0 |  | 2.547% | 0
49 | Pharmaceuticals | Orals | Ferrous salts (60mg Fe), pack of 50, 100, 250, 500 or 1000 sugar coated tablets | [11 empty] | 0 |  | 0.2209% | 0
50 | Pharmaceuticals | Orals | Folic acid, 5mg, pack of 7, 10, 14, 21, 28, 30, 56, 60, 84, 90, 100, 112, 500 or 1000s tablets | [11 empty] | 0 |  | 0.2209% | 0
51 | Pharmaceuticals | Parenterals | Gentamicin sulphate, 10mg/ml injection in 2ml ampoule, pack of 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
52 | Pharmaceuticals | Parenterals | Gentamicin sulphate, 40mg base/ml injection in 2ml ampoule, pack of 10, 50 or 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
54 | Pharmaceuticals | Parenterals | Glucose, 10% injection in pack of 20 or 50 ampoules | [11 empty] | 0 |  | 0.2209% | 0
55 | Pharmaceuticals | Parenterals | Glucose, solution for infusion, 5% (iso-osmotic), pack of 30 or 36 bags of 250 ml, pack of 20 or 24 bags for 500 ml or pack of 10 or 12 bags for 1000 ml+ Infusion-giving set, sterile, single use | [11 empty] | 0 |  | 0.2209% | 0
56 | Pharmaceuticals | Parenterals | Hartmann's sol. (Ringer's lactate) intravenous infusion, 0.5L + giving set, 1 set or pack of 20 sets | [11 empty] | 0 |  | 0.2209% | 0
57 | Pharmaceuticals | Parenterals | Hartmann's sol. (Ringer's lactate) intravenous infusion, 1L + giving set, 1 set or pack of 12 sets | [11 empty] | 0 |  | 0.2209% | 0
58 | Pharmaceuticals | Parenterals | Heparin sodium 5000IU/ml neutral glass ampoules (Type I Ph Eur) of 1ml capacity OR 5ml capacity containing 1ml OR 5ml of solution respectively | [11 empty] | 0 |  | 0.2209% | 0
59 | Pharmaceuticals | Parenterals | Hydralazine hydrochloride 20mg/2ml powder for injection in 2ml ampoule | [11 empty] | 0 |  | 0.2209% | 0
60 | Pharmaceuticals | Orals | Hydralazine hydrochloride, 25mg, pack of 28, 30, 56, 60, 84, 90, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | [11 empty] | 0 |  | 0.2209% | 0
61 | Pharmaceuticals | Orals | Hydralazine hydrochloride, 50mg, pack of 28, 30, 56, 60, 84, 90, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | [11 empty] | 0 |  | 0.2209% | 0
62 | Pharmaceuticals | Orals | Ibuprofen, 400mg, pack of 6, 10, 12, 15, 18, 20, 24, 30, 40, 50, 56, 60, 84 or 100 | [11 empty] | 0 |  | 0.2209% | 0
63 | Pharmaceuticals | Parenterals | Ketamine hydrochloride, 50mg base/ml injection in 10ml vial, pack of 1, 5 or 25 vials | [11 empty] | 0 |  | 0.2209% | 0
64 | Pharmaceuticals | Orals | Lamivudine, 150mg, pack of 10 or 60 tablets | [11 empty] | 0 |  | 0.2209% | 0
65 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 1% injection in 10ml vial, pack of 20 vials | [11 empty] | 0 |  | 0.2209% | 0
66 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 1% injection in 20ml vial/ampoule, pack of 1, 10 or 20 vials/ampoules | [11 empty] | 0 |  | 0.2209% | 0
67 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 1% injection in 50ml vial, pack of 10 or 25 vials | [11 empty] | 0 |  | 0.2209% | 0
68 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 2% injection in 20ml vial, pack of 10 or 20 vials | [11 empty] | 0 |  | 0.2209% | 0
69 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 2% injection in 50ml vial, pack of 25 vials | [11 empty] | 0 |  | 0.2209% | 0
70 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 2% injection in 5ml vial, pack of 10, 20 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
71 | Pharmaceuticals | Parenterals | Lidocaine hydrochloride 5% (50mg/ml) injection in 2ml ampoule, pack of 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
72 | Pharmaceuticals | Parenterals | Magnesium sulphate 500mg/ml injection in 10ml ampoule, pack of 1, 10, 50 or 100 ampoules | [11 empty] | 0 |  | 18.48% | 0
73 | Pharmaceuticals | Parenterals | Magnesium Sulphate 500mg/ml injection in 2ml ampoule, pack of 10 ampoules | [11 empty] | 0 |  | 2.845% | 0
74 | Pharmaceuticals | Orals | Methyldopa, 250mg, packs of 28, 56, 60, 84, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | [11 empty] | 0 |  | 0.2209% | 0
75 | Pharmaceuticals | Orals | Metronidazole, 250mg, pack of 50, 100, 250 or 500 tablets | [11 empty] | 0 |  | 0.2209% | 0
76 | Pharmaceuticals | Parenterals | Metronidazole, for intravenous infusion, 5mg/ml, 100-ml, pack of 1 or 20 PVC free flexi-bags | [11 empty] | 0 |  | 0.2209% | 0
77 | Pharmaceuticals | Parenterals | Metronidazole, for intravenous infusion, 5mg/ml, 100-ml, semi-rigid PVC free bottle | [11 empty] | 0 |  | 0.2209% | 0
78 | Pharmaceuticals | Orals | Mifepristone 200mg + 4 misoprostol 200mcg tablets (blister), pack of 5 (1mife+4miso) tablets | [11 empty] | 0 |  | 4.14% | 0
79 | Pharmaceuticals | Orals | Mifepristone, 200mg, pack of 1, 3, 15, 30 tablets | [11 empty] | 0 |  | 0.2209% | 0
80.1 | Pharmaceuticals | Orals | Misoprostol, 200mcg, pack of 3 tablets | [11 empty] | 0 |  | 6.358% | 0
80.2 | Pharmaceuticals | Orals | Misoprostol, 200mcg, pack of 4 tablets | [11 empty] | 0 |  | 6.358% | 0
81 | Pharmaceuticals | Orals | Multivitamin coated, pack of 20, 84, 90, 100 or 500 tablets Vitamin A (retinol acetate) -800 micrograms retinol activity equivalent Vitamin C (ascorbic acid) -70 mg Vitamin D (cholecalciferol) - 5 micrograms (200 IU) Vitamin E (alpha tocopherol succinate) -10 mg alpha tocopherol equivalent Vitamin B1 (thiamine mononitrate) -1.4 mg Vitamin B2 (riboflavin) -1.4 mg Vitamin B3 (niacinamide) -18 mg niacin equivalent Vitamin B6 (pyridoxine hydrochloride) -1.9 mg Folic acid (folic acid) - 680 micrograms dietary folate equivalent (400 micrograms) Vitamin B12 (cyanocobalamin) - 2.6 micrograms Iron (ferrous fumarate) - 30 mg Iodine (potassium iodide) - 150 micrograms Zinc (zinc oxide) - 15 mg Selenium (sodium selenite) - 65 micrograms Copper (cupric oxide) - 2 mg | [11 empty] | 0 |  | 0.2209% | 0
82 | Pharmaceuticals | Orals | Nifedipine, 10mg, pack of 90 or 100 immediate release capsule | [11 empty] | 0 |  | 0.2209% | 0
83 | Pharmaceuticals | Orals | Oxybutynin hydrochloride, 5mg, pack of 6, 20, 21, 28, 30, 50, 56, 60, 84 or 100 tablets | [11 empty] | 0 |  | 0.2209% | 0
84 | Pharmaceuticals | Parenterals | Oxytocin 10 I.U. base/ml injection in 1ml ampoule, pack of 5, 10 or 100 ampoules | [11 empty] | 0 |  | 30.32% | 0
85 | Pharmaceuticals | Orals | Paracetamol 500mg, tablet | [11 empty] | 0 |  | 0.2209% | 0
86 | Pharmaceuticals | External Use | Povidone iodine 10% solution for cutaneous use, 1000ml bottle | [11 empty] | 0 |  | 0.2209% | 0
87 | Pharmaceuticals | External Use | Povidone iodine 10% solution for cutaneous use, 500ml bottle | [11 empty] | 0 |  | 0.2209% | 0
88 | Pharmaceuticals | Parenterals | Quinine dihydrochloride 300 mg/ml, solution for dilution for infusion in 2 ml ampoule, pack of 10, 50 or 100 ampoules | [11 empty] | 0 |  | 0.2209% | 0
89 | Pharmaceuticals | Parenterals | Sodium Chloride 0.9% injection 10 ml ampoule | [11 empty] | 0 |  | 0.2209% | 0
90 | Pharmaceuticals | Parenterals | Sodium chloride 0.9% isotonic IV infusion in 1000ml bag + Infusion-giving set, sterile, single use, pack of 1 or 20 sets | [11 empty] | 0 |  | 0.2209% | 0
91 | Pharmaceuticals | Parenterals | Sodium chloride 0.9% isotonic IV infusion in 500ml bag + Infusion-giving set, sterile, single use, pack of 1, 10, 12 or 20 sets | [11 empty] | 0 |  | 0.2209% | 0
92 | Pharmaceuticals | Orals | Tenofovir Disoproxil Fumarate 300mg + Lamivudine 300mg + Dolutegravir 50mg, pack of 30, 90 or 180 tablets | [11 empty] | 0 |  | 0.2209% | 0
93 | Pharmaceuticals | External Use | Tetracycline hydrochloride eye-ointment 1% in tube of 5g - CANCELLED: TO BE REPLACED (no need to quote) | [11 empty] | 0 |  | 0.2209% | 0
94 | Pharmaceuticals | Parenterals | Tranexamic acid injection 100mg/ml 10ml ampoule, pack of 1, 5 or 10 ampoules | [11 empty] | 0 |  | 2.268% | 0
95 | Pharmaceuticals | Parenterals | Tranexamic acid injection 100mg/ml 5ml ampoule, pack of 1, 5 or 10 ampoules | [11 empty] | 0 |  | 0.2209% | 0
96 | Pharmaceuticals | Parenterals | Vitamin K injection 2mg/0.2ml in 0.2 ml vial, pack of 5 vials | [11 empty] | 0 |  | 0.2209% | 0
97 | Pharmaceuticals | Parenterals | Vitamin K-1 (Phytomenadione) 1mg/ml, 1ml vial, pack of 10 or 100 vials | [11 empty] | 0 |  | 0.2209% | 0
98 | Pharmaceuticals | Parenterals | Water for injection in 10ml, pack of 20 or 50 plastic ampoules | [11 empty] | 0 |  | 0.2209% | 0
100 | Pharmaceuticals | Orals | Zidovudine 300MG + 150mg Lamivudine, pack of 60 tablets | [11 empty] | 0 |  | 0.2209% | 0
101 | Medical devices | Anaesthesia & Resuscitation Equipment | Anaesthesia machine | [11 empty] | 0 |  | 11.34% | 0
102 | Medical devices | Anaesthesia & Resuscitation Equipment | Aspirator, surgical, suction pump, w/accessories | [11 empty] | 0 |  | 1.13% | 0
103 | Medical devices | Anaesthesia & Resuscitation Equipment | Incubator, Infant, automatic | [11 empty] | 0 |  | 3.29% | 0
104 | Medical devices | Anaesthesia & Resuscitation Equipment | Infusion pump, with accessories | [11 empty] | 0 |  | 0.55% | 0
105 | Medical devices | Anaesthesia & Resuscitation Equipment | Pulse oximeter, portable and accessories | [11 empty] | 0 |  | 1.72% | 0
106 | Medical devices | Anaesthesia & Resuscitation Equipment | Oxygen concentrator, set | [11 empty] | 0 |  | 2.05% | 0
107 | Medical devices | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operating, child/newborn | [11 empty] | 0 |  | 0.49% | 0
108 | Medical devices | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operating, neonate, set | [11 empty] | 0 |  | 0.57% | 0
109 | Medical devices | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operated, adult, set | [11 empty] | 0 |  | 0.05654% | 0
110 | Medical devices | Anaesthesia & Resuscitation Equipment | Table, resuscitation, newborn, w/accessories | [11 empty] | 0 |  | 2.15% | 0
111 | Medical devices | Anaesthesia & Resuscitation Equipment | Pump, suction, electric, aspirator, portable, w/accessories | [11 empty] | 0 |  | 1.24% | 0
112 | Medical devices | Anaesthesia & Resuscitation Equipment | Pump, suction, foot-operated, aspirator | [11 empty] | 0 |  | 0.5% | 0
113 | Medical devices | Anaesthesia & Resuscitation Equipment | Syringe pump | [11 empty] | 0 |  | 0.05654% | 0
114 | Medical devices | Medical Electrical Equipment | Cryosurgical Unit | [11 empty] | 0 |  | 0.05654% | 0
115 | Medical devices | Medical Electrical Equipment | Electrosurgical unit, w/accessories | [11 empty] | 0 |  | 1.4% | 0
116 | Medical devices | Medical Electrical Equipment | Light, examination, mobile | [11 empty] | 0 |  | 3.25% | 0
117 | Medical devices | Medical Electrical Equipment | Torch, flashlight, pre-focused, angled, medical | [11 empty] | 0 |  | 0.05654% | 0
118 | Medical devices | Medical Electrical Equipment | Light, operating, mobile | [11 empty] | 0 |  | 0.82% | 0
119 | Medical devices | Medical Electrical Equipment | Monitor, patient, portable,w/access | [11 empty] | 0 |  | 4.87% | 0
120 | Medical devices | Medical Electrical Equipment | Refrigerator, compression | [11 empty] | 0 |  | 0.58% | 0
121 | Medical devices | Medical Electrical Equipment | Warmer, baby, electric | [11 empty] | 0 |  | 2.12% | 0
122 | Medical devices | Medical Sterilization Equipment | Indicator, TST control, spot, adhesive, P1 | [11 empty] | 0 |  | 0.05654% | 0
123 | Medical devices | Medical Sterilization Equipment | Indicator, TST control, spot, adhesive, P1050 | [11 empty] | 0 |  | 0.05654% | 0
124 | Medical devices | Medical Sterilization Equipment | Basket, sterilizing, approx. 120 x 250 x 60 mm | [11 empty] | 0 |  | 0.05654% | 0
125 | Medical devices | Medical Sterilization Equipment | Basket, instruments, for sterilization, wired, 400 x 200 x 90 mm | [11 empty] | 0 |  | 0.05654% | 0
126 | Medical devices | Medical Sterilization Equipment | Drum, sterilizing, diameter 165mm | [11 empty] | 0 |  | 0.05654% | 0
127 | Medical devices | Medical Sterilization Equipment | Drum, sterilizing, diameter 260mm | [11 empty] | 0 |  | 0.05654% | 0
128 | Medical devices | Medical Sterilization Equipment | Drum, sterilizing, diameter 290mm | [11 empty] | 0 |  | 0.05654% | 0
129 | Medical devices | Medical Sterilization Equipment | Drum, sterilizing, diameter 340mm | [11 empty] | 0 |  | 0.05654% | 0
130 | Medical devices | Medical Sterilization Equipment | Sterilizer, steam, pressure-type, 24 litres | [11 empty] | 0 |  | 1.89% | 0
131 | Medical devices | Medical Sterilization Equipment | Sterilizer, steam, pressure-type, 39 litres | [11 empty] | 0 |  | 2.57% | 0
132 | Medical devices | Medical Sterilization Equipment | Stove, kerosene, single-burner, pressure-type | [11 empty] | 0 |  | 0.05654% | 0
133 | Medical devices | Medical Sterilization Equipment | Timer, 60 min, mechanical | [11 empty] | 0 |  | 0.05654% | 0
134 | Medical devices | PPEs | Apron, surgical, disposable, water resistant, non-sterile | [11 empty] | 0 |  | 0.05654% | 0
135 | Medical devices | PPEs | Gloves, exam, nitrile, powder free, non-sterile, large, P100 | [11 empty] | 0 |  | 0.05654% | 0
136 | Medical devices | PPEs | Gloves, exam, nitrile, powder free, non-sterile, medium,P100 | [11 empty] | 0 |  | 0.05654% | 0
137 | Medical devices | PPEs | Gown, surgical, non-woven, sterile, disposable | [11 empty] | 0 |  | 0.05654% | 0
138 | Medical devices | PPEs | Medical Mask, type I, for patients, disposable | [11 empty] | 0 |  | 0.05654% | 0
139 | Medical devices | PPEs | Respirator, FFP2/N95/KN95, mask, no valve, disposable | [11 empty] | 0 |  | 0.05654% | 0
140 | Medical devices | PPEs | Surgical Mask, type IIR, healthcare worker, disposable | [11 empty] | 0 |  | 0.05654% | 0
141 | Medical devices | PPEs | Goggles, panoramic, regular nose, indirect ventilation | [11 empty] | 0 |  | 0.05654% | 0
142 | Medical devices | PPEs | Coverall, L, 3b/4b/5b/6b, disposable Coverall, M, 3b/4b/5b/6b, disposable Coverall, XL, 3b/4b/5b/6b, disposable | [11 empty] | 0 |  | 0.05654% | 0
143 | Medical devices | PPEs | Faceshield, fog-resistant, disposable | [11 empty] | 0 |  | 0.05654% | 0
144 | Medical devices | PPEs | Cap,surgical,bouffant,non-woven | [11 empty] | 0 |  | 0.05654% | 0
145 | Medical devices | Medical supplies | Airway, Guedel, paediatric, size 0 (50mm) | [11 empty] | 0 |  | 0.05654% | 0
146 | Medical devices | Medical supplies | Airway, Guedel, neonate, size 00 (35-40mm) | [11 empty] | 0 |  | 0.05654% | 0
147 | Medical devices | Medical supplies | Airway, Guedel, paediatric, size 1 (60mm) | [11 empty] | 0 |  | 0.05654% | 0
148 | Medical devices | Medical supplies | Airway, Guedel, child, size 2 (70mm) | [11 empty] | 0 |  | 0.05654% | 0
149 | Medical devices | Medical supplies | Airway, Guedel, adolescent, size 3 (80mm) | [11 empty] | 0 |  | 0.05654% | 0
150 | Medical devices | Medical supplies | Airway, Guedel, adult, size 4 (90mm) | [11 empty] | 0 |  | 0.05654% | 0
151 | Medical devices | Medical supplies | Airway, Guedel, adult, size 5 (100mm) | [11 empty] | 0 |  | 0.05654% | 0
152 | Medical devices | Medical supplies | Amniohook, plastic, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
153 | Medical devices | Medical supplies | Amniohook, plastic, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
154 | Medical devices | Medical supplies | Non-pneumatic anti-shock garment | [11 empty] | 0 |  | 1.24% | 0
155 | Medical devices | Medical supplies | Bag, Biohazard, yellow, 50 liters, P50 | [11 empty] | 0 |  | 0.05654% | 0
156 | Medical devices | Medical supplies | Blood-giving set | [11 empty] | 0 |  | 0.05654% | 0
157 | Medical devices | Medical supplies | Safety box, disposal of used syringes & needles, 5L, P1 | [11 empty] | 0 |  | 0.05654% | 0
158 | Medical devices | Medical supplies | Safety box, disposal of used syringes & needles, 5L, P25 | [11 empty] | 0 |  | 0.05654% | 0
159 | Medical devices | Medical supplies | Cannula, IV short, 16G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
160 | Medical devices | Medical supplies | Cannula, IV short, 18G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
161 | Medical devices | Medical supplies | Cannula, IV short, 20G, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
162 | Medical devices | Medical supplies | Cannula, IV short, 20G, sterile, single use, P50 | [11 empty] | 0 |  | 0.05654% | 0
163 | Medical devices | Medical supplies | Cannula, IV short, 22G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
164 | Medical devices | Medical supplies | Cannula, IV short, 24G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
165 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
166 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P10 | [11 empty] | 0 |  | 0.05654% | 0
167 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
168 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P10 | [11 empty] | 0 |  | 0.05654% | 0
169 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
170 | Medical devices | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P10 | [11 empty] | 0 |  | 0.05654% | 0
171 | Medical devices | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
172 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
173 | Medical devices | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
174 | Medical devices | Medical supplies | Catheter, urethral, CH12, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
175 | Medical devices | Medical supplies | Catheter, urethral, CH12, sterile, P10 | [11 empty] | 0 |  | 0.05654% | 0
176 | Medical devices | Medical supplies | Catheter, ureteral, CH12, sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
177 | Medical devices | Medical supplies | Catheter, ureteral, Nelaton, CH12, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
178 | Medical devices | Medical supplies | Catheter, urethral, CH5, with metal guide wire, sterile | [11 empty] | 0 |  | 0.05654% | 0
179 | Medical devices | Medical supplies | Catheter, urethral, size CH6, with metal guide wire, sterile | [11 empty] | 0 |  | 0.05654% | 0
180 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH16, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
181 | Medical devices | Medical supplies | Catheter, urethral, Foley, CH20, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
182 | Medical devices | Medical supplies | Colostomy bag, transparent, drainable, w/filter, P10 | [11 empty] | 0 |  | 0.05654% | 0
183 | Medical devices | Medical supplies | Colostomy bag, transparent, drainable, w/filter, P20 | [11 empty] | 0 |  | 0.05654% | 0
184 | Medical devices | Medical supplies | Gauze, compress, 10 x 10cm, non-sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
185 | Medical devices | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
186 | Medical devices | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P20 | [11 empty] | 0 |  | 0.05654% | 0
187 | Medical devices | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P5 | [11 empty] | 0 |  | 0.05654% | 0
188 | Medical devices | Medical supplies | Cotton wool, 500g, roll, non-sterile | [11 empty] | 0 |  | 0.55% | 0
189 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | [11 empty] | 0 |  | 0.05654% | 0
190 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | [11 empty] | 0 |  | 0.05654% | 0
191 | Medical devices | Medical supplies | Bloodbag+CPDA-1(citrate-phosphate-dextrose-adenine),250ml,P5 | [11 empty] | 0 |  | 0.05654% | 0
192 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),350ml,P1 | [11 empty] | 0 |  | 0.05654% | 0
193 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 350ml | [11 empty] | 0 |  | 0.05654% | 0
194 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),450ml,P1 | [11 empty] | 0 |  | 0.05654% | 0
195 | Medical devices | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 450ml | [11 empty] | 0 |  | 0.05654% | 0
196 | Medical devices | Medical supplies | Bag (envelope), plastic, for drugs, 10 x 15cm | [11 empty] | 0 |  | 0.05654% | 0
197 | Medical devices | Medical supplies | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P1 | [11 empty] | 0 |  | 0.05654% | 0
198 | Medical devices | Medical supplies | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P100 | [11 empty] | 0 |  | 0.05654% | 0
199 | Medical devices | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
200 | Medical devices | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P10 | [11 empty] | 0 |  | 0.05654% | 0
201 | Medical devices | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
202 | Medical devices | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P50 | [11 empty] | 0 |  | 0.05654% | 0
203 | Medical devices | Medical supplies | Gloves, exam, latex, large, non-sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
204 | Medical devices | Medical supplies | Gloves, examination, latex, medium, non-sterile | [11 empty] | 0 |  | 0.05654% | 0
205 | Medical devices | Medical supplies | Gloves, examination, latex, medium, non-sterile, P100 | Pair | [10 empty] | 0 |  | 0.05654% | 0
206 | Medical devices | Medical supplies | Gloves, exam, latex, small, non-sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
207 | Medical devices | Medical supplies | Gloves, gynaecological, large, powder-free, sterile, 1 pair | [11 empty] | 0 |  | 0.05654% | 0
208 | Medical devices | Medical supplies | Gloves, gynaecological, large, powder-free, sterile | [11 empty] | 0 |  | 0.05654% | 0
209 | Medical devices | Medical supplies | Gloves, gynaecological, medium, powder-free, sterile, 1 pair | [11 empty] | 0 |  | 0.05654% | 0
210 | Medical devices | Medical supplies | Gloves, gynaecological, medium, powder-free, sterile | [11 empty] | 0 |  | 0.05654% | 0
211 | Medical devices | Medical supplies | Gloves, gynaecological, small, powder-free, sterile, 1 pair | [11 empty] | 0 |  | 0.05654% | 0
212 | Medical devices | Medical supplies | Gloves, gynaecological, small,powder-free, sterile, P25 | [11 empty] | 0 |  | 0.05654% | 0
213 | Medical devices | Medical supplies | Gloves, gynaecological, small,powder-free, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
214 | Medical devices | Medical supplies | Gloves, surgical, size 7, powder-free, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
215 | Medical devices | Medical supplies | Gloves, surgical, size 7, powder-free, sterile, single use, P50 | [11 empty] | 0 |  | 0.72% | 0
216 | Medical devices | Medical supplies | Gloves, surgical, size 8, powder-free, sterile, single use, P50 | [11 empty] | 0 |  | 0.7% | 0
217 | Medical devices | Medical supplies | Gloves, surgical, size 8, powder-free, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
218 | Medical devices | Medical supplies | Gloves, surgical, size 7,5, powder-free, sterile, single use | [11 empty] | 0 |  | 0.99% | 0
219 | Medical devices | Medical supplies | Gloves, surgical, size 8.5, powder-free, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
220 | Medical devices | Medical supplies | Brush, hand, scrubbing, plastic, P1 | [11 empty] | 0 |  | 0.05654% | 0
221 | Medical devices | Medical supplies | Infusion set, sterile, single use, 1 piece | [11 empty] | 0 |  | 0.05654% | 0
222 | Medical devices | Medical supplies | Infusion set, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
223 | Medical devices | Medical supplies | Infusion set, sterile, single use, P400 | [11 empty] | 0 |  | 0.05654% | 0
224 | Medical devices | Medical supplies | Needle, spinal, 25G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
225 | Medical devices | Medical supplies | Needle, luer, 19G, sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
226 | Medical devices | Medical supplies | Needle, luer, 21G, sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
227 | Medical devices | Medical supplies | Needle, luer, 23G, sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
228 | Medical devices | Medical supplies | Needle, scalp vein, 21G, sterile, P100 | [11 empty] | 0 |  | 0.05654% | 0
229 | Medical devices | Medical supplies | Needle, scalp vein, 21G, sterile, single use, P50 | [11 empty] | 0 |  | 0.05654% | 0
230 | Medical devices | Medical supplies | Needle, scalp vein, 21G, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
231 | Medical devices | Medical supplies | Needle, spinal, 22G, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
232 | Medical devices | Medical supplies | Needle, spinal, 22G, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
233 | Medical devices | Medical supplies | Needle, spinal, 22G, sterile, single use, P25 | [11 empty] | 0 |  | 0.05654% | 0
234 | Medical devices | Medical supplies | Needle, spinal, 22G, sterile, single use, P50 | [11 empty] | 0 |  | 0.05654% | 0
235 | Medical devices | Medical supplies | Surgical needle, 1/2 circle, round, spring eye, size14, sterile | [11 empty] | 0 |  | 0.05654% | 0
236 | Medical devices | Medical supplies | Needle, surgical, 1/2, 27mm, round, spring eye, P24 | [11 empty] | 0 |  | 0.05654% | 0
237 | Medical devices | Medical supplies | Needle, surgical, 1/2, 63mm, round, spring eye, P24 | [11 empty] | 0 |  | 0.05654% | 0
239 | Medical devices | Medical supplies | Adhesive bandage, wound plaster, waterproof | [11 empty] | 0 |  | 0.05654% | 0
240 | Medical devices | Medical supplies | Pressure cuff, infusion/transfusion, for bags 500-1000 ml | [11 empty] | 0 |  | 0.05654% | 0
241 | Medical devices | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P10 | [11 empty] | 0 |  | 0.05654% | 0
242 | Medical devices | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
243 | Medical devices | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P15 | [11 empty] | 0 |  | 0.05654% | 0
244 | Medical devices | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P5 | [11 empty] | 0 |  | 0.05654% | 0
245 | Medical devices | Medical supplies | Tube, suction, Yankauer, 28cm | [11 empty] | 0 |  | 0.05654% | 0
246 | Medical devices | Medical supplies | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile | [11 empty] | 0 |  | 0.05654% | 0
247 | Medical devices | Medical supplies | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
248 | Medical devices | Medical supplies | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
249 | Medical devices | Medical supplies | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
250 | Medical devices | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
251 | Medical devices | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
252 | Medical devices | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
253 | Medical devices | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
254 | Medical devices | Medical supplies | Suture, absorbable, DEC 3 (2-0), spool, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
255 | Medical devices | Medical supplies | Suture, absorbable, DEC 3 (2-0), spool, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
256 | Medical devices | Medical supplies | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
257 | Medical devices | Medical supplies | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
258 | Medical devices | Medical supplies | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P12 | [11 empty] | 0 |  | 0.05654% | 0
259 | Medical devices | Medical supplies | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
260 | Medical devices | Medical supplies | Suture, absorbable, DEC 4 (1), 3/8, 50mm, round, sterile, P36 | [11 empty] | 0 |  | 0.05654% | 0
261 | Medical devices | Medical supplies | Suture, absorbable, DEC 1.5 (4/0), with needle, fistula repair | [11 empty] | 0 |  | 0.05654% | 0
262 | Medical devices | Medical supplies | Suture, absorbable, DEC 2 (3/0), with needle, fistula repair | [11 empty] | 0 |  | 0.05654% | 0
263 | Medical devices | Medical supplies | Suture, absorbable, DEC 3 (2/0), with needle, fistula repair | [11 empty] | 0 |  | 0.05654% | 0
264 | Medical devices | Medical supplies | Suture, absorbable, DEC 3.5 (0), with needle, fistula repair | [11 empty] | 0 |  | 0.05654% | 0
265 | Medical devices | Medical supplies | Suture, absorbable, DEC 4 (1), with needle, fistula repair | [11 empty] | 0 |  | 0.05654% | 0
266 | Medical devices | Medical supplies | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | [11 empty] | 0 |  | 0.05654% | 0
267 | Medical devices | Medical supplies | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | [11 empty] | 0 |  | 0.05654% | 0
268 | Medical devices | Medical supplies | Suture, non-absorbable PA, DEC 3 (2/0), 1/2 circle, round | [11 empty] | 0 |  | 0.05654% | 0
269 | Medical devices | Medical supplies | Suture, non-absorbable PA, DEC 4 (1), 1/2 circle, round | [11 empty] | 0 |  | 0.05654% | 0
270 | Medical devices | Medical supplies | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | [11 empty] | 0 |  | 0.05654% | 0
271 | Medical devices | Medical supplies | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | [11 empty] | 0 |  | 0.05654% | 0
272 | Medical devices | Medical supplies | Swab, alcohol, sterile, single use | [11 empty] | 0 |  | 0.05654% | 0
273 | Medical devices | Medical supplies | Syringe, 0.5ml, w/permanently att. needle,sterile,single use | [11 empty] | 0 |  | 0.05654% | 0
274 | Medical devices | Medical supplies | Auto Disabled Syringe, 0.5ml, w/permanently att. needle,sterile,single use | [11 empty] | 0 |  | 0.05654% | 0
275 | Medical devices | Medical supplies | Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
276 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
277 | Medical devices | Medical supplies | Syringe, luer, 2 ml, w/21G needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
278 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 2ml, w/21G needle, sterile, single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
279 | Medical devices | Medical supplies | Syringe, feeding, catheter tip, 50 ml, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
280 | Medical devices | Medical supplies | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
281 | Medical devices | Medical supplies | Syringe, feeding, catheter tip, 50 ml, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
282 | Medical devices | Medical supplies | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P50 | [11 empty] | 0 |  | 0.05654% | 0
283 | Medical devices | Medical supplies | Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.49% | 0
284 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
285 | Medical devices | Medical supplies | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
286 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
287 | Medical devices | Medical supplies | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
288 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
289 | Medical devices | Medical supplies | Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | [11 empty] | 0 |  | 0.05654% | 0
290 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | [11 empty] | 0 |  | 0.05654% | 0
291 | Medical devices | Medical supplies | Syringe, luer lock, 20 ml, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
292 | Medical devices | Medical supplies | Auto Disabled Syringe, luer lock, 20 ml, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
293 | Medical devices | Medical supplies | Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
294 | Medical devices | Medical supplies | Auto Disabled Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
295 | Medical devices | Medical supplies | Bladder syringe 60-100 mls (with long nozzle, not luer lock) | [11 empty] | 0 |  | 0.05654% | 0
296 | Medical devices | Medical supplies | Auto Disabled Bladder syringe 60-100 mls (with long nozzle, not luer lock) | [11 empty] | 0 |  | 0.05654% | 0
297 | Medical devices | Medical supplies | Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
298 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
299 | Medical devices | Medical supplies | Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
300 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
301 | Medical devices | Medical supplies | Syringe, luer, 2ml, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
302 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 2ml, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
303 | Medical devices | Medical supplies | Syringe, luer, 5ml, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
304 | Medical devices | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/o needle, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
305 | Medical devices | Medical supplies | Tape, adhesive, zinc oxide, 2.5cm x 5m, P1 | [11 empty] | 0 |  | 0.05654% | 0
306 | Medical devices | Medical supplies | Tape, adhesive, zinc oxide, 2.5cm x 5m, P10 | [11 empty] | 0 |  | 0.05654% | 0
307 | Medical devices | Medical supplies | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P1 | [11 empty] | 0 |  | 0.05654% | 0
308 | Medical devices | Medical supplies | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P10 | [11 empty] | 0 |  | 0.05654% | 0
309 | Medical devices | Medical supplies | Tape, umbilical, 3mmx50m, non-sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
310 | Medical devices | Medical supplies | Tape, umbilical, 3mm x 15cm, non-sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
311 | Medical devices | Medical supplies | Tourniquet, latex rubber, 75cm | [11 empty] | 0 |  | 0.05654% | 0
312 | Medical devices | Medical supplies | Tube, suction, CH10, 50 cm long, conical tip, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
313 | Medical devices | Medical supplies | Tube, suction, CH14, 50 cm long, conical tip, sterile, P1 | [11 empty] | 0 |  | 0.05654% | 0
314 | Medical devices | Medical supplies | Bag, urine, collecting, 2 litres, P10 | [11 empty] | 0 |  | 0.05654% | 0
315 | Medical devices | Medical supplies | Bag, urine, collecting, 2 litres, tap for emptying | [11 empty] | 0 |  | 0.05654% | 0
316 | Medical devices | Medical supplies | Sound, uterine, Martin, 32cm | [11 empty] | 0 |  | 0.05654% | 0
317 | Medical devices | Medical utensils | Basin, kidney, stainless steel, 1400ml | [11 empty] | 0 |  | 0.05654% | 0
318 | Medical devices | Medical utensils | Basin, kidney, polypropylene, 475ml | [11 empty] | 0 |  | 0.05654% | 0
319 | Medical devices | Medical utensils | Basin, kidney, stainless steel, 825ml | [11 empty] | 0 |  | 0.05654% | 0
320 | Medical devices | Medical utensils | Bowl, stainless steel, 180ml | [11 empty] | 0 |  | 0.05654% | 0
321 | Medical devices | Medical utensils | Bowl, stainless steel, 4L | [11 empty] | 0 |  | 0.05654% | 0
322 | Medical devices | Medical utensils | Bowl, stainless steel, 500 ml | [11 empty] | 0 |  | 0.05654% | 0
323 | Medical devices | Medical utensils | Bowl, polypropylene, 6L | [11 empty] | 0 |  | 0.05654% | 0
324 | Medical devices | Medical utensils | Jar, forceps, 180 mm, no cover | [11 empty] | 0 |  | 0.05654% | 0
325 | Medical devices | Medical utensils | Jar, thermometer, plastic, 11cm, w/o cover | [11 empty] | 0 |  | 0.05654% | 0
326 | Medical devices | Medical utensils | Bowl, sponge, stainless steel, 1.2 L | [11 empty] | 0 |  | 0.05654% | 0
327 | Medical devices | Medical utensils | Ruler, metal, in centimeters, 30cm | [11 empty] | 0 |  | 0.05654% | 0
328 | Medical devices | Medical utensils | Tray, dressing, stainless steel, 30 x 20 x 3cm | [11 empty] | 0 |  | 0.05654% | 0
329 | Medical devices | Medical utensils | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm, with cover | [11 empty] | 0 |  | 0.05654% | 0
330 | Medical devices | Medical utensils | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover | [11 empty] | 0 |  | 0.84% | 0
331 | Medical devices | Medical Attire & Linen | Apron, protection, plastic, reusable | [11 empty] | 0 |  | 0.05654% | 0
332 | Medical devices | Medical Attire & Linen | Apron, surgical, rubber, washable | [11 empty] | 0 |  | 0.05654% | 0
333 | Medical devices | Medical Attire & Linen | Sheet, bed, 289 x 178 cm, washable | [11 empty] | 0 |  | 0.05654% | 0
334 | Medical devices | Medical Attire & Linen | Blanket, survival, 220 x 140cm | [11 empty] | 0 |  | 0.05654% | 0
335 | Medical devices | Medical Attire & Linen | Cotton cloth, “tetra”, 100 x 100cm | [11 empty] | 0 |  | 0.05654% | 0
336 | Medical devices | Medical Attire & Linen | Drape, surgical, woven, 100x100 cm, reusable | [11 empty] | 0 |  | 0.05654% | 0
337 | Medical devices | Medical Attire & Linen | Drape, surgical woven, 100 x 150cm | [11 empty] | 0 |  | 0.05654% | 0
338 | Medical devices | Medical Attire & Linen | Drape, surgical, woven, 90x150 cm, reusable | [11 empty] | 0 |  | 0.05654% | 0
339 | Medical devices | Medical Attire & Linen | Draw sheet, waterproof, 80 x 120cm, single use | [11 empty] | 0 |  | 0.05654% | 0
340 | Medical devices | Medical Attire & Linen | Draw sheet, plastic, 90 x 180cm, reusable | [11 empty] | 0 |  | 0.05654% | 0
341 | Medical devices | Medical Attire & Linen | Draw sheet, plastic, 100 x 100 cm | [11 empty] | 0 |  | 0.05654% | 0
342 | Medical devices | Medical Attire & Linen | Draw sheet, rubber, 90x150 cm, reusable | [11 empty] | 0 |  | 0.05654% | 0
343 | Medical devices | Medical Attire & Linen | Glasses, safety, regular size, disposable | [11 empty] | 0 |  | 0.05654% | 0
344 | Diagnostic Test Kits | In Vitro Diagnostics | Tile, porcelain, with depressions for blood grouping | [11 empty] | 0 |  | 0.05654% | 0
345 | Medical devices | Med. Diagnostic Equip&Supplies | Dopplerfetalheart | [11 empty] | 0 |  | 0.05654% | 0
346 | Medical devices | Med. Diagnostic Equip&Supplies | Haemoglobin photometer handheld w/accessories | [11 empty] | 0 |  | 1.39% | 0
347 | Medical devices | Med. Diagnostic Equip&Supplies | Height measuring instrument (0-2.2m) | [11 empty] | 0 |  | 0.05654% | 0
348 | Medical devices | Med. Diagnostic Equip&Supplies | Capillary tubes for HemoCue Hb 301, P200 | [11 empty] | 0 |  | 0.05654% | 0
349 | Medical devices | Med. Diagnostic Equip&Supplies | Pelvimeter, Collyer type, external | [11 empty] | 0 |  | 0.05654% | 0
350 | Medical devices | Med. Diagnostic Equip&Supplies | Pregnancy wheel | [11 empty] | 0 |  | 0.05654% | 0
351 | Medical devices | Med. Diagnostic Equip&Supplies | Scale, infant, beam type, 16kg x 5g | [11 empty] | 0 |  | 0.8% | 0
352 | Medical devices | Med. Diagnostic Equip&Supplies | Sling, for use with infant scale | [11 empty] | 0 |  | 0.05654% | 0
353 | Medical devices | Med. Diagnostic Equip&Supplies | Scale, mother/child, electronic, 200kg | [11 empty] | 0 |  | 0.05654% | 0
354 | Medical devices | Med. Diagnostic Equip&Supplies | Scale, physician, adult beam balance, 0-200kg | [11 empty] | 0 |  | 1.37% | 0
355 | Medical devices | Med. Diagnostic Equip&Supplies | Scale, infant, spring-type, 5kg x 25g | [11 empty] | 0 |  | 0.05654% | 0
356 | Medical devices | Med. Diagnostic Equip&Supplies | Sphygmomanometer, aneroid, adult | [11 empty] | 0 |  | 0.05654% | 0
357 | Medical devices | Med. Diagnostic Equip&Supplies | Stethoscope, binaural, complete | [11 empty] | 0 |  | 0.05654% | 0
358 | Medical devices | Med. Diagnostic Equip&Supplies | Stethoscope, obstetrical, Pinard, monoaural | [11 empty] | 0 |  | 0.05654% | 0
359 | Medical devices | Med. Diagnostic Equip&Supplies | Thermometer, clinical, digital 32°C-43°C | [11 empty] | 0 |  | 0.05654% | 0
360 | Medical devices | Med. Diagnostic Equip&Supplies | Timer, for respiration, ARI | [11 empty] | 0 |  | 0.05654% | 0
361 | Medical devices | Med. Diagnostic Equip&Supplies | Ultrasound, scanner, mobile w/ printer and accessories | [11 empty] | 0 |  | 4.28% | 0
362 | Medical devices | Medical & Surgical Instruments | Spatula, Ribbon retractor, malleable, 27 mm x 250 mm | [11 empty] | 0 |  | 0.05654% | 0
363 | Medical devices | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no.10 | [11 empty] | 0 |  | 0.05654% | 0
364 | Medical devices | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 11, P1 | [11 empty] | 0 |  | 0.05654% | 0
365 | Medical devices | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 11, P100 | [11 empty] | 0 |  | 0.05654% | 0
366 | Medical devices | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 15, P1 | [11 empty] | 0 |  | 0.05654% | 0
367 | Medical devices | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 22, P100 | [11 empty] | 0 |  | 0.05654% | 0
368 | Medical devices | Medical & Surgical Instruments | Forceps, intestinal clamp, Doyen, 23cm, curved | [11 empty] | 0 |  | 0.05654% | 0
369 | Medical devices | Medical & Surgical Instruments | Forceps, intestinal, clamp, Doyen, 23cm, straight | [11 empty] | 0 |  | 0.05654% | 0
370 | Medical devices | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P100 | [11 empty] | 0 |  | 0.86% | 0
371 | Medical devices | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P1 | [11 empty] | 0 |  | 0.05654% | 0
372 | Medical devices | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P50 | [11 empty] | 0 |  | 0.05654% | 0
373 | Medical devices | Medical & Surgical Instruments | Cranioclast, Braun, 420 mm | [11 empty] | 0 |  | 0.05654% | 0
374 | Medical devices | Medical & Surgical Instruments | Curette, uterine, Sims, 12 mm, sharp | [11 empty] | 0 |  | 0.05654% | 0
375 | Medical devices | Medical & Surgical Instruments | Curette, uterine, Sims, 7 mm, sharp | [11 empty] | 0 |  | 0.05654% | 0
376 | Medical devices | Medical & Surgical Instruments | Curette, uterine, Sims, 8 mm, blunt | [11 empty] | 0 |  | 0.05654% | 0
377 | Medical devices | Medical & Surgical Instruments | Curette, uterine, Sims, 9 mm, sharp | [11 empty] | 0 |  | 0.05654% | 0
378 | Medical devices | Medical & Surgical Instruments | Hook, decapitation, Braun, 31cm | [11 empty] | 0 |  | 0.05654% | 0
379 | Medical devices | Medical & Surgical Instruments | Dilator, uterine, Hegar, set of 16 dilators, size 3-18 | [11 empty] | 0 |  | 0.05654% | 0
380 | Medical devices | Medical & Surgical Instruments | Forceps, dressing, Cheron, 25cm | [11 empty] | 0 |  | 0.05654% | 0
381 | Medical devices | Medical & Surgical Instruments | Forceps, dressing, standard, 145 mm, straight | [11 empty] | 0 |  | 0.05654% | 0
382 | Medical devices | Medical & Surgical Instruments | Forceps, uterine, Duplay, 28cm, curved | [11 empty] | 0 |  | 0.05654% | 0
383 | Medical devices | Medical & Surgical Instruments | Forceps, catheter introducing, Magill, child | [11 empty] | 0 |  | 0.05654% | 0
384 | Medical devices | Medical & Surgical Instruments | Forceps, obstetric, Simpson-Braun, 35 cm | [11 empty] | 0 |  | 0.05654% | 0
385 | Medical devices | Medical & Surgical Instruments | Forceps, uterine, Phaneuf, 21.5 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
386 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 20 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
387 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 22cm, straight | [11 empty] | 0 |  | 0.05654% | 0
388 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 24 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
389 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 16 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
390 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 16 cm, straight | [11 empty] | 0 |  | 0.05654% | 0
391 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Bengolea, 20cm, curved | [11 empty] | 0 |  | 0.05654% | 0
392 | Medical devices | Medical & Surgical Instruments | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 16mm, S | [11 empty] | 0 |  | 0.05654% | 0
393 | Medical devices | Medical & Surgical Instruments | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 19mm, L | [11 empty] | 0 |  | 0.05654% | 0
394 | Medical devices | Medical & Surgical Instruments | Forceps, sterilizer, Cheatle, 27 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
395 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Crafoord, 24cm, curved | [11 empty] | 0 |  | 0.05654% | 0
396 | Medical devices | Medical & Surgical Instruments | Forceps, dressing, standard, fine serrated jaws, 20cm | [11 empty] | 0 |  | 0.05654% | 0
397 | Medical devices | Medical & Surgical Instruments | Forceps, dressing, standard, 250 mm, straight | [11 empty] | 0 |  | 0.05654% | 0
398 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Kelly, 14cm, curved | [11 empty] | 0 |  | 0.05654% | 0
399 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Kocher, 140 mm, 1x2 teeth, curved | [11 empty] | 0 |  | 0.05654% | 0
400 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Kocher, 14cm, straight | [11 empty] | 0 |  | 0.05654% | 0
401 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Mixter, 14cm | [11 empty] | 0 |  | 0.05654% | 0
402 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Mixter, 23cm | [11 empty] | 0 |  | 0.05654% | 0
403 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Halsted-Mosquito, 12.5cm, curved | [11 empty] | 0 |  | 0.05654% | 0
404 | Medical devices | Medical & Surgical Instruments | Forceps, uterine, Museux, 24cm, curved | [11 empty] | 0 |  | 0.05654% | 0
405 | Medical devices | Medical & Surgical Instruments | Forceps, artery, Spencer-Wells, 20.5 cm, curved | [11 empty] | 0 |  | 0.05654% | 0
406 | Medical devices | Medical & Surgical Instruments | Forceps, peritoneal, Faure, 21cm, slightly curved | [11 empty] | 0 |  | 0.05654% | 0
407 | Medical devices | Medical & Surgical Instruments | Forceps, sponge-holding, Foerster, 20 cm, straight | [11 empty] | 0 |  | 0.69% | 0
408 | Medical devices | Medical & Surgical Instruments | Forceps, sponge-holding, Foerster, 24 cm, straight | [11 empty] | 0 |  | 0.05654% | 0
409 | Medical devices | Medical & Surgical Instruments | Forceps, tenaculum, vulsellum, Schröder, straight, 25cm | [11 empty] | 0 |  | 0.05654% | 0
410 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Allis, 3x4 teeth, 20 cm | [11 empty] | 0 |  | 0.05654% | 0
411 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, standard, 145mm, straight | [11 empty] | 0 |  | 0.05654% | 0
412 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Allis, 4x5 teeth, 20 cm | [11 empty] | 0 |  | 0.05654% | 0
413 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Babcock, 20 cm | [11 empty] | 0 |  | 0.05654% | 0
414 | Medical devices | Medical & Surgical Instruments | Forceps,tissue,standard,dissecting,1x2 teeth,20cm, straight | [11 empty] | 0 |  | 0.05654% | 0
415 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, standard, 250 mm, straight | [11 empty] | 0 |  | 0.05654% | 0
416 | Medical devices | Medical & Surgical Instruments | Forceps, Schroeder, Vulsellum, 25 cm, 2x2 prongs, curved | [11 empty] | 0 |  | 0.05654% | 0
417 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Allis, 4x5 teeth, 15 cm | [11 empty] | 0 |  | 0.05654% | 0
418 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Allis, 3x4 teeth, 15 cm | [11 empty] | 0 |  | 0.05654% | 0
419 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Colin, 16 cm | [11 empty] | 0 |  | 0.05654% | 0
420 | Medical devices | Medical & Surgical Instruments | Forceps, tissue, Duval, 23cm | [11 empty] | 0 |  | 0.05654% | 0
421 | Medical devices | Medical & Surgical Instruments | Hook, fistula, Emmet, 22 cm | [11 empty] | 0 |  | 0.05654% | 0
422 | Medical devices | Medical & Surgical Instruments | Accesory kit for MVA, including (1) O-ring, (1) collar stop, (1), cap, (1) 2cc silicone | [11 empty] | 0 |  | 0.05654% | 0
423 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 10mm, integrated base, SR10 | [11 empty] | 0 |  | 0.05654% | 0
424 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 12mm, integrated base, SR12 | [11 empty] | 0 |  | 0.05654% | 0
425 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 4mm, integrated base, SR4 | [11 empty] | 0 |  | 0.05654% | 0
426 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 5mm, integrated base, SR5 | [11 empty] | 0 |  | 0.05654% | 0
427 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 6mm, integrated base, SR6 | [11 empty] | 0 |  | 0.05654% | 0
428 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 7mm, integrated base, SR7 | [11 empty] | 0 |  | 0.05654% | 0
429 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 8mm, integrated base, SR8 | [11 empty] | 0 |  | 0.05654% | 0
430 | Medical devices | Medical & Surgical Instruments | Cannulae, Easygrip, 9mm, integrated base, SR9 | [11 empty] | 0 |  | 0.05654% | 0
431 | Medical devices | Medical & Surgical Instruments | Dilators, Denniston, polymer, P5 | [11 empty] | 0 |  | 3.13% | 0
432 | Medical devices | Medical & Surgical Instruments | MVA Ipas kitting service | [11 empty] | 0 |  | 0.05654% | 0
433 | Medical devices | Medical & Surgical Instruments | MVA with 2cc silicone | [11 empty] | 0 |  | 2.88% | 0
434 | Medical devices | Medical & Surgical Instruments | Silicone, 2cc, P10 | [11 empty] | 0 |  | 0.73% | 0
435 | Medical devices | Medical & Surgical Instruments | Needle, aneurysm, Deschamps, 20cm, sharp, left hand | [11 empty] | 0 |  | 0.05654% | 0
436 | Medical devices | Medical & Surgical Instruments | Needle, aneurysm, Deschamps, 20cm, sharp, right hand | [11 empty] | 0 |  | 0.05654% | 0
437 | Medical devices | Medical & Surgical Instruments | Needle holder, Derf, 11.5 cm, straight | [11 empty] | 0 |  | 0.05654% | 0
438 | Medical devices | Medical & Surgical Instruments | Needle holder, Baby-Crile-Wood, 15cm, delicate | [11 empty] | 0 |  | 0.05654% | 0
439 | Medical devices | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 18cm, curved | [11 empty] | 0 |  | 0.05654% | 0
440 | Medical devices | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 18cm, straight | [11 empty] | 0 |  | 0.05654% | 0
441 | Medical devices | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 20cm, curved | [11 empty] | 0 |  | 0.05654% | 0
442 | Medical devices | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 20cm, straight | [11 empty] | 0 |  | 0.05654% | 0
443 | Medical devices | Medical & Surgical Instruments | Needle, luer, 25G, sterile, single use, P100 | [11 empty] | 0 |  | 0.05654% | 0
444 | Medical devices | Medical & Surgical Instruments | Needle, scalp vein, butterfly, 25G, sterile, single use,P100 | [11 empty] | 0 |  | 0.05654% | 0
445 | Medical devices | Medical & Surgical Instruments | Perforator, Smellie, 25cm | [11 empty] | 0 |  | 0.05654% | 0
446 | Medical devices | Medical & Surgical Instruments | Probe, surgical, double-ended, 14.5 cm | [11 empty] | 0 |  | 0.05654% | 0
447 | Medical devices | Medical & Surgical Instruments | Probe, surgical, with eye, malleable, 20 cm | [11 empty] | 0 |  | 0.05654% | 0
448 | Medical devices | Medical & Surgical Instruments | Retractor, double-ended, Farabeuf, 15 cm, pair | [11 empty] | 0 |  | 0.05654% | 0
449 | Medical devices | Medical & Surgical Instruments | Retractor, Langenbeck, single-ended, medium, 21cm | [11 empty] | 0 |  | 0.05654% | 0
450 | Medical devices | Medical & Surgical Instruments | Retractor, abdominal, Balfour, 3 blades | [11 empty] | 0 |  | 0.05654% | 0
451 | Medical devices | Medical & Surgical Instruments | Retractor, abdominal, Collin, 3 blades | [11 empty] | 0 |  | 0.05654% | 0
452 | Medical devices | Medical & Surgical Instruments | Retractor, vaginal, Doyen, 8.5 x 4.5cm | [11 empty] | 0 |  | 0.05654% | 0
453 | Medical devices | Medical & Surgical Instruments | Retractor, abdominal, Gosset, 3 blades | [11 empty] | 0 |  | 0.05654% | 0
454 | Medical devices | Medical & Surgical Instruments | Safety Pin, forceps holder, Mayo, 11.5cm | [11 empty] | 0 |  | 0.05654% | 0
455 | Medical devices | Medical & Surgical Instruments | Safety Pin, forceps holder, Schaedel, 9cm | [11 empty] | 0 |  | 0.05654% | 0
456 | Medical devices | Medical & Surgical Instruments | Scissors for umbilical cord cutting, reusable, sterile | [11 empty] | 0 |  | 0.05654% | 0
457 | Medical devices | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 15 cm | [11 empty] | 0 |  | 0.05654% | 0
458 | Medical devices | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 16 cm | [11 empty] | 0 |  | 0.05654% | 0
459 | Medical devices | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 18 cm | [11 empty] | 0 |  | 0.05654% | 0
460 | Medical devices | Medical & Surgical Instruments | Scalpel handle no. 3, blade holder, 13 cm | [11 empty] | 0 |  | 0.05654% | 0
461 | Medical devices | Medical & Surgical Instruments | Scalpel handle no. 4, blade holder, 13 cm | [11 empty] | 0 |  | 0.05654% | 0
462 | Medical devices | Medical & Surgical Instruments | Scalpel handle no. 4, blade holder, long 21cm | [11 empty] | 0 |  | 0.05654% | 0
463 | Medical devices | Medical & Surgical Instruments | Scalpel, knife, blade no. 22, with plastic handle, sterile | [11 empty] | 0 |  | 0.05654% | 0
464 | Medical devices | Medical & Surgical Instruments | Scalpel, knife, blade no. 22, w/ plastic handle, sterile,P10 | [11 empty] | 0 |  | 0.05654% | 0
465 | Medical devices | Medical & Surgical Instruments | Scissors, Deaver, 14cm, curved, sharp/blunt | [11 empty] | 0 |  | 0.05654% | 0
466 | Medical devices | Medical & Surgical Instruments | Scissors, fistula, 20cm curved | [11 empty] | 0 |  | 0.05654% | 0
467 | Medical devices | Medical & Surgical Instruments | Scissors, Episiotomy, Braun, 14.5, angular | [11 empty] | 0 |  | 0.05654% | 0
468 | Medical devices | Medical & Surgical Instruments | Scissors, Bandage, Lister, 14 cm, angular, sharp/blunt | [11 empty] | 0 |  | 0.05654% | 0
469 | Medical devices | Medical & Surgical Instruments | Scissors, Bandage, Lister, 18 cm, angular, sharp / blunt | [11 empty] | 0 |  | 0.05654% | 0
470 | Medical devices | Medical & Surgical Instruments | Scissors, Boyd, 17cm, double curved | [11 empty] | 0 |  | 0.05654% | 0
471 | Medical devices | Medical & Surgical Instruments | Scissors, Deaver, 14cm, straight, sharp/blunt | [11 empty] | 0 |  | 0.05654% | 0
472 | Medical devices | Medical & Surgical Instruments | Scissors, gynaecological, 20cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
473 | Medical devices | Medical & Surgical Instruments | Scissors, Mayo, 14cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
474 | Medical devices | Medical & Surgical Instruments | Scissors, Mayo, 17cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
475 | Medical devices | Medical & Surgical Instruments | Scissors, Mayo, 17 cm, straight, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
476 | Medical devices | Medical & Surgical Instruments | Scissors, Mayo, 23cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
477 | Medical devices | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 14cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
478 | Medical devices | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 18cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
479 | Medical devices | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 23cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
480 | Medical devices | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 24 cm, curved, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
481 | Medical devices | Medical & Surgical Instruments | Scissors, stitch, 15cm | [11 empty] | 0 |  | 0.05654% | 0
482 | Medical devices | Medical & Surgical Instruments | Scissors, suture stich, 18cm, straight, sharp/blunt | [11 empty] | 0 |  | 0.05654% | 0
483 | Medical devices | Medical & Surgical Instruments | Scissors, Thorek, 20cm, curved | [11 empty] | 0 |  | 0.05654% | 0
484 | Medical devices | Medical & Surgical Instruments | Scissors, Thorek, 18cm, curved, sharp/blunt | [11 empty] | 0 |  | 0.05654% | 0
485 | Medical devices | Medical & Surgical Instruments | Scissors, umbilical, USA-model, 10 cm | [11 empty] | 0 |  | 0.05654% | 0
486 | Medical devices | Medical & Surgical Instruments | Scissors, uterine, Sims, curved, 20 cm, blunt/blunt | [11 empty] | 0 |  | 0.05654% | 0
487 | Medical devices | Medical & Surgical Instruments | Scissors, Thorek, 18cm, curved 90 degrees | [11 empty] | 0 |  | 0.05654% | 0
488 | Medical devices | Medical & Surgical Instruments | Scoop, uterine, Simon, 6mm, sharp | [11 empty] | 0 |  | 0.05654% | 0
489 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Auvard, retractor, weighted, 8 x 3.8cm | [11 empty] | 0 |  | 0.05654% | 0
490 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Auvard, retractor, weighted, 10 x 4 cm | [11 empty] | 0 |  | 0.05654% | 0
491 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Sims, retractor, double-ended, large | [11 empty] | 0 |  | 0.05654% | 0
492 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Sims, retractor, double-ended, medium | [11 empty] | 0 |  | 0.05654% | 0
493 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 11.5 x 3.5 cm | [11 empty] | 0 |  | 0.05654% | 0
494 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 9.5 x 3.5 cm | [11 empty] | 0 |  | 0.92% | 0
495 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 7.5 x 2 cm | [11 empty] | 0 |  | 0.05654% | 0
496 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Graves, 115 x 35mm | [11 empty] | 0 |  | 0.05654% | 0
497 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Graves, 95 x 35mm | [11 empty] | 0 |  | 0.05654% | 0
498 | Medical devices | Medical & Surgical Instruments | Speculum, vaginal, Graves, 75 x 20mm | [11 empty] | 0 |  | 0.05654% | 0
499 | Medical devices | Medical & Surgical Instruments | Clamp, towel, Backhaus, 120 mm | [11 empty] | 0 |  | 0.05654% | 0
500 | Medical devices | Medical & Surgical Instruments | Clamp, towel, Backhaus, 90 mm | [11 empty] | 0 |  | 0.05654% | 0
501 | Medical devices | Medical & Surgical Instruments | Sound, uterine, Martin, malleable, 30 cm | [11 empty] | 0 |  | 0.05654% | 0
502 | Medical devices | Medical & Surgical Instruments | Vacuum extractor, Bird, anterior and posterior cups, manual, set | [11 empty] | 0 |  | 0.53% | 0
503 | Medical devices | Medical & Surgical Instruments | Vacuum extractor, Bird, manual, complete set | [11 empty] | 0 |  | 3.35% | 0
504 | Medical devices | Medical & Surgical Instruments | Vacuum extractor, Vacca OmniCup, hand-held, 1 person, reusable | [11 empty] | 0 |  | 0.91% | 0
505 | Medical devices | Medical & Surgical Instruments | Forceps, Dressing, Standard Pattern, 12.7 cm | [11 empty] | 0 |  | 0.05654% | 0
506 | Medical devices | Medical & Surgical Instruments | Forceps, Intestinal, Allis, Delicate, 15 cm (5x6 teeth) | [11 empty] | 0 |  | 0.05654% | 0
507 | Medical devices | Medical & Surgical Instruments | Retractor, double-ended, Farabeuf, 120 mm, pair | [11 empty] | 0 |  | 0.05654% | 0
508 | Medical devices | Medical & Surgical Instruments | Army-Navy retractors, double-ended (for subumbilical procedure) | [11 empty] | 0 |  | 0.05654% | 0
509 | Medical devices | Medical & Surgical Instruments | Forceps, Artery, Kelly, Straight, 14 cm | [11 empty] | 0 |  | 0.05654% | 0
510 | Medical devices | Medical & Surgical Instruments | Forceps mosquito, delicate, curved, 14 cm | [11 empty] | 0 |  | 0.05654% | 0
511 | Medical devices | Medical & Surgical Instruments | Scissors, Tonsil, Metzenbaum, Curved, 18 cm | [11 empty] | 0 |  | 0.05654% | 0
512 | Medical devices | Medical & Surgical Instruments | Ramathobodi tubal hook | [11 empty] | 0 |  | 0.05654% | 0
513 | Medical devices | Medical & Surgical Instruments | Ramathobodi uterine elevator, 28 cm | [11 empty] | 0 |  | 0.05654% | 0
514 | Medical devices | Medical & Surgical Instruments | Forceps, uterine, tenaculum, Duplay, 248 mm, curved | [11 empty] | 0 |  | 0.05654% | 0
515 | Medical devices | Medical & Surgical Instruments | Forceps, sponge, Foerster, straight, 24 cm | [11 empty] | 0 |  | 0.05654% | 0
516 | Medical devices | Medical & Surgical Instruments | 1 Jackson vaginal retractor, 38mm X 76mm (deep blade) (for alternative method for inserting the uterine elevator) | [11 empty] | 0 |  | 0.05654% | 0
517 | Medical devices | Medical Electrical Equipment | ECG Recorder,portable,w/access | [11 empty] | 0 |  | 0.05654% | 0
518 | Medical devices | Med. Diagnostics Equipment & Supplies | Sphygmomanaometer, Aneroid, Adult | [11 empty] | 0 |  | 0.05654% | 0
519 | Medical devices | Anaesthesia & Resus. Equip | Anaesthesia machine Portable | [11 empty] | 0 |  | 0.05654% | 0
520 | Medical devices | Medical Electrical Equipment | Electrocardiograph with 12 channels | [11 empty] | 0 |  | 0.05654% | 0
521 | Medical devices | Medical Electrical Equipment | CTG (Cardiotocography) instrument | [11 empty] | 0 |  | 0.05654% | 0
522 | Medical devices | Medical Electrical Equipment | Automated external defibrillator system | [11 empty] | 0 |  | 0.05654% | 0
523 | Medical devices | Med. Diagnostics & Supplies | Handheld breast ultrasound | [11 empty] | 0 |  | 0.05654% | 0
524 | Medical devices | Medical Electrical Equipment | Ultrasound, Scanner, Digital w/access | [11 empty] | 0 |  | 0.05654% | 0
525 | Medical devices | Med. Diagnostics & Supplies | Digital blood pressure monitor | [11 empty] | 0 |  | 0.05654% | 0
526 | Medical devices | Medical Sterilization Equipmnt | Autoclave | [11 empty] | 0 |  | 0.05654% | 0
527 | Medical devices | Medical Electrical Diagnostic Equipment | Colposcope | [11 empty] | 0 |  | 0.06% | 0
528 | Diagnostic Test Kits | In Vitro Diagnostics | HIV 1+2 rapid test (Colloidal Test) | [11 empty] | 0 |  | 4.93% | 0
529 | Diagnostic Test Kits | In Vitro Diagnostics | HIV 1/2 Rapid Antibody Test | [11 empty] | 0 |  | 1.64% | 0
530 | Diagnostic Test Kits | In Vitro Diagnostics | Anti-HIV 1+2 ELISA | [11 empty] | 0 |  | 0.05654% | 0
531 | Diagnostic Test Kits | In Vitro Diagnostics | Anti - HCV Rapid Test | [11 empty] | 0 |  | 0.05654% | 0
532 | Diagnostic Test Kits | In Vitro Diagnostics | HCV Antibody Test | [11 empty] | 0 |  | 0.05654% | 0
533 | Diagnostic Test Kits | In Vitro Diagnostics | Test Hepatitis B surface antigen (HBsAg) | [11 empty] | 0 |  | 0.05654% | 0
534 | Diagnostic Test Kits | In Vitro Diagnostics | Test Hepatitis C virus (HCV) | [11 empty] | 0 |  | 0.05654% | 0
535 | Diagnostic Test Kits | In Vitro Diagnostics | HPV Test (Professional Use Only) | [11 empty] | 0 |  | 0.05654% | 0
536 | Diagnostic Test Kits | In Vitro Diagnostics | HPV Test | [11 empty] | 0 |  | 0.05654% | 0
537 | Diagnostic Test Kits | In Vitro Diagnostics | HIV Blot Elisa Test | [11 empty] | 0 |  | 0.05654% | 0
538 | Diagnostic Test Kits | In Vitro Diagnostics | Anti-HIV 1+2 ELISA Blot | [11 empty] | 0 |  | 0.05654% | 0
539 | Diagnostic Test Kits | In Vitro Diagnostics | HIV Ag/Ab Combination | [11 empty] | 0 |  | 0.05654% | 0
540 | Diagnostic Test Kits | In Vitro Diagnostics | HIV (1+2) antibody (colloidal gold) | [11 empty] | 0 |  | 0.05654% | 0
541 | Diagnostic Test Kits | In Vitro Diagnostics | HIV1+2/Syphilis | [11 empty] | 0 |  | 0.05654% | 0
542 | Diagnostic Test Kits | In Vitro Diagnostics | Syphilis Rapid Test | [11 empty] | 0 |  | 0.05654% | 0
543 | Diagnostic Test Kits | In Vitro Diagnostics | HIV Self Test | [11 empty] | 0 |  | 0.05654% | 0
544 | Diagnostic Test Kits | In Vitro Diagnostics | HIV AntibodyTest | [11 empty] | 0 |  | 0.05654% | 0
545 | Diagnostic Test Kits | In Vitro Diagnostics | Urine Test – strips | [11 empty] | 0 |  | 0.05654% | 0
546 | Diagnostic Test Kits | In Vitro Diagnostics | Blood Test - Anti A | [11 empty] | 0 |  | 0.05654% | 0
547 | Diagnostic Test Kits | In Vitro Diagnostics | Blood Test - Anti B | [11 empty] | 0 |  | 0.05654% | 0
548 | Diagnostic Test Kits | In Vitro Diagnostics | Blood Test - Anti A+B | [11 empty] | 0 |  | 0.05654% | 0
549 | Diagnostic Test Kits | In Vitro Diagnostics | Test Rhesus, anti D | [11 empty] | 0 |  | 0.05654% | 0
550 | Diagnostic Test Kits | In Vitro Diagnostics | Test pregnancy, strip, temperature stable | [11 empty] | 0 |  | 0.46% | 0
Secondary sources (please add lines as needed if you propose secondary sources)
Bidders Full Name:
Please indicate any cumulative volume percentage discount levels that increase as the cumulative value/volume increases throughout the validity of the LTA:
Please indicate any early payment percentage discount within a specified period of time faster than UNFPA's standard payment of net 30 days:
Name and Title
Signature
Date

## 2. Technical Info Pharma
Tab 2. Technical Information Form for the Pharmaceuticals
| Name of Company (Please enter in the cell on the right): | to be filled-in
Other comments (if any):
Important guidance to Bidders:
1. Please fill in the blanks under orange columns.
2. Technical information Form should be submitted in excel form.
3. Complete every single row following Technical Requirements for Pharmaceutical Products.
4. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font (red).
5. Please list the secondary sources, where applicable, in the end under sub-section "secondary source".
6. Bidders are not allowed to copy and paste UNFPA generic specifications or to use general statements like: “offer as per specifications”, “comply”, “offered as requested”, “as per requirement”, “yes, but different pack size instead - pack of 1000“ or similar general statements/information. Similar statements might invalidate the bid. Deviations between this sheet and product literature (e.g, manufacturer brochures/webpage) may invalidate the bid.
PHARMACEUTICALS PRODUCT LIST | [9 empty] | TECHNICAL INFORMATION
[7 empty] | Product Identification Annex A | [6 empty] | Product packaging Type Annex B & C |  | Manufacturer identification |  |  | Registration (licencing) status Annex D - H | [8 empty] | Packaging and Labelling Annex I - J |  |  |  |  | Active Pharmaceutical Ingredient (API) Annex K - P2 | [9 empty] | Status of Finished Pharmaceutical Product (FPP) Annex Q |  |  |  |  | Validation of Analytical methods - FPP Annex R - S | [7 empty] | Process validation and sterilization methods Annex T - U | [6 empty] | Safety & Efficacy and/or Therapeutic Equivalence Annex Z - AD
Bid Item No. | Product Sub-Category | UNFPA Short Item Description | UNFPA Item Description | Packaging | Storage conditions | Transportation conditions | Total shelf life | Product HS Code | International Nonproprietary Name (INN) | 1.1 Active Pharmaceutical Ingredient | 1.2 Strength | 1.3 Dosage form | 1.4 Route of administration (Oral, I.M, I.V., S.C, other - please specify) | 1.5 Pack size (please describe primary and secondary packaging) | 1.6 Secondary packaging material and pack size (e.g unit carton of 100s) | 2.1 Manufacturer's name | 2.2 Manufacturing Country (Country of origin) | 2.3 Complete address(es) of the manufacturing site(s) of the finished product(s), including address of site of final product release if different from the manufacturing site address. | 3.1 Is the product registered and currently marketed in a SRA country? YES/NO | 3.2 Country of SRA issuing authorty and SRA issuing authority Please specify | 3.3 Commercialization/ marketing authorisation status (SRA approval for marketing within SRA country or for export purpose only) | 3.4 Certificate of Pharmaceutical Product (CPP), Validity | 3.5 Registration in other countries/ Licensing status in other countries Describe | 3.6 This product is Prequalified by WHO? YES/NO, please provide reference and vaiidity | 3.7. Expert Review Panel (ERP) recommended, YES/NO, please provide reference and vaiidity | 3.8 Has the dossier been submitted to ERP, ICRC, MSF, GDF, WHO, or UNICEF? Describe YES, name of Agency or NO. | Countries of registration | 4.1 Primary packaging label language (English/French/Spanish) | 4.2 Primary packaging information (State compliance with Pharmacopeias, and describe material) | 4.3 Secondary packaging label language (English/French/Spanish) | 4.4 Patient Information Leaflet (PIL) and cartons artworks according to WHO guidelines? YES/NO | 4.5 Labelling and storage conditions in line with WHO guidelines for stability of FPP? YES/NO | 5.1 API Manufacturer Name(s) | 5.2 Address(es) of API manufacturer(s) (For more than 1, use different colours to identify) | 5.3 Agency/Authority issuing the approval Please specify | 5.4 Date of last inspection Please specify | 5.5 GMP Certificate validity Please specify | 5.6 Scope (Production Line) Describe | 5.7 Is/are the API(s) used to manufacture this product WHO-prequalified? YES/NO | 5.8 Are the technical specification parameters according to pharmacopeia monographs? YES/NO | 5.9 Summary of results from 3 Certificate of Analysis (CoAs) Please describe | 5.10 If CEP available – compare with specifications, validity and site of manufacture | 6. 1 All manufacturing sites involved into production Please specify | 6. 2 GMP inspecting agency for each manufacturing site Please specify | 6.2 Date of last inspection for each manufacturing site Please specify | 6.3 GMP Certificate validity for each manufacturing site Please specify | 6.4 Scope (Production Line) Describe | 7.1 Finished Product Analytical Reference Adequacy (BP/USP/Int.Ph., additional tests and in house methods) | 7.2 Are the technical specifications according to pharmacopeia monographs? YES/NO | 7.3 Summary of the CoAs results from the three last batches released Describe | 7.4 Are the analytical methods validated? YES/NO | 8.1 Have the manufacturing methods for each standard batch size been validated? YES/NO | 8.2 Product related critical quality attributes Describe | 8.3 Process related critical quality attributes Describe | 8.4 Is there validation of sterilization methods? (For sterile products) YES/NO | 9.1 Long-term stability studies conditions Describe | 9.2 Long-term stability studies duration Describe | 9.3 Long-term stability studies periods of testing Describe | 9.4 Long-term stability studies packaging material Describe | 9.5 Storage conditions and shelf-life supported by data Please specify | 9.6 Product suitable for use in which ICH Climatic Zones? (Zone I, II, II, IVa, IVb) | 9.7 Special storage conditions. Describe, if applicable. | 10. 1 Bioequivalence (BE) study design and conduct Describe, if applicable. | 10.2 Reference product Describe, if applicable. | 10.3 For in-vitro studies, BE study design and conduct Describe, if applicable. | 10.4 Reference product for in-vitro studies Describe, if applicable. | 10.5 AUC Ratio Describe, if applicable. | 10.6 Cmax Ratio Describe, if applicable. | 11.1. Specify whether it or the manufacturer has been involved with product recall or in litigation and arbitration, or whether any FDA 483 warning letters have been issued against the company in the past five years. | 11.2 Specific shipping conditions, if any. Describe, if applicable.
1 | Orals | Amoxicillin, powder oral suspension, 125mg/5ml, 100 ml, pack of 10 or 40 bottles | Amoxicillin Strength: 125mg/5ml Dosage form: powder for suspension Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 100ml Primary pack: HDPE bottle with suitable measuring device Secondary pack: 10 bottles or 40 bottles in cardboard boxes Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
2 | Orals | Amoxicillin trihydrate, 250mg, packs of 15, 20, 21, 50, 100, 500 or 1000 capsules | Amoxicillin Trihydrate Strength: 250mg Dosage form: Capsule Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 15, 20, 21, 50, 100, 500 and 1000 capsules Primary pack: Polypropylene/polyethylene containers with 1000, 500, 100, 50, 21, 20 and 15s or blister foil packs 15s and 21s Secondary pack: Cardboard boxes Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Store in tightly closed original container, protect from light | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
3 | Orals | Amoxicillin trihydrate, 500mg, packs of 15, 20, 21, 100, 500 or 1000 bottles or 15 or 21 blister strips | Amoxicillin Trihydrate Strength: 500mg Dosage form: capsule Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 15,20,21,100,500,1000s in HDPE and 15,21 blister strips Primary pack: HDPE bottle and blister strips Secondary pack: Cardboard boxes Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
4 | Orals | Amoxicillin trihydrate suspension, 250mg/5ml, in 100ml, pack of 10 or 40 bottles | Amoxicillin Trihydrate Strength: 250mg/5ml Dosage form: suspension Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 100ml Primary pack: HDPE bottle with suitable measuring device Secondary pack: Packs of 10 bottles or 40 bottles in cardboard boxes Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
5 | Parenterals | Ampicillin sodium, powder for injection, 1000mg (1g), pack of 1, 2, 5, 10, 20, 25 or 50 vials | Ampicillin Sodium Strength: 1000mg Dosage form: Powder for injection Route of administration: IV or IM | Primary pack : Type I/III glass Vials Secondary pack: Pack of 1/2/5/10/20/ 25/ 50 Vials with package insert in card board box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
6 | Parenterals | Ampicillin sodium, powder for injection, 500mg, pack of 50 vials | Ampicillin sodium Strength: 500mg Dosage form: Powder for injection Route of administration: IV or IM | Primary pack: Type 1 colorless borosilicate glass vials Secondary pack: 50 vials in a card board box with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
7 | Parenterals | Anti-D (RhO) immunoglobulin, 300mcg (winrho SDF), pack of 1 or 5 vials | Anti-D (RhO) immunoglobulin Strength: 300mcg/2ml Dosage form: Solution for injection Route of administration: IV or IM | Primary pack: Type 1 glass vial (amber colored vial or packaged to protect from light) Secondary pack: 1 vial in a pack or 5 vials in a multi pack with package insert Label & Carton text: Trilingual Package Insert: Trilingual | 2 to 8°C. Do not store freeze. Protect from light | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight (for up to 3 days) and smaller aircrafts Active cooled temperature for airfreight (longer than 3 days) Active container for sea freight
8 | Orals | Atazanavir (ATV) 300mg + Ritonavir (r) 100mg, pack of 30 tablets | Atazanavir (ATV) + Ritonavir (r) Strength: Atazanavir 300mg + Rito 100mg Dosage form: Tablets Route of administration: oral BCS Class: Class 4/2 Bioequivalence required | Pack size: 30s Primary pack: Alu/Alu Blisters HDPE Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage requirements | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
9 | Parenterals | Atropine sulphate 1mg/ml injection in 1ml ampoule, pack or 10, 20, 50 or 100 ampoules | Atropine Sulphate Strength: 1mg/ml in 1ml ampoule Dosage form: Solution for injection Route of administration: IV, IM or SC | Primary pack: Type 1 glass ampoules (amber colored or packaged to protect from light) Secondary pack: Pack or 10, 20, 50, 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
10 | Orals | Abacavir (as sulphate) 120mg + Lamivudine 60mg, pack of 60 dispersible tablets | Abacavir (as sulphate) + Lamivudine Dispersible Tablets Strength: 120mg Abacavir (as sulphate) + 60 mg Lamivudine Dosage form: dispersible tablet Route of administration: oral Additional requirements: Tablets should be scored with at least one break line to allow for division of tablets into two parts, each containing the same amount of active ingredient BCS Class: Lamivudine - Class 3 Abacavir (as sulphate) - Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 60s Primary packaging: White HDPE bottles Secondary packaging: Cardboard boxes Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
11 | Orals | Azithromycin Anhydrous, 250mg, pack of 1, 6, 10 or 12 capsules | Azithromycin Anhydrous (if it comes with salts, salts must be clearly indicated) Strength: 250mg Dosage form: capsule Route of administration: oral | Pack size: 1s 6s, 10s, 12s Primary pack: Aluminum PVC Blister pack, Alu foil Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
12 | Orals | Azithromycin, 250mg, pack of 4 or 6 tablets | Azithromycin Strength: 250mg Dosage form: tablet Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 4s and 6s Primary pack: Aluminum/PVC blister strips Secondary pack: 1 strip in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
13 | Orals | Azithromycin, 500mg, pack of 3 or 500 tablets | Azithromycin Strength: 500mg Dosage form: tablet Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 3s and 500s Primary pack: Aluminum/PVC blister strips HDPE bottle / Polythene container Secondary pack: 1 strip in a cardboard box 1 bottle in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
14 | Orals | Azithromycin dihydrate, 200mg base/ 5ml suspension, 15ml bottle | Azithromycin Dihydrate Strength: 200mg base / 5ml Dosage form: powder for oral suspension Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 15ml Primary pack: Polypropylene bottle with child resistant screw cap (with or without a tamper evident seal). Secondary pack: Bottle in a cardboard box containing a double-ended multi-dosing spoon and 10 ml oral dosing syringe with detachable adaptor. Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
15 | Parenterals | Benzathine benzylpenicillin, 1.44G (2.4 M.I.U) powder for injection, 5ml or 10ml, pack of 1, 10, 50 or 100 vials | Benzathine Benzylpenicillin Strength: 1.44G (2.4 M.I.U) in 5ml or 10ml vial Dosage form: Powder for injection Route of administration: deep IM | Primary pack: Type 1 glass vial Secondary pack: Pack of 1, 10, 50 or 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
16 | Parenterals | Benzathine benzylpenicillin, 900mg (1.2 M.I.U) powder for injection, 5ml or 10ml, pack of 1, 10, 50 or 100 vials | Benzathine Benzylpenicillin Strength: 900mg (1.2 M.I.U) in 5ml or 10ml vial Dosage form: Powder for injection Route of administration: deep IM | Primary pack: Type 1 glass vial Secondary pack: Pack of 1, 10, 50 or 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
17 | Parenterals | Betamethasone sodium phosphate 3mg base/ml + betamethasone acetate 3mg base/ml in an aqueous vehicle, pack of 20 vials | Betamethasone 6 mg/ml injection Strength: betamethasone sodium phosphate 3mg base/ml + betamethasone acetate 3mg base/ml in an aqueous vehicle vial Dosage form: Solution for injection Route of administration: percutaneous infiltration | Primary pack: Type 1 glass vial (amber colored or packaged to protect from light) Secondary pack: 20 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
18 | Parenterals | Bupivacaine hydrochloride 0.5%, 10 ml, pack of 10, 20, 50 or 100 vials | Bupivacaine hydrochloride 0.5% Strength: 0.5% w/v Dosage form: Solution for injection Route of administration: Infiltration by injection. Epidural. Caudal | Primary pack: 10ml translucent vials Secondary pack: 10, 20,50, 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
19 | Parenterals | Bupivacaine hydrochloride 0.5%, 20 ml, pack of 10, 20, 50 or 100 vials | Bupivacaine hydrochloride 0.5% Strength: 0.5% w/v Dosage form: Solution for injection Route of administration: Infiltration by injection. Epidural. Caudal | Primary pack: 20ml translucent vials Secondary pack: 10, 20,50, 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
20 | Parenterals | Bupivacaine hydrochloride 0.5% 4ml ampoule + 7.5% glucose solution, pack of 5, 10 or 20 ampoules | Bupivacaine hydrochloride 0.5% 4ml ampoule + 7.5% glucose solution Strength: 0.5 mg/ml in 4ml ampoule to be mixed with 7.5% glucose Dosage form: Solution for injection Route of administration: percutaneous infiltration | Primary pack: Type 1 glass ampoule (amber colored or packaged to protect from light) Secondary pack: 5, 10, 20 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
21 | Parenterals | Calcium gluconate, 100mg base/ml injection, 10ml, pack of 10, 20 or 50 ampoules | Calcium Gluconate Strength: 100mg base/ml in 10ml ampoule Dosage form: solution for injection Route of administration: IV | Primary pack: Type 1 glass ampoule Secondary pack: Pack of 10, 20, 50 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
22 | Parenterals | Carbetocin, 100mcg/ml in 1 ml ampoule, pack of 10, 20 or 50 ampoules | Carbetocin Strength: 100mcg/ml Dosage form: Solution for injection Route of administration: IV or IM | Primary pack: Type 1 colorless borosilicate glass ampoules Secondary pack: 5 or 10x1 ml with pack insert in cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accomodate zove IVb climatic conditions. Store below 30°C, protect from light | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
23 | Orals | Cefixime, 200mg, pack of 8 or 10 tablets | Cefixime Strength: 200mg Dosage form: tablet Route of administration: oral BCS Class: Class 4 Bioequivalence (BE) is required | Pack size: 8s and 10s Primary pack: Aluminum/PVC blister strips Secondary pack: 1 strip in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
24 | Orals | Cefixime, 400mg, pack of 8, 10 or 100 tablets | Cefixime Strength: 400mg Dosage form: tablet Route of administration: oral BCS Class: Class 4 Bioequivalence (BE) is required | Pack size: 8s, 10s and 100s Primary pack: Aluminum/PVC blister strips HDPE bottle / Polythene / Polypropylene container Secondary pack: 1 strip in a cardboard box 1 bottle in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
25 | Orals | Cefixime trihydrate, 100mg/5ml, powder oral suspension, bottle of 30ml, 40ml or 50ml | Cefixime Trihydrate Strength: 100mg / 5ml Dosage form: powder for oral suspension Route of administration: oral BCS Class: Class 4 Bioequivalence (BE) is required | Pack size: 30ml, 40ml, 50ml Primary pack: Polypropylene bottle with child resistant screw cap (with or without a tamper evident seal). Secondary pack: Bottle in a cardboard box with appropriate measuring device. Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
26 | Orals | Cephalexin, 500mg, pack of 21, 100 or 500 capsules | Cephalexin Strength: 500mg Dosage form: capsule Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 21s, 100s and 500s Primary pack: Aluminum - PVC/PVDC blister strips HDPE bottle / Polythene / Polypropylene container Secondary pack: 1 strip in a cardboard box 1 bottle in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
27 | External Use | Chlorhexidine digluconate 7.1% w/v solution, bottle of 10ml with a dripper | Chlorhexidine Digluconate Product name composition: Chlorhexidine digluconate (if it comes with salts, salts must be clearly indicated) Strength:7.1% w/v solution Dosage form: Solution Route of administration: Topical | Pack size: 10ml nozzle/dropper plastic bottle Primary pack: 10ml nozzle/dropper plastic bottle Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zone IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products. Sunlight adversely affects the stability of chlorhexidine digluconate, transparent primary containers should be avoided | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
28 | External Use | Chlorhexidine digluconate solution 5%, 1 litre bottle | Chlorhexidine Digluconate Product name composition: Chlorhexidine digluconate (if it comes with salts, salts must abe clearly indicated) Strength: 5% w/v solution Dosage form: Solution Route of administration: Topical | Pack size: 1000ml Primary pack: Plastic bottle Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products. Sunlight adversely affects the stability of chlorhexidine digluconate, transparent primary containers should be avoided | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
29 | External Use | Chlorhexidine digluconate 7.1% (equivalent to 4% chlorhexidine) gel, 10g or 20g, aluminium tube | Chlorhexidine Digluconate Product name composition: Chlorhexidine gluconate 7.1% (equivalent to 4% chlorhexidine) gel (if it comes with salts, salts must be clearly indicated) Strength: 7.1% Dosage form: Gel Route of administration: Topical | Pack size: 10g or 20g Primary pack: Aluminum Tube Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products. Store away from direct sunlight | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
30 | External Use | Chlorhexidine gluconate 7.1% (equivalent to 4% chlorhexidine) gel, 3g, pack of 7 plastic individual sachets | Chlorhexidine Gluconate Product name composition: Chlorhexidine gluconate 7.1% (equivalent to 4% chlorhexidine) gel (if it comes with salts, salts must be clearly indicated) Strength: 7.1%, 3g Dosage form: Gel Route of administration: Topical | Pack size: 7 sachets Primary pack: Plastic laminated individual sachets Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions.30 °C storage conditions. Light protection for light sensitive products. Store away from direct sunlight | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
31 | External Use | Chlorhexidine gluconate solution 4.0% (detergent solution), 500ml bottle | Chlorhexidine Gluconate (detergent solution) Product name composition: Chlorhexidine gluconate (if it comes with salts, salts must be clearly indicated) Strength: 4% w/v solution Dosage form: Solution Route of administration: Topical | Pack size: 500ml Primary pack: Plastic bottle Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
32 | Orals | Ciprofloxacin, 500mg, pack of 10 or 100 tablets | Ciprofloxacin Strength: 500mg Dosage form: tablet Route of administration: oral BCS Class: Class 4 Bioequivalence (BE) is required | Pack size: 10s, 100s Primary pack: Aluminum - PVC/PVDC blister strips HDPE bottle / Polythene / Polypropylene container Secondary pack: 1 strip in a cardboard box 1 bottle in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
33 | Parenterals | Clindamycin injection, 125mg/ml in 4 ml ampoule/vial (600mg/4ml), pack of 1, 5, 10, 20 or 50 vials/ampoules | Clindamycin Strength: 125mg/ml in 4 ml ampoule/vial (600mg/4ml ) Dosage form: solution for injection Route of administration: IV, IM | Primary pack: Type 1 glass ampoule / vials Secondary pack: 1, 5, 10, 20 or 50 vials/ ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
34 | Vaginal | Clotrimazole, 100mg, vaginal tablet | Clotrimazole Strength: 100mg Dosage form: tablet Route of administration: vaginal BCS Class: Class 2 Exempt from bioequivalence but should comply with pharmacopeial tests | Pack size: 1 Primary pack: Aluminum foil blister Secondary pack: 1 tablet and applicator in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
35 | Vaginal | Clotrimazole, 500mg, vaginal tablet | Clotrimazole Strength: 500mg Dosage form: tablet Route of administration: vaginal BCS Class: Class 2 Exempt from bioequivalence but should comply with pharmacopeial tests | Pack size: 1 Primary pack: Aluminum foil blister Secondary pack: 1 tablet and applicator in a cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30°C storage conditions | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
36 | Parenterals | Cloxacillin, as sodium salt, 500mg, pack of 10 or 50 vials | Cloxacillin as Sodium salt Strength: 500mg Dosage form: powder for solution for injection Route of administration: IV, IM | Primary pack: Type 1 glass vials Secondary pack: Pack of 10 and 50 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
37 | Parenterals | Dexamethasone sodium phosphate, 4mg base/ml, 1ml, pack of 10, 50 and 100 ampoules | Dexamethasone Sodium Phosphate Strength: 4 mg base/ml in 1ml ampoule Dosage form: solution for injection Route of administration: IV, IM, local injection | Primary pack: Type 1 glass ampoule Secondary pack: Pack of 10, 50 and 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
39 | External Use | Disinfectant tablets for water, 1.67g NaDCC, pack of 10, 100 or 200 tablets | Disinfectant tablets for water Product name composition: 1.67 g NaDCC effervescent tablet, releasing 1 g active chlorine when dissolved in water. (if it comes with salts, salts must be clearly indicated) Strength: 1.67g Dosage form: Tablet Route of administration: Topical | Pack size: 10s, 100s or 200s Primary pack: Alu/Alu Blister Packs Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light, heat and moisture protection for light, heat and moisture sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
40 | Orals | Doxycycline, 100mg, pack of 8, 10 or 12 tablets | Doxycycline Product name composition: Available as hydrochloride, hyclate for capsules or monohydrate for dispersible tablets. Specific salt to be specified. Strength: 100mg as doxycycline monohydrate Dosage form: dispersible tablets Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 8s, 10s, 12s Primary pack: Aluminum/ PVC Blister pack Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
41 | Orals | Dolutegravir, 50mg, pack of 30 film coated tablets | Dolutegravir 50mg Strength: 50mg Dosage form: film coated tablet Route of administration: oral BCS Class: Class 2 Bioequivalence is required | Pack size: 30 Primary packaging: HDPE (high density polyethylene) bottles closed with child resistant polypropylene screw closures, with a polyethylene faced induction heat seal liner with dessicant Secondary packaging: Cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
42 | Orals | Dolutegravir, 10mg, pack of 90 dispersible tablets | Dolutegravir 10mg Strength: 10mg Dosage form: dispersible tablets Route of administration: oral BCS Class: Class 2 Bioequivalence is required | Pack size: 90 Primary packaging: HDPE (high density polyethylene) bottles closed with child resistant polypropylene screw closures, with a polyethylene faced induction heat seal liner with dessicant Secondary packaging: Cardboard box Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
43 | Orals | Efavirenz 600mg+Emtricitabine 200mg+Tenofovir 300mg, pack of 28, 30 or 500 tablets | Efavirenz 600mg+Emtricitabine 200mg+Tenofovir 300mg Strength: 600 mg + 200 mg + 300 mg (tenofovir disoproxil fumarate – equivalent to 245 mg tenofovir disoproxil). Dosage form: tablets Route of administration: oral BCS Class: Class II/I/III Bioequivalence required | Pack size: 28's, 30's and 500's Primary pack: HDPE bottles with silica gel Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
44 | Orals | Emtricitabine 200mg + Tenofovir 300mg, pack of 30 tablets | Emtricitabine + Tenofovir Product Name composition: Tenofovir available as tenofovir disoproxil fumarate, specific salt to be specified. Strength: 200 mg Emtricitabine and 300 mg Tenofovir Disoproxil Fumarate (TDF) equivalent to 245 mg of tenofovir disoproxil or 136 mg of tenofovir. Dosage form: tablets Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 30s Primary pack: HDPE Bottle with silica gel desiccant sachet. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light, heat and moisture protection for light, heat and moisture sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
45 | Parenterals | Ephedrine hydrochloride injection, 30 mg/ml, pack of 5 or 10 ampoules | Ephedrine hydrochloride Strength: 30 mg/ml Dosage form: solution for injection Route of administration: IM | Primary pack: 1 ml colorless glass one-point-cut (OPC) ampoules, type I containing 1 ml solution for injection Secondary pack: Cartons of 5 or 10 ampoules Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
46 | Parenterals | Epinephrine (Adrenaline) as hydrochloride or hydrogen tartrate injection, 1mg/ml in 0.5, 1, 2, 5 or 10 ml ampoule, pack of 10 ampoules | Epinephrine (Adrenaline) as hydrochloride or hydrogen tartrate Strength: 1mg/ml Dosage form: solution for injection Route of administration: IM | Primary pack: 1 ml in type 1 colorless neutral OPC (one point cut) glass ampoules Secondary pack: Cartons of 10 ampoules Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
47 | Parenterals | Ergometrine maleate, 0.2mg base/ml injection in 1ml ampoule, pack of 6 or 10 ampoules | Ergometrine Maleate Strength: 0.2 mg base/ml in 1ml ampoule Dosage form: Solution for injection Route of administration: IV or IM | Primary pack: Type 1 glass ampoule (amber colored or packaged to protect from light) Secondary pack: 6, 10 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
48 | Orals | Ferrous salts/Folic acid coated (60mg/0.4mg), pack of 100 or 1000 tablets | Ferrous salts/Folic acid coated (60mg/0.4mg) Product name composition: Ferrous fumarate or gluconate or sulphate (if it comes with other salts, salts must be clearly indicated) Strength: 60mg elemental iron and 400mcg folic acid Dosage form: tablets Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 100s, 1000s Primary pack: HDPE securitainer or Aluminum PVC Blister pack, Alu foil, Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
49 | Orals | Ferrous salts (60mg Fe), pack of 50, 100, 250, 500 or 1000 sugar coated tablets | Ferrous salts (60mg elemental iron Fe) Product name composition: Ferrous fumarate or gluconate or sulphate (if it comes with other salts, salts must be clearly indicated) Strength: 60/65mg elemental iron Dosage form: sugar coated tablets Route of administration: oral BCS Class: Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 50s, 100s, 250s, 500s and 1000s. Primary pack: Screw cap plastic containers or white opaque PVC/PVDC blister Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No further special conditions required. | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
50 | Orals | Folic acid, 5mg, pack of 7, 10, 14, 21, 28, 30, 56, 60, 84, 90, 100, 112, 500 or 1000s tablets | Folic acid Strength: 5mg Dosage form: tablets Route of administration: oral BCS Class: Class 1/3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size:7s, 10s, 14s, 21s, 28s, 30s, 56s, 60s, 84s, 90s, 100s, 112s, 500s, 1000s Primary pack: rigid injection moulded polypropylene or injection blow-moulded polyethylene containers and snap-on polyethylene lids; in case any supply difficulties should arise the alternative is amber glass containers with screw caps., PVC/PVDC/Alu blisters Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
51 | Parenterals | Gentamicin sulphate, 10mg/ml injection in 2ml ampoule, pack of 100 ampoules | Gentamicin Sulphate Strength: 10mg/ml in 2ml Dosage form: solution for injection Route of administration: IV, IM | Primary pack: Type 1 glass ampoules Secondary pack: 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
52 | Parenterals | Gentamicin sulphate, 40mg base/ml injection in 2ml ampoule, pack of 10, 50 or 100 ampoules | Gentamicin Sulphate Strength: 40mg base/ml in 2ml Dosage form: solution for injection or infusion Route of administration: IV, IM | Primary pack: Type 1 glass ampoules Secondary pack: 10, 50 or100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
54 | Parenterals | Glucose, 10% injection in pack of 20 or 50 ampoules | Glucose Strength: 10% Dosage form: solution for injection Route of administration: IV | Primary pack: 10 ml polypropylene ampoule Secondary pack: 20, 50 ampoules Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
55 | Parenterals | Glucose, solution for infusion, 5% (iso-osmotic), pack of 30 or 36 bags of 250 ml, pack of 20 or 24 bags for 500 ml or pack of 10 or 12 bags for 1000 ml+ Infusion-giving set, sterile, single use | Glucose + Infusion-giving set, sterile, single use Strength: 5% (iso-osmotic) Dosage form: solution for infusion Route of administration: IV | Primary pack: 250 ml, 500 ml and 1000 ml bag (e.g Viaflo) Secondary pack: 30 or 36 bags for 250ml per carton/ 20 or 24 bags for 500ml per carton/ 10 or 12 bags of 1000ml per carton including the infusion giving set with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
56 | Parenterals | Hartmann's sol. (Ringer's lactate) intravenous infusion, 0.5L + giving set, 1 set or pack of 20 sets | Hartmann's sol. (Ringer's lactate), 0.5L + giving set Strength: 5% Dosage form: solution for infusion Route of administration: IV | Primary pack: 0.5L bag (e.g Viaflo) Secondary pack: 1 set or pack of 20 including the infusion giving sets with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
57 | Parenterals | Hartmann's sol. (Ringer's lactate) intravenous infusion, 1L + giving set, 1 set or pack of 12 sets | Hartmann's sol. (Ringer's lactate), 1L + giving set Strength: 5% Dosage form: solution for infusion Route of administration: IV | Primary pack: 1 litre bag (e.g Viaflo) Secondary pack: Packs of 10 or 12 including the infusion giving sets with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
58 | Parenterals | Heparin sodium 5000IU/ml neutral glass ampoules (Type I Ph Eur) of 1ml capacity OR 5ml capacity containing 1ml OR 5ml of solution respectively | Heparin Sodium Strength: 5000IU in 1ml or 25000IU in 5ml Dosage form: Solution for injection or concentrate for solution for infusion Route of administration: Continuous intravenous infusion in 5% glucose or 0.9% sodium chloride or by intermittent intravenous injection, or by subcutaneous injection. | Primary pack: 1ml capacity OR 5ml capacity containing 1ml OR 5ml of solution respective Secondary pack: 1 x 10 ampoules Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
59 | Parenterals | Hydralazine hydrochloride 20mg/2ml powder for injection in 2ml ampoule | Hydralazine Hydrochloride Strength: 20mg/2ml in 2ml ampoule Dosage form: powder for solution for injection Route of administration: IV, IM | Primary pack: Type 1 glass ampoules (amber colored or packaged to protect from light) Secondary pack: 5 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
60 | Orals | Hydralazine hydrochloride, 25mg, pack of 28, 30, 56, 60, 84, 90, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | Hydralazine Hydrochloride Product name composition: Hydralazine available as hydrochloride salt. Strength: 25mg Dosage form: tablets Route of administration: oral BCS Class: Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 28s, 30s, 56s, 60s, 84s, 90s, 100s, 112s, 120s, 168s, 180s, 250s, 500s, 1000s. Primary pack: PVC/PVDC/Alu blisters Polyethylene container with a polypropylene lid. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
61 | Orals | Hydralazine hydrochloride, 50mg, pack of 28, 30, 56, 60, 84, 90, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | Hydralazine Hydrochloride Product name composition: Hydralazine available as hydrochloride salt. Strength: 50mg Dosage form: tablets Route of administration: oral BCS Class: Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 28s, 30s, 56s, 60s, 84s, 90s, 100s, 112s, 120s, 168s, 180s, 250s, 500s, 1000s. Primary pack: PVC/PVDC/Alu blisters Polyethylene container with a polypropylene lid. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
62 | Orals | Ibuprofen, 400mg, pack of 6, 10, 12, 15, 18, 20, 24, 30, 40, 50, 56, 60, 84 or 100 | Ibuprofen Strength: 400mg Dosage form: tablets Route of administration: oral BCS Class: Class 2 Bioequivalence (BE) required. | Pack size: 6s, 10s, 12s, 15s, 18s, 20s, 24s, 30s, 40s, 50s, 56s, 60s, 84s and 100s Primary pack: PVC/PVdC/Alu blisters PVC/ Alu blisters. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. No special storage conditions required | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
63 | Parenterals | Ketamine hydrochloride, 50mg base/ml injection in 10ml vial, pack of 1, 5 or 25 vials | Ketamine Hydrochloride Strength: 50mg base/ml in 10ml vial Dosage form: solution for injection Route of administration: IV, IM | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light) Secondary pack: 1, 5, 25 vials with package insert per pack Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
64 | Orals | Lamivudine, 150mg, pack of 10 or 60 tablets | Lamivudine Strength: 150mg Dosage form: tablets Route of administration: oral BCS Class: Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 10s, 60s Primary pack: PVC/PVdC/Alu blisters HDPE botlles with polypropylene child resistant caps. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
65 | Parenterals | Lidocaine hydrochloride 1% injection in 10ml vial, pack of 20 vials | Lidocaine hydrochloride 1% Strength: 10mg/ml in 10ml vial Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light) Secondary pack: 20 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
66 | Parenterals | Lidocaine hydrochloride 1% injection in 20ml vial/ampoule, pack of 1, 10 or 20 vials/ampoules | Lidocaine hydrochloride 1% Strength: 10mg/ml in 20ml vial/ ampoule Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light)// ampoules Secondary pack: 1, 10 or 20 vials// ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
67 | Parenterals | Lidocaine hydrochloride 1% injection in 50ml vial, pack of 10 or 25 vials | Lidocaine hydrochloride 1% Strength: 10mg/ml in 50ml vial Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light) Secondary pack: 10 or 25 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
68 | Parenterals | Lidocaine hydrochloride 2% injection in 20ml vial, pack of 10 or 20 vials | Lidocaine hydrochloride 2% Strength: 20mg/ml in 20ml vial Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light Secondary pack: 10 or 20 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
69 | Parenterals | Lidocaine hydrochloride 2% injection in 50ml vial, pack of 25 vials | Lidocaine hydrochloride 2% Strength: 20mg/ml in 50ml vial Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light) Secondary pack: 25 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
70 | Parenterals | Lidocaine hydrochloride 2% injection in 5ml vial, pack of 10, 20 or 100 vials | Lidocaine hydrochloride 2% Strength: 20mg/ml in 5ml vial Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass vials (amber colored or packaged to protect from light) Secondary pack: 10, 20 or 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
71 | Parenterals | Lidocaine hydrochloride 5% (50mg/ml) injection in 2ml ampoule, pack of 100 ampoules | Lidocaine hydrochloride 5% Strength: 50mg/ml in 2ml ampoule Dosage form: solution for infiltration, injection or infusion Route of administration: percutaneous, epidural, IV, IM or SC | Primary pack: Type 1 glass ampoules (amber colored or packaged to protect from light) Secondary pack: 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
72 | Parenterals | Magnesium sulphate 500mg/ml injection in 10ml ampoule, pack of 1, 10, 50 or 100 ampoules | Magnesium Sulphate Strength: 500mg/ml in 10ml ampoule Dosage form: solution for injection or infusion Route of administration: IV or IM | Primary pack: type 1 glass ampoules Secondary pack: Pack of 1, 10, 50 or 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
73 | Parenterals | Magnesium Sulphate 500mg/ml injection in 2ml ampoule, pack of 10 ampoules | Magnesium Sulphate Strength: 500mg/ml in 2ml ampoule Dosage form: solution for injection or infusion Route of administration: IV or IM | Primary pack: type 1 glass ampoules Secondary pack: Pack of 10 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
74 | Orals | Methyldopa, 250mg, packs of 28, 56, 60, 84, 100, 112, 120, 168, 180, 250, 500 or 1000 tablets | Methyldopa Strength: 250mg Dosage form: tablets Route of administration: oral BCS Class: Class 3 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 28s, 56s, 60s, 84s, 100s, 112s, 120s, 168s, 180s, 250s, 500s, 1000s. Primary pack: PVC/PVdC/Alu blisters HDPE container Polypropylene securitainer. PVC/ Alu Blisters Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
75 | Orals | Metronidazole, 250mg, pack of 50, 100, 250 or 500 tablets | Metronidazole Strength: 250mg Dosage form: tablets Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 50s, 100s, 250s,500s. Primary pack: Polyethylene container with pilfer-proof polyethylene closure, Amber glass bottles, PVC/ Alu Blisters, Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
76 | Parenterals | Metronidazole, for intravenous infusion, 5mg/ml, 100-ml, pack of 1 or 20 PVC free flexi-bags | Metronidazole Strength: 5mg/ml Dosage form: solution for infusion Route of administration: IV | Primary pack: 100ml PVC free flexi-bags (keep in original packaging to protect from light) Secondary pack: 1 or 20 bags with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
77 | Parenterals | Metronidazole, for intravenous infusion, 5mg/ml, 100-ml, semi-rigid PVC free bottle | Metronidazole Strength: 5mg/ml Dosage form: solution for infusion Route of administration: IV | Primary pack: 100ml semi-rigid PVC free bottle (keep in the outer packaging to protect from light) Secondary pack: 1 with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
78 | Orals | Mifepristone 200mg + 4 misoprostol 200mcg tablets (blister), pack of 5 (1mife + 4miso) tablets | Mifepristone + 4 misoprostol Strength: 200mg mife + 4 200mcg miso Dosage form: tablets Route of administration: oral for mife; oral, vaginal, buccal or sublingual route for miso BCS Class: Class 2/4 BE required | Pack size: 5s (1 mife+4 miso). Primary pack: OPA/Al/PVC/Al blister PVC/ Alu Blisters, Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
79 | Orals | Mifepristone, 200mg, pack of 1, 3, 15, 30 tablets | Mifepristone Strength: 200mg Dosage form: tablets Route of administration: oral BCS Class: Class 2/4 Bioequivalence required | Pack size: 1s,3s,15s,30s. Primary pack: PVC/ Alu Blisters Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
80.1 | Orals | Misoprostol, 200mcg, pack of 3 tablets | Misoprostol Strength: 200mcg Dosage form: tablets Route of administration: oral BCS Class: Class 2/4 Bioequivalence required | Pack size: 3s. Primary pack: Alu/ Alu Blisters Secondary pack: one strip per box, 10 boxes in a shrink wrap and 60 shrink wraps in a carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
80.2 | Orals | Misoprostol, 200mcg, pack of 4 tablets | Misoprostol Strength: 200mcg Dosage form: tablets Route of administration: oral BCS Class: Class 2/4 Bioequivalence required | Pack size: 4s. Primary pack: Alu/ Alu Blisters Secondary pack: one strip per box, 10 boxes in a shrink wrap and 60 shrink wraps in a carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
81 | Orals | Multivitamin coated, pack of 20, 84, 90, 100 or 500 tablets Vitamin A (retinol acetate) -800 micrograms retinol activity equivalent Vitamin C (ascorbic acid) -70 mg Vitamin D (cholecalciferol) - 5 micrograms (200 IU) Vitamin E (alpha tocopherol succinate) -10 mg alpha tocopherol equivalent Vitamin B1 (thiamine mononitrate) -1.4 mg Vitamin B2 (riboflavin) -1.4 mg Vitamin B3 (niacinamide) -18 mg niacin equivalent Vitamin B6 (pyridoxine hydrochloride) -1.9 mg Folic acid (folic acid) - 680 micrograms dietary folate equivalent (400 micrograms) Vitamin B12 (cyanocobalamin) - 2.6 micrograms Iron (ferrous fumarate) - 30 mg Iodine (potassium iodide) - 150 micrograms Zinc (zinc oxide) - 15 mg Selenium (sodium selenite) - 65 micrograms Copper (cupric oxide) - 2 mg | Multivitamin coated tablet Vitamin A (retinol acetate) -800 micrograms retinol activity equivalent Vitamin C (ascorbic acid) -70 mg Vitamin D (cholecalciferol) - 5 micrograms (200 IU) Vitamin E (alpha tocopherol succinate) -10 mg alpha tocopherol equivalent Vitamin B1 (thiamine mononitrate) -1.4 mg Vitamin B2 (riboflavin) -1.4 mg Vitamin B3 (niacinamide) -18 mg niacin equivalent Vitamin B6 (pyridoxine hydrochloride) -1.9 mg Folic acid (folic acid) - 680 micrograms dietary folate equivalent (400 micrograms) Vitamin B12 (cyanocobalamin) - 2.6 micrograms Iron (ferrous fumarate) - 30 mg Iodine (potassium iodide) - 150 micrograms Zinc (zinc oxide) - 15 mg Selenium (sodium selenite) - 65 micrograms Copper (cupric oxide) - 2 mg Strength: Multistrength Dosage form: tablets Route of administration: oral BCS Class: N/A | Pack size: 20s, 84s, 90s, 100s, 500s. Primary pack: Alu/ PVC Blisters. Secondary pack: Carton with package insert. Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Protect from light, heat & moisture | Freight condition Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
82 | Orals | Nifedipine, 10mg, pack of 90 or 100 immediate release capsule | Nifedipine immediate release Strength: 10mg Dosage form: capsules Route of administration: oral BCS Class: Class II Bioequivalence required | Pack size: 90s, 100s. Primary pack: Polypropylene/ Alu Blisters, HDPE Bottles Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
83 | Orals | Oxybutynin hydrochloride, 5mg, pack of 6, 20, 21, 28, 30, 50, 56, 60, 84 or 100 tablets | Oxybutynin Hydrochloride Strength: 5mg Dosage form: Tablets Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 6, 20, 21, 28, 30, 50, 56, 60, 84 and 100s. Primary pack: PVC/PVdC-Alu blister Clear PVC –Plain Alu blister pack Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions . 30 °C storage conditions. Moisture protection for moisture sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
84 | Parenterals | Oxytocin 10 I.U. base/ml injection in 1ml ampoule, pack of 5, 10 or 100 ampoules | Oxytocin Strength: 10 I.U base/ml, 1 ml ampoule Dosage form: Solution for injection Route of administration: I.V or I.M | Primary pack: Ampoules (amber colored packaged to protect from light) Secondary pack: 5, 10 or 100 ampoules in a multi pack with package insert Label & Carton text: Trilingual Package Insert: Trilingual | 2 to 8°C. Do not store freeze. Protect from light | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
85 | Orals | Paracetamol 500mg, tablet | Paracetamol Strength: 500mg Dosage form: tablets Route of administration: oral BCS Class: Class 1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size:10s, 20s, 50s, 100s, 300s. Primary pack: PVC/ Alu Blisters, HDPE Bottles Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light and moisture protection for light and moisture sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freigh
86 | External Use | Povidone iodine 10% solution for cutaneous use, 1000ml bottle | Povidone Iodine Product name composition: Povidone Iodine (if it comes with salts, salts must be clearly indicated) Strength: 10% w/v solution Dosage form: solution for cutaneous use Route of administration: Topical | Pack size: 1000ml Primary pack: Plastic bottle Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
87 | External Use | Povidone iodine 10% solution for cutaneous use, 500ml bottle | Povidone Iodine Product name composition: Povidone Iodine (if it comes with salts, salts must abe clearly indicated) Strength: 10% w/v solution Dosage form: solution for cutaneous use Route of administration: Topical | Pack size: 500ml Primary pack: Plastic bottle Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
88 | Parenterals | Quinine dihydrochloride 300 mg/ml, solution for dilution for infusion in 2 ml ampoule, pack of 10, 50 or 100 ampoules | Quinine Dihydrochloride Strength: 300 mg/ml in 2ml ampoule Dosage form: Solution for dilution for infusion or IM injection Route of administration: slow IV infusion or IM | Primary pack: Type 1 glass ampoules Secondary pack: 10, 50, 100 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
89 | Parenterals | Sodium Chloride 0.9% injection 10 ml ampoule | Sodium Chloride 0.9% injection Strength: 0.9% (9g/l) Dosage form: solution for infusion Route of administration: IV | Primary pack: 10ml ampoule Secondary pack: Pack of 20, 50, 100 ampoules, with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
90 | Parenterals | Sodium chloride 0.9% isotonic IV infusion in 1000ml bag + Infusion-giving set, sterile, single use, pack of 1 or 20 sets | Sodium chloride in 1000ml bag + Infusion-giving set, sterile, single use Strength: 0.9% (9g/l) Dosage form: solution for infusion Route of administration: IV | Primary pack: 1000ml bag (e.g Viaflo) Secondary pack: Pack of 1, 10, 12 or 20 including infusion giving sets with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
91 | Parenterals | Sodium chloride 0.9% isotonic IV infusion in 500ml bag + Infusion-giving set, sterile, single use, pack of 1, 10, 12 or 20 sets | Sodium Chloride in 500ml bag + Infusion-giving set, sterile, single use Strength: 0.9% (9g/l) isotonic Dosage form: solution for infusion Route of administration: IV | Primary pack: 500ml bag (e.g Viaflo) Secondary pack: Pack of 1 or 20 including infusion giving sets with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
92 | Orals | Tenofovir Disoproxil Fumarate 300mg + Lamivudine 300mg + Dolutegravir 50mg, pack of 30, 90 or 180 tablets | Tenofovir Disoproxil Fumarate 300mg + Lamivudine 300mg + Dolutegravir 50mg Strength: 300 mg + 300 mg + 50 mg (tenofovir disoproxil fumarate – equivalent to 245 mg tenofovir disoproxil). Dosage form: tablets Route of administration: oral BCS Class: Class III/III/II Bioequivalence required | Pack size: 30's, 90's and 180's Primary pack: HDPE bottles with silica gel Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
93 | External Use | Tetracycline hydrochloride eye-ointment 1% in tube of 5g - CANCELLED: TO BE REPLACED (no need to quote) | Tetracycline Hydrochloride Eye-ointment Product name composition: Tetracycline Hydrochloride (if it comes with salts, salts must be clearly indicated) Strength: 1% w/w Dosage form: Eye ointment Route of administration: Ophthalmic | Pack size: 5g Primary pack: Aluminum tube with epoxy-phenolic-ureic resin internal coating. Polyethylene screw cap and nozzle. Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
94 | Parenterals | Tranexamic acid injection 100mg/ml 10ml ampoule, pack of 1, 5 or 10 ampoules | Tranexamic acid, 10ml ampoules Strength: 100mg/ml Dosage form: Solution for injection Route of administration: Slow Intravenous Injection | Primary pack: Type I clear glass 10ml ampoules Secondary pack: 1 x 10ml; 5 x 10ml; 10 x 10ml. Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
95 | Parenterals | Tranexamic acid injection 100mg/ml 5ml ampoule, pack of 1, 5 or 10 ampoules | Tranexamic acid, 5 ml ampoules Strength: 100mg/ml Dosage form: Solution for injection Route of administration: Slow Intravenous Injection | Primary pack: Type I clear glass 5ml,ampoules Secondary pack: 1 x 5ml; 5 x 5ml; 10 x 5ml. Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
96 | Parenterals | Vitamin K injection 2mg/0.2ml in 0.2 ml vial, pack of 5 vials | Vitamin K Strength: 2mg/0.2ml in 0.2ml vial Dosage form: Solution for injection Route of administration: IV/ IM | Primary pack: Amber glass vial Secondary pack: 5 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
97 | Parenterals | Vitamin K-1 (Phytomenadione) 1mg/ml, 1ml vial, pack of 10 or 100 vials | Vitamin K-1 (Phytomenadione) Strength: 1mg/ml in 1ml vial Dosage form: Solution for injection Route of administration: IV | Primary pack: Amber glass vial Secondary pack: 10 or 100 vials with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
98 | Parenterals | Water for injection in 10ml, pack of 20 or 50 plastic ampoules | Water for injection in 10ml plastic ampoule Strength: Each 1ml contains 1m water for injection Dosage form: Solvent/solution for parenteral use Route of administration: parenteral | Primary pack: polypropylene plastic ampoules Secondary pack: 20 or 50 ampoules with package insert Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. Do not store above 30°C | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
100 | Orals | Zidovudine 300MG + 150mg Lamivudine, pack of 60 tablets | Zidovudine + Lamivudine Strength: 300mg Zidovudine + 150mg Lamivudine Dosage form: tablets Route of administration: oral BCS Class: Class 1/1 BCS based biowaiver can be accepted in place of bioequivalence (BE) | Pack size: 60s Primary pack: HDPE bottles Secondary pack: Unit carton Label & Carton text: Trilingual Package Insert: Trilingual | To accommodate Zove IVb climatic conditions. 30 °C storage conditions. Light protection for light sensitive products | Controlled temperature throughout transit Data logger to be included with shipment Passive cooled temperature for airfreight Active container for sea freight
Secondary sources (please add lines as needed if you propose secondary sources)

## 3. Technical Info Medical Dev
Tab 3. Technical Information Form for the Medical Devices
| Name of Company (Please enter in the cell on the right): | to be filled-in
Other comments (if any):
Important guidance to Bidders:
1. Please fill in the blanks under green columns.
2. Technical information Form should be submitted in excel form.
3. Complete every single row following Technical Requirements for Medical Devices.
4. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font (red).
5. To enable product identification for medical devices, it is essential to indicate and accurate Manufacturer`s Product Code in column 1.2.
6. Please list the secondary sources, where applicable, in the end under sub-section "secondary source".
7. Bidders are not allowed to copy and paste UNFPA generic specifications or to use general statements like: “offer as per specifications”, “comply”, “offered as requested”, “as per requirement”, “yes, but different pack size instead - pack of 1000“ or similar general statements/information. Similar statements might invalidate the bid. Deviations between this sheet and product literature (e.g, manufacturer brochures/webpage) may invalidate the bid.
MEDICAL DEVICES PRODUCT LIST |  |  |  | TECHNICAL INFORMATION
Bid Item No. | Product Sub- category | UNFPA Item Description | Manufacturer's name | Manufacturer's full address | Address of the actual manufacturing site | Manufacturer's website | Country of Origin | Product Shelf life (months) | Intended Application | Medical Device Class - is it correctly identified based on EU MDR (YES/NO) | 1.1 Manufacturer's Product Name or Brand name or description If multiple, describe all | 1.2 Manufacturer's Product code, Reference number, Model number, Serial number If multiple, describe all | 1.3 Date of last inspection / audit by a Notified Body or a regulatory agency | 1.4 Identification of Subcontractor (name and address) (IF e.g. final assembly, final sterilization, sub part manufacturing is done by other party than 1.1) | 2.0 EC Class Describe | 2.1 Purpose, Dimensions, Size, Diameter, Volume, Type, Shape, Model meet the UNFPA requirements. (YES/NO) | 2.1.1 Clarification if Failed Compliance in 2.1 Describe | 2.1.2 Weight of the item (grams) Describe | 2.2 Material, Raw Material, Material Features meet the UNFPA requirements. (YES/NO) | 2.2.1 Clarification if Failed Compliance in 2.2 Describe | 2.2.2 Item is/has stainless steel? Describe the AISI type of stainless steel | 2.3 Identify Documents in the Dossier: Product Brochure, Leaflet, Technical Data Sheet and Instruction Manual is provided (YES/NO) | 2.4 Are the Technical Specifications from the Supplier 'exact replicate or identical' to UNFPA Technical Specificatons (YES/NO) | 2.5 Photos of product are provided (YES/NO) | 3.1 Is Item prequalified by WHO or approved by other UN procurement agencies for long term contract (e.g. WHO Procurement, UNICEF) Describe agency and provide details like year of approval | 3.2.1 Valid EC certificate (= CE mark)? YES/NO | 3.2.2 Describe expiry date of the EC certificate MM/YYYY | 3.2.3 The name and country of EC certifying authority (e.g. notifying body) Describe | 3.2.4 EC Representative's name and country. Describe | 3.3 Valid FDA approval and 510k (or PMA) clearance? YES/NO | 3.4 ISO 17050 Declaration of Conformity Signed and dated and with correct scope? YES/NO | 3.5.1 ISO 13485 - QMS, MDD 93/42/EEC, MDR 2017/745 Valid - Yes/No Valid? YES/NO | 3.5.2 Describe expiry date of the QMS 13485 certificate MM/YYYY | 3.5.3 The name and country of QMS 13485 certifying authority (e.g. notifying body) Describe | 3.6.1 ISO 9001 - QMS Valid? YES/NO | 3.6.2 Describe expiry date of the QMS 9001 certificate MM/YYYY | 3.6.3 The name and country of QMS 9001 certifying authority (e.g. notifying body) Describe | 3.7 Free Sale Certificate, Certificate for Export or similar YES/NO | 3.8 Compliance to Other Item-specific Safety Standards Describe Standard | 3.9 ISO 14001 - Environment Valid YES/NO | 3.10 ISO 50001 - Energy Management Valid YES/NO | 3.11 List of all other documentation provided by the Supplier or Manufacturer Describe Name the document | 3.12 Manufacturing Licence (1.1) Valid YES/NO | 3.13 If applicaple - Safety documents of Subcontractor (1.5) (e.g. ISO 13485) Describe | 3.14 Test Reports - List down the test reports submitted | 3.15 Test Reports - Sample name | 3.16 Test Report - Verified authenticity of the report (YES/NO) | 3.17 Test Report - is the testing facility ISO 17025 accredited to perform the testing? (YES/NO) | 4.1 Reusable, re-sterilized Item? YES/NO | 4.1.1 Reusable, re-sterilized Item - Instructions are provided for efficient cleaning, desinfecting and sterilization, and certificate to prove that material endures sterilization process. YES/NO/not applicable | 4.2 Sterile or single use Item? YES/NO | 4.2.1 Sterile or single use Item - Method of sterilization Describe | 4.2.2 Sterile or single use Item - ISO 13485 for specific sterilization process: ISO 11135 (ETO) ISO 11137 (gamma radiat) ISO 17665 (steam) ISO 20857 (dry heat) ISO 14937 (other) YES/NO/not applicable | 4.2.3 Sterile or single use Item - Three (3) latest, individual batch sterilization certificates are submitted YES/NO/not applicable | 4.3.1 Packaging requirements are efficient enough to protect the product? YES/NO | 4.3.2 Is the cardboard carton FSC marked? YES/NO | 4.4 Photos of primary and secondary LABELS are submitted YES/NO | 4.5 Photos of primary and secondary packaging are submitted YES/NO | 4.6 Storage conditions Describe | 4.7 Installation manual, spares provided and service/repair information correspond to given UNFPA specifications Describe | 4.8 Training and Support Needed for device operation YES/NO/not applicable | 4.9 Warranty - sufficient warranty and copy of it is provided YES/NO/not applicable | 4.10 Electrical Equipment - Large variety of voltages, frequencies and plugs quaranteed for various customers globally. YES/NO/not applicable | 4.11 Information of Safe Disposal was provided and approved YES/NO | 4.12 Instructions for use are provided in all 3 languages (EN, FR, SP). YES/NO/not applicable
101 | Anaesthesia & Resuscitation Equipment | Anaesthesia machine
102 | Anaesthesia & Resuscitation Equipment | Aspirator, surgical, suction pump, w/accessories
103 | Anaesthesia & Resuscitation Equipment | Incubator, Infant, automatic
104 | Anaesthesia & Resuscitation Equipment | Infusion pump, with accessories
105 | Anaesthesia & Resuscitation Equipment | Pulse oximeter, portable and accessories
106 | Anaesthesia & Resuscitation Equipment | Oxygen concentrator, set
107 | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operating, child/newborn
108 | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operating, neonate, set
109 | Anaesthesia & Resuscitation Equipment | Resuscitator, hand-operated, adult, set
110 | Anaesthesia & Resuscitation Equipment | Table, resuscitation, newborn, w/accessories
111 | Anaesthesia & Resuscitation Equipment | Pump, suction, electric, aspirator, portable, w/accessories
112 | Anaesthesia & Resuscitation Equipment | Pump, suction, foot-operated, aspirator
113 | Anaesthesia & Resuscitation Equipment | Syringe pump
114 | Medical Electrical Equipment | Cryosurgical Unit
115 | Medical Electrical Equipment | Electrosurgical unit, w/accessories
116 | Medical Electrical Equipment | Light, examination, mobile
117 | Medical Electrical Equipment | Torch, flashlight, pre-focused, angled, medical
118 | Medical Electrical Equipment | Light, operating, mobile
119 | Medical Electrical Equipment | Monitor, patient, portable,w/access
120 | Medical Electrical Equipment | Refrigerator, compression
121 | Medical Electrical Equipment | Warmer, baby, electric
122 | Medical Sterilization Equipment | Indicator, TST control, spot, adhesive, P1
123 | Medical Sterilization Equipment | Indicator, TST control, spot, adhesive, P1050
124 | Medical Sterilization Equipment | Basket, sterilizing, approx. 120 x 250 x 60 mm
125 | Medical Sterilization Equipment | Basket, instruments, for sterilization, wired, 400 x 200 x 90 mm
126 | Medical Sterilization Equipment | Drum, sterilizing, diameter 165mm
127 | Medical Sterilization Equipment | Drum, sterilizing, diameter 260mm
128 | Medical Sterilization Equipment | Drum, sterilizing, diameter 290mm
129 | Medical Sterilization Equipment | Drum, sterilizing, diameter 340mm
130 | Medical Sterilization Equipment | Sterilizer, steam, pressure-type, 24 litres
131 | Medical Sterilization Equipment | Sterilizer, steam, pressure-type, 39 litres
132 | Medical Sterilization Equipment | Stove, kerosene, single-burner, pressure-type
133 | Medical Sterilization Equipment | Timer, 60 min, mechanical
134 | PPEs | Apron, surgical, disposable, water resistant, non-sterile
135 | PPEs | Gloves, exam, nitrile, powder free, non-sterile, large, P100
136 | PPEs | Gloves, exam, nitrile, powder free, non-sterile, medium,P100
137 | PPEs | Gown, surgical, non-woven, sterile, disposable
138 | PPEs | Medical Mask, type I, for patients, disposable
139 | PPEs | Respirator, FFP2/N95/KN95, mask, no valve, disposable
140 | PPEs | Surgical Mask, type IIR, healthcare worker, disposable
141 | PPEs | Goggles, panoramic, regular nose, indirect ventilation
142 | PPEs | Coverall, L, 3b/4b/5b/6b, disposable Coverall, M, 3b/4b/5b/6b, disposable Coverall, XL, 3b/4b/5b/6b, disposable
143 | PPEs | Faceshield, fog-resistant, disposable
144 | PPEs | Cap,surgical,bouffant,non-woven
145 | Medical supplies | Airway, Guedel, paediatric, size 0 (50mm)
146 | Medical supplies | Airway, Guedel, neonate, size 00 (35-40mm)
147 | Medical supplies | Airway, Guedel, paediatric, size 1 (60mm)
148 | Medical supplies | Airway, Guedel, child, size 2 (70mm)
149 | Medical supplies | Airway, Guedel, adolescent, size 3 (80mm)
150 | Medical supplies | Airway, Guedel, adult, size 4 (90mm)
151 | Medical supplies | Airway, Guedel, adult, size 5 (100mm)
152 | Medical supplies | Amniohook, plastic, sterile, single use, P1
153 | Medical supplies | Amniohook, plastic, sterile, single use, P100
154 | Medical supplies | Non-pneumatic anti-shock garment
155 | Medical supplies | Bag, Biohazard, yellow, 50 liters, P50
156 | Medical supplies | Blood-giving set
157 | Medical supplies | Safety box, disposal of used syringes & needles, 5L, P1
158 | Medical supplies | Safety box, disposal of used syringes & needles, 5L, P25
159 | Medical supplies | Cannula, IV short, 16G, sterile, single use
160 | Medical supplies | Cannula, IV short, 18G, sterile, single use
161 | Medical supplies | Cannula, IV short, 20G, sterile, single use, P1
162 | Medical supplies | Cannula, IV short, 20G, sterile, single use, P50
163 | Medical supplies | Cannula, IV short, 22G, sterile, single use
164 | Medical supplies | Cannula, IV short, 24G, sterile, single use
165 | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P1
166 | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P10
167 | Medical supplies | Catheter, urethral, Foley, CH12, sterile, P50
168 | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P10
169 | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P50
170 | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P10
171 | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P50
172 | Medical supplies | Catheter, urethral, Foley, CH14, sterile, P1
173 | Medical supplies | Catheter, urethral, Foley, adult, CH18, sterile, P1
174 | Medical supplies | Catheter, urethral, CH12, sterile, P1
175 | Medical supplies | Catheter, urethral, CH12, sterile, P10
176 | Medical supplies | Catheter, ureteral, CH12, sterile, P100
177 | Medical supplies | Catheter, ureteral, Nelaton, CH12, sterile, P50
178 | Medical supplies | Catheter, urethral, CH5, with metal guide wire, sterile
179 | Medical supplies | Catheter, urethral, size CH6, with metal guide wire, sterile
180 | Medical supplies | Catheter, urethral, Foley, CH16, sterile, P1
181 | Medical supplies | Catheter, urethral, Foley, CH20, sterile, P1
182 | Medical supplies | Colostomy bag, transparent, drainable, w/filter, P10
183 | Medical supplies | Colostomy bag, transparent, drainable, w/filter, P20
184 | Medical supplies | Gauze, compress, 10 x 10cm, non-sterile, P100
185 | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P50
186 | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P20
187 | Medical supplies | Gauze, compress, 10 x 10cm, sterile, P5
188 | Medical supplies | Cotton wool, 500g, roll, non-sterile
189 | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml
190 | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml
191 | Medical supplies | Bloodbag+CPDA-1(citrate-phosphate-dextrose-adenine),250ml,P5
192 | Medical supplies | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),350ml,P1
193 | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 350ml
194 | Medical supplies | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),450ml,P1
195 | Medical supplies | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 450ml
196 | Medical supplies | Bag (envelope), plastic, for drugs, 10 x 15cm
197 | Medical supplies | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P1
198 | Medical supplies | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P100
199 | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P1
200 | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P10
201 | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P100
202 | Medical supplies | Extractor, mucus, 20ml, sterile, single use, P50
203 | Medical supplies | Gloves, exam, latex, large, non-sterile, single use, P100
204 | Medical supplies | Gloves, examination, latex, medium, non-sterile
205 | Medical supplies | Gloves, examination, latex, medium, non-sterile, P100
206 | Medical supplies | Gloves, exam, latex, small, non-sterile, single use, P100
207 | Medical supplies | Gloves, gynaecological, large, powder-free, sterile, 1 pair
208 | Medical supplies | Gloves, gynaecological, large, powder-free, sterile
209 | Medical supplies | Gloves, gynaecological, medium, powder-free, sterile, 1 pair
210 | Medical supplies | Gloves, gynaecological, medium, powder-free, sterile
211 | Medical supplies | Gloves, gynaecological, small, powder-free, sterile, 1 pair
212 | Medical supplies | Gloves, gynaecological, small,powder-free, sterile, P25
213 | Medical supplies | Gloves, gynaecological, small,powder-free, sterile, P50
214 | Medical supplies | Gloves, surgical, size 7, powder-free, sterile, single use, P1
215 | Medical supplies | Gloves, surgical, size 7, powder-free, sterile, single use, P50
216 | Medical supplies | Gloves, surgical, size 8, powder-free, sterile, single use, P50
217 | Medical supplies | Gloves, surgical, size 8, powder-free, sterile, single use, P1
218 | Medical supplies | Gloves, surgical, size 7,5, powder-free, sterile, single use
219 | Medical supplies | Gloves, surgical, size 8.5, powder-free, sterile, single use
220 | Medical supplies | Brush, hand, scrubbing, plastic, P1
221 | Medical supplies | Infusion set, sterile, single use, 1 piece
222 | Medical supplies | Infusion set, sterile, single use, P100
223 | Medical supplies | Infusion set, sterile, single use, P400
224 | Medical supplies | Needle, spinal, 25G, sterile, single use
225 | Medical supplies | Needle, luer, 19G, sterile, P100
226 | Medical supplies | Needle, luer, 21G, sterile, P100
227 | Medical supplies | Needle, luer, 23G, sterile, P100
228 | Medical supplies | Needle, scalp vein, 21G, sterile, P100
229 | Medical supplies | Needle, scalp vein, 21G, sterile, single use, P50
230 | Medical supplies | Needle, scalp vein, 21G, sterile, single use
231 | Medical supplies | Needle, spinal, 22G, sterile, single use, P100
232 | Medical supplies | Needle, spinal, 22G, sterile, single use, P1
233 | Medical supplies | Needle, spinal, 22G, sterile, single use, P25
234 | Medical supplies | Needle, spinal, 22G, sterile, single use, P50
235 | Medical supplies | Surgical needle, 1/2 circle, round, spring eye, size14, sterile
236 | Medical supplies | Needle, surgical, 1/2, 27mm, round, spring eye, P24
237 | Medical supplies | Needle, surgical, 1/2, 63mm, round, spring eye, P24
239 | Medical supplies | Adhesive bandage, wound plaster, waterproof
240 | Medical supplies | Pressure cuff, infusion/transfusion, for bags 500-1000 ml
241 | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P10
242 | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P12
243 | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P15
244 | Medical supplies | Gauze, 8cm x 4m, roll, non-sterile, P5
245 | Medical supplies | Tube, suction, Yankauer, 28cm
246 | Medical supplies | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile
247 | Medical supplies | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile, P36
248 | Medical supplies | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P12
249 | Medical supplies | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P36
250 | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P12
251 | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P36
252 | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P12
253 | Medical supplies | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P36
254 | Medical supplies | Suture, absorbable, DEC 3 (2-0), spool, sterile, P12
255 | Medical supplies | Suture, absorbable, DEC 3 (2-0), spool, sterile, P36
256 | Medical supplies | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P12
257 | Medical supplies | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P36
258 | Medical supplies | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P12
259 | Medical supplies | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P36
260 | Medical supplies | Suture, absorbable, DEC 4 (1), 3/8, 50mm, round, sterile, P36
261 | Medical supplies | Suture, absorbable, DEC 1.5 (4/0), with needle, fistula repair
262 | Medical supplies | Suture, absorbable, DEC 2 (3/0), with needle, fistula repair
263 | Medical supplies | Suture, absorbable, DEC 3 (2/0), with needle, fistula repair
264 | Medical supplies | Suture, absorbable, DEC 3.5 (0), with needle, fistula repair
265 | Medical supplies | Suture, absorbable, DEC 4 (1), with needle, fistula repair
266 | Medical supplies | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile,
267 | Medical supplies | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile,
268 | Medical supplies | Suture, non-absorbable PA, DEC 3 (2/0), 1/2 circle, round
269 | Medical supplies | Suture, non-absorbable PA, DEC 4 (1), 1/2 circle, round
270 | Medical supplies | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile
271 | Medical supplies | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile
272 | Medical supplies | Swab, alcohol, sterile, single use
273 | Medical supplies | Syringe, 0.5ml, w/permanently att. needle,sterile,single use
274 | Medical supplies | Auto Disabled Syringe, 0.5ml, w/permanently att. needle,sterile,single use
275 | Medical supplies | Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100
276 | Medical supplies | Auto Disabled Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100
277 | Medical supplies | Syringe, luer, 2 ml, w/21G needle, sterile, single use, P100
278 | Medical supplies | Auto Disabled Syringe, luer, 2ml, w/21G needle, sterile, single use,P100
279 | Medical supplies | Syringe, feeding, catheter tip, 50 ml, sterile, P1
280 | Medical supplies | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P1
281 | Medical supplies | Syringe, feeding, catheter tip, 50 ml, sterile, P50
282 | Medical supplies | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P50
283 | Medical supplies | Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100
284 | Medical supplies | Auto Disabled Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100
285 | Medical supplies | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P100
286 | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P100
287 | Medical supplies | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P1
288 | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P1
289 | Medical supplies | Syringe, luer, 5ml, w/23G needle, sterile, single use, P25
290 | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P25
291 | Medical supplies | Syringe, luer lock, 20 ml, sterile, single use, P1
292 | Medical supplies | Auto Disabled Syringe, luer lock, 20 ml, sterile, single use, P1
293 | Medical supplies | Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100
294 | Medical supplies | Auto Disabled Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100
295 | Medical supplies | Bladder syringe 60-100 mls (with long nozzle, not luer lock)
296 | Medical supplies | Auto Disabled Bladder syringe 60-100 mls (with long nozzle, not luer lock)
297 | Medical supplies | Syringe, luer, 10 ml, w/o needle, sterile, single use, P100
298 | Medical supplies | Auto Disabled Syringe, luer, 10 ml, w/o needle, sterile, single use, P100
299 | Medical supplies | Syringe, luer, 1ml, w/o needle, sterile, single use, P100
300 | Medical supplies | Auto Disabled Syringe, luer, 1ml, w/o needle, sterile, single use, P100
301 | Medical supplies | Syringe, luer, 2ml, sterile, single use, P100
302 | Medical supplies | Auto Disabled Syringe, luer, 2ml, sterile, single use, P100
303 | Medical supplies | Syringe, luer, 5ml, sterile, single use, P100
304 | Medical supplies | Auto Disabled Syringe, luer, 5ml, w/o needle, sterile, single use, P100
305 | Medical supplies | Tape, adhesive, zinc oxide, 2.5cm x 5m, P1
306 | Medical supplies | Tape, adhesive, zinc oxide, 2.5cm x 5m, P10
307 | Medical supplies | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P1
308 | Medical supplies | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P10
309 | Medical supplies | Tape, umbilical, 3mmx50m, non-sterile, single use, P1
310 | Medical supplies | Tape, umbilical, 3mm x 15cm, non-sterile, P1
311 | Medical supplies | Tourniquet, latex rubber, 75cm
312 | Medical supplies | Tube, suction, CH10, 50 cm long, conical tip, sterile, P1
313 | Medical supplies | Tube, suction, CH14, 50 cm long, conical tip, sterile, P1
314 | Medical supplies | Bag, urine, collecting, 2 litres, P10
315 | Medical supplies | Bag, urine, collecting, 2 litres, tap for emptying
316 | Medical supplies | Sound, uterine, Martin, 32cm
317 | Medical utensils | Basin, kidney, stainless steel, 1400ml
318 | Medical utensils | Basin, kidney, polypropylene, 475ml
319 | Medical utensils | Basin, kidney, stainless steel, 825ml
320 | Medical utensils | Bowl, stainless steel, 180ml
321 | Medical utensils | Bowl, stainless steel, 4L
322 | Medical utensils | Bowl, stainless steel, 500 ml
323 | Medical utensils | Bowl, polypropylene, 6L
324 | Medical utensils | Jar, forceps, 180 mm, no cover
325 | Medical utensils | Jar, thermometer, plastic, 11cm, w/o cover
326 | Medical utensils | Bowl, sponge, stainless steel, 1.2 L
327 | Medical utensils | Ruler, metal, in centimeters, 30cm
328 | Medical utensils | Tray, dressing, stainless steel, 30 x 20 x 3cm
329 | Medical utensils | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm, with cover
330 | Medical utensils | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover
331 | Medical Attire & Linen | Apron, protection, plastic, reusable
332 | Medical Attire & Linen | Apron, surgical, rubber, washable
333 | Medical Attire & Linen | Sheet, bed, 289 x 178 cm, washable
334 | Medical Attire & Linen | Blanket, survival, 220 x 140cm
335 | Medical Attire & Linen | Cotton cloth, “tetra”, 100 x 100cm
336 | Medical Attire & Linen | Drape, surgical, woven, 100x100 cm, reusable
337 | Medical Attire & Linen | Drape, surgical woven, 100 x 150cm
338 | Medical Attire & Linen | Drape, surgical, woven, 90x150 cm, reusable
339 | Medical Attire & Linen | Draw sheet, waterproof, 80 x 120cm, single use
340 | Medical Attire & Linen | Draw sheet, plastic, 90 x 180cm, reusable
341 | Medical Attire & Linen | Draw sheet, plastic, 100 x 100 cm
342 | Medical Attire & Linen | Draw sheet, rubber, 90x150 cm, reusable
343 | Medical Attire & Linen | Glasses, safety, regular size, disposable
345 | Med. Diagnostic Equip&Supplies | Dopplerfetalheart
346 | Med. Diagnostic Equip&Supplies | Haemoglobin photometer handheld w/accessories
347 | Med. Diagnostic Equip&Supplies | Height measuring instrument (0-2.2m)
348 | Med. Diagnostic Equip&Supplies | Capillary tubes for HemoCue Hb 301, P200
349 | Med. Diagnostic Equip&Supplies | Pelvimeter, Collyer type, external
350 | Med. Diagnostic Equip&Supplies | Pregnancy wheel
351 | Med. Diagnostic Equip&Supplies | Scale, infant, beam type, 16kg x 5g
352 | Med. Diagnostic Equip&Supplies | Sling, for use with infant scale
353 | Med. Diagnostic Equip&Supplies | Scale, mother/child, electronic, 200kg
354 | Med. Diagnostic Equip&Supplies | Scale, physician, adult beam balance, 0-200kg
355 | Med. Diagnostic Equip&Supplies | Scale, infant, spring-type, 5kg x 25g
356 | Med. Diagnostic Equip&Supplies | Sphygmomanometer, aneroid, adult
357 | Med. Diagnostic Equip&Supplies | Stethoscope, binaural, complete
358 | Med. Diagnostic Equip&Supplies | Stethoscope, obstetrical, Pinard, monoaural
359 | Med. Diagnostic Equip&Supplies | Thermometer, clinical, digital 32°C-43°C
360 | Med. Diagnostic Equip&Supplies | Timer, for respiration, ARI
361 | Med. Diagnostic Equip&Supplies | Ultrasound, scanner, mobile w/ printer and accessories
362 | Medical & Surgical Instruments | Spatula, Ribbon retractor, malleable, 27 mm x 250 mm
363 | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no.10
364 | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 11, P1
365 | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 11, P100
366 | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 15, P1
367 | Medical & Surgical Instruments | Blade, scalpel, sterile, single use, no. 22, P100
368 | Medical & Surgical Instruments | Forceps, intestinal clamp, Doyen, 23cm, curved
369 | Medical & Surgical Instruments | Forceps, intestinal, clamp, Doyen, 23cm, straight
370 | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P100
371 | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P1
372 | Medical & Surgical Instruments | Clamp, umbilical, 5.2cm, sterile, single use, P50
373 | Medical & Surgical Instruments | Cranioclast, Braun, 420 mm
374 | Medical & Surgical Instruments | Curette, uterine, Sims, 12 mm, sharp
375 | Medical & Surgical Instruments | Curette, uterine, Sims, 7 mm, sharp
376 | Medical & Surgical Instruments | Curette, uterine, Sims, 8 mm, blunt
377 | Medical & Surgical Instruments | Curette, uterine, Sims, 9 mm, sharp
378 | Medical & Surgical Instruments | Hook, decapitation, Braun, 31cm
379 | Medical & Surgical Instruments | Dilator, uterine, Hegar, set of 16 dilators, size 3-18
380 | Medical & Surgical Instruments | Forceps, dressing, Cheron, 25cm
381 | Medical & Surgical Instruments | Forceps, dressing, standard, 145 mm, straight
382 | Medical & Surgical Instruments | Forceps, uterine, Duplay, 28cm, curved
383 | Medical & Surgical Instruments | Forceps, catheter introducing, Magill, child
384 | Medical & Surgical Instruments | Forceps, obstetric, Simpson-Braun, 35 cm
385 | Medical & Surgical Instruments | Forceps, uterine, Phaneuf, 21.5 cm, curved
386 | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 20 cm, curved
387 | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 22cm, straight
388 | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 24 cm, curved
389 | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 16 cm, curved
390 | Medical & Surgical Instruments | Forceps, artery, Rochester-Pean, 16 cm, straight
391 | Medical & Surgical Instruments | Forceps, artery, Bengolea, 20cm, curved
392 | Medical & Surgical Instruments | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 16mm, S
393 | Medical & Surgical Instruments | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 19mm, L
394 | Medical & Surgical Instruments | Forceps, sterilizer, Cheatle, 27 cm, curved
395 | Medical & Surgical Instruments | Forceps, artery, Crafoord, 24cm, curved
396 | Medical & Surgical Instruments | Forceps, dressing, standard, fine serrated jaws, 20cm
397 | Medical & Surgical Instruments | Forceps, dressing, standard, 250 mm, straight
398 | Medical & Surgical Instruments | Forceps, artery, Kelly, 14cm, curved
399 | Medical & Surgical Instruments | Forceps, artery, Kocher, 140 mm, 1x2 teeth, curved
400 | Medical & Surgical Instruments | Forceps, artery, Kocher, 14cm, straight
401 | Medical & Surgical Instruments | Forceps, artery, Mixter, 14cm
402 | Medical & Surgical Instruments | Forceps, artery, Mixter, 23cm
403 | Medical & Surgical Instruments | Forceps, artery, Halsted-Mosquito, 12.5cm, curved
404 | Medical & Surgical Instruments | Forceps, uterine, Museux, 24cm, curved
405 | Medical & Surgical Instruments | Forceps, artery, Spencer-Wells, 20.5 cm, curved
406 | Medical & Surgical Instruments | Forceps, peritoneal, Faure, 21cm, slightly curved
407 | Medical & Surgical Instruments | Forceps, sponge-holding, Foerster, 20 cm, straight
408 | Medical & Surgical Instruments | Forceps, sponge-holding, Foerster, 24 cm, straight
409 | Medical & Surgical Instruments | Forceps, tenaculum, vulsellum, Schröder, straight, 25cm
410 | Medical & Surgical Instruments | Forceps, tissue, Allis, 3x4 teeth, 20 cm
411 | Medical & Surgical Instruments | Forceps, tissue, standard, 145mm, straight
412 | Medical & Surgical Instruments | Forceps, tissue, Allis, 4x5 teeth, 20 cm
413 | Medical & Surgical Instruments | Forceps, tissue, Babcock, 20 cm
414 | Medical & Surgical Instruments | Forceps,tissue,standard,dissecting,1x2 teeth,20cm, straight
415 | Medical & Surgical Instruments | Forceps, tissue, standard, 250 mm, straight
416 | Medical & Surgical Instruments | Forceps, Schroeder, Vulsellum, 25 cm, 2x2 prongs, curved
417 | Medical & Surgical Instruments | Forceps, tissue, Allis, 4x5 teeth, 15 cm
418 | Medical & Surgical Instruments | Forceps, tissue, Allis, 3x4 teeth, 15 cm
419 | Medical & Surgical Instruments | Forceps, tissue, Colin, 16 cm
420 | Medical & Surgical Instruments | Forceps, tissue, Duval, 23cm
421 | Medical & Surgical Instruments | Hook, fistula, Emmet, 22 cm
422 | Medical & Surgical Instruments | Accesory kit for MVA, including (1) O-ring, (1) collar stop, (1), cap, (1) 2cc silicone
423 | Medical & Surgical Instruments | Cannulae, Easygrip, 10mm, integrated base, SR10
424 | Medical & Surgical Instruments | Cannulae, Easygrip, 12mm, integrated base, SR12
425 | Medical & Surgical Instruments | Cannulae, Easygrip, 4mm, integrated base, SR4
426 | Medical & Surgical Instruments | Cannulae, Easygrip, 5mm, integrated base, SR5
427 | Medical & Surgical Instruments | Cannulae, Easygrip, 6mm, integrated base, SR6
428 | Medical & Surgical Instruments | Cannulae, Easygrip, 7mm, integrated base, SR7
429 | Medical & Surgical Instruments | Cannulae, Easygrip, 8mm, integrated base, SR8
430 | Medical & Surgical Instruments | Cannulae, Easygrip, 9mm, integrated base, SR9
431 | Medical & Surgical Instruments | Dilators, Denniston, polymer, P5
432 | Medical & Surgical Instruments | MVA Ipas kitting service
433 | Medical & Surgical Instruments | MVA with 2cc silicone
434 | Medical & Surgical Instruments | Silicone, 2cc, P10
435 | Medical & Surgical Instruments | Needle, aneurysm, Deschamps, 20cm, sharp, left hand
436 | Medical & Surgical Instruments | Needle, aneurysm, Deschamps, 20cm, sharp, right hand
437 | Medical & Surgical Instruments | Needle holder, Derf, 11.5 cm, straight
438 | Medical & Surgical Instruments | Needle holder, Baby-Crile-Wood, 15cm, delicate
439 | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 18cm, curved
440 | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 18cm, straight
441 | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 20cm, curved
442 | Medical & Surgical Instruments | Needle holder, Mayo-Hegar, 20cm, straight
443 | Medical & Surgical Instruments | Needle, luer, 25G, sterile, single use, P100
444 | Medical & Surgical Instruments | Needle, scalp vein, butterfly, 25G, sterile, single use,P100
445 | Medical & Surgical Instruments | Perforator, Smellie, 25cm
446 | Medical & Surgical Instruments | Probe, surgical, double-ended, 14.5 cm
447 | Medical & Surgical Instruments | Probe, surgical, with eye, malleable, 20 cm
448 | Medical & Surgical Instruments | Retractor, double-ended, Farabeuf, 15 cm, pair
449 | Medical & Surgical Instruments | Retractor, Langenbeck, single-ended, medium, 21cm
450 | Medical & Surgical Instruments | Retractor, abdominal, Balfour, 3 blades
451 | Medical & Surgical Instruments | Retractor, abdominal, Collin, 3 blades
452 | Medical & Surgical Instruments | Retractor, vaginal, Doyen, 8.5 x 4.5cm
453 | Medical & Surgical Instruments | Retractor, abdominal, Gosset, 3 blades
454 | Medical & Surgical Instruments | Safety Pin, forceps holder, Mayo, 11.5cm
455 | Medical & Surgical Instruments | Safety Pin, forceps holder, Schaedel, 9cm
456 | Medical & Surgical Instruments | Scissors for umbilical cord cutting, reusable, sterile
457 | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 15 cm
458 | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 16 cm
459 | Medical & Surgical Instruments | Scalpel handle, no. 7, blade holder, 18 cm
460 | Medical & Surgical Instruments | Scalpel handle no. 3, blade holder, 13 cm
461 | Medical & Surgical Instruments | Scalpel handle no. 4, blade holder, 13 cm
462 | Medical & Surgical Instruments | Scalpel handle no. 4, blade holder, long 21cm
463 | Medical & Surgical Instruments | Scalpel, knife, blade no. 22, with plastic handle, sterile
464 | Medical & Surgical Instruments | Scalpel, knife, blade no. 22, w/ plastic handle, sterile,P10
465 | Medical & Surgical Instruments | Scissors, Deaver, 14cm, curved, sharp/blunt
466 | Medical & Surgical Instruments | Scissors, fistula, 20cm curved
467 | Medical & Surgical Instruments | Scissors, Episiotomy, Braun, 14.5, angular
468 | Medical & Surgical Instruments | Scissors, Bandage, Lister, 14 cm, angular, sharp/blunt
469 | Medical & Surgical Instruments | Scissors, Bandage, Lister, 18 cm, angular, sharp / blunt
470 | Medical & Surgical Instruments | Scissors, Boyd, 17cm, double curved
471 | Medical & Surgical Instruments | Scissors, Deaver, 14cm, straight, sharp/blunt
472 | Medical & Surgical Instruments | Scissors, gynaecological, 20cm, curved, blunt/blunt
473 | Medical & Surgical Instruments | Scissors, Mayo, 14cm, curved, blunt/blunt
474 | Medical & Surgical Instruments | Scissors, Mayo, 17cm, curved, blunt/blunt
475 | Medical & Surgical Instruments | Scissors, Mayo, 17 cm, straight, blunt/blunt
476 | Medical & Surgical Instruments | Scissors, Mayo, 23cm, curved, blunt/blunt
477 | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 14cm, curved, blunt/blunt
478 | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 18cm, curved, blunt/blunt
479 | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 23cm, curved, blunt/blunt
480 | Medical & Surgical Instruments | Scissors, Metzenbaum/Nelson, 24 cm, curved, blunt/blunt
481 | Medical & Surgical Instruments | Scissors, stitch, 15cm
482 | Medical & Surgical Instruments | Scissors, suture stich, 18cm, straight, sharp/blunt
483 | Medical & Surgical Instruments | Scissors, Thorek, 20cm, curved
484 | Medical & Surgical Instruments | Scissors, Thorek, 18cm, curved, sharp/blunt
485 | Medical & Surgical Instruments | Scissors, umbilical, USA-model, 10 cm
486 | Medical & Surgical Instruments | Scissors, uterine, Sims, curved, 20 cm, blunt/blunt
487 | Medical & Surgical Instruments | Scissors, Thorek, 18cm, curved 90 degrees
488 | Medical & Surgical Instruments | Scoop, uterine, Simon, 6mm, sharp
489 | Medical & Surgical Instruments | Speculum, vaginal, Auvard, retractor, weighted, 8 x 3.8cm
490 | Medical & Surgical Instruments | Speculum, vaginal, Auvard, retractor, weighted, 10 x 4 cm
491 | Medical & Surgical Instruments | Speculum, vaginal, Sims, retractor, double-ended, large
492 | Medical & Surgical Instruments | Speculum, vaginal, Sims, retractor, double-ended, medium
493 | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 11.5 x 3.5 cm
494 | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 9.5 x 3.5 cm
495 | Medical & Surgical Instruments | Speculum, vaginal, Cusco, 7.5 x 2 cm
496 | Medical & Surgical Instruments | Speculum, vaginal, Graves, 115 x 35mm
497 | Medical & Surgical Instruments | Speculum, vaginal, Graves, 95 x 35mm
498 | Medical & Surgical Instruments | Speculum, vaginal, Graves, 75 x 20mm
499 | Medical & Surgical Instruments | Clamp, towel, Backhaus, 120 mm
500 | Medical & Surgical Instruments | Clamp, towel, Backhaus, 90 mm
501 | Medical & Surgical Instruments | Sound, uterine, Martin, malleable, 30 cm
502 | Medical & Surgical Instruments | Vacuum extractor, Bird, anterior and posterior cups, manual, set
503 | Medical & Surgical Instruments | Vacuum extractor, Bird, manual, complete set
504 | Medical & Surgical Instruments | Vacuum extractor, Vacca OmniCup, hand-held, 1 person, reusable
505 | Medical & Surgical Instruments | Forceps, Dressing, Standard Pattern, 12.7 cm
506 | Medical & Surgical Instruments | Forceps, Intestinal, Allis, Delicate, 15 cm (5x6 teeth)
507 | Medical & Surgical Instruments | Retractor, double-ended, Farabeuf, 120 mm, pair
508 | Medical & Surgical Instruments | Army-Navy retractors, double-ended (for subumbilical procedure)
509 | Medical & Surgical Instruments | Forceps, Artery, Kelly, Straight, 14 cm
510 | Medical & Surgical Instruments | Forceps mosquito, delicate, curved, 14 cm
511 | Medical & Surgical Instruments | Scissors, Tonsil, Metzenbaum, Curved, 18 cm
512 | Medical & Surgical Instruments | Ramathobodi tubal hook
513 | Medical & Surgical Instruments | Ramathobodi uterine elevator, 28 cm
514 | Medical & Surgical Instruments | Forceps, uterine, tenaculum, Duplay, 248 mm, curved
515 | Medical & Surgical Instruments | Forceps, sponge, Foerster, straight, 24 cm
516 | Medical & Surgical Instruments | 1 Jackson vaginal retractor, 38mm X 76mm (deep blade) (for alternative method for inserting the uterine elevator)
517 | Medical Electrical Equipment | ECG Recorder,portable,w/access
518 | Med. Diagnostics Equipment & Supplies | Sphygmomanaometer, Aneroid, Adult
519 | Anaesthesia & Resus. Equip | Anaesthesia machine Portable
520 | Medical Electrical Equipment | Electrocardiograph with 12 channels
521 | Medical Electrical Equipment | CTG (Cardiotocography) instrument
522 | Medical Electrical Equipment | Automated external defibrillator system
523 | Med. Diagnostics & Supplies | Handheld breast ultrasound
524 | Medical Electrical Equipment | Ultrasound, Scanner, Digital w/access
525 | Med. Diagnostics & Supplies | Digital blood pressure monitor
526 | Medical Sterilization Equipmnt | Autoclave
527 | Medical Electrical Diagnostic Equipment | Colposcope
Secondary sources (please add lines as needed if you propose secondary sources)

## Specifications Pharmaceuticals (hidden)
Important guidance to Bidders: Kindly pay attention to complete the below information as per the specifications confirmed by the manufacturer and yourself. 1. Complete every single row following the model of UNFPA technical specifications. 2. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font. Notes: - If there is a deviation between this information and the supplied device, this will create a PSI discrepancy and initiate a potential 'Change of Specification' process (part of supplier's performance evaluation). - Below information will go unedited into the UNFPA online catalog if the LTA is signed. If you have comments for UNFPA, please write them in column E. Do not type phrases such as 'confirmed' or 'as per catalog' in this column D.
Bid Item No. | UNFPA Item Description | UNFPA Detailed Technical Specifications | Bidders Technical Specification of the Offered Product | Bidders Comments

## 4. Technical InfoIVDs
Tab 4. Technical Information Form for In Vitro Diagnostic Tests
| Name of Company (Please enter in the cell on the right): | to be filled-in
Other comments (if any):
Important guidance to Bidders:
1. Please fill in the blanks under orange columns
2. Technical information Form should be submitted in excel form
3. Complete every single row following Technical Requirements for IVDs
4. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font (red).
5. Bidders are not allowed to copy and paste UNFPA generic specifications or to use general statements like: “offer as per specifications”, “comply”, “offered as requested”, “as per requirement”, “yes, but different pack size instead - pack of 1000“ or similar general statements/information. Similar statements might invalidate the bid. Deviations between this sheet and product literature (e.g, manufacturer brochures/webpage) may invalidate the bid.
IN VITRO DIAGNOSTICS PRODUCT LIST |  |  | Technical Information
|  |  | Part I. Submitter and Manufacturer Information | [7 empty] | Part II. Product Information | [9 empty] | Part III. Product – Disease Category, Analyte and Method |  |  |  |  | Part IV. Regulatory and Commercial Status of the Product | [12 empty] | Part V. Manufacturer – Quality Management System | [8 empty] | Part VI. Product –Safety, efficacy, Operation | [15 empty] | Part VII. Environmental Management System |  | Part VIII. Other requirements
Bid Item No. | Product Sub- category | UNFPA Item Description | 1.1.1 Name of submitter | 1.1.3 Contact person’s name | 1.1.6 Status of the submitter: Legal manufacturer or Distributor- Trader | 1.2.1 Name of manufacturer | 1.2.2 Manufacturer physical address | 1.2.7 List of all manufacturing sites (subcontractor)/addresses | 1.2.8 Type work conducted by subcontractor e.g final assembly, final sterilization, sub part manufacturing is done by other party than 1.2) | 1.2.9 Date of last inspection of the manufacturer | Pack Size | 2.1.1 Product name | 2.1.2 Full product name | 2.1.3a Kit(s) Configuration | 2.1.3.b Product code | 2.1.3.c Type of Component and 2.1.3.d Description of Components | 2.1.4 Other accessories required but not provided (name and product code(s) where applicable | 2.2 EC Class Describe | 2.3 Product utilization (2.3.1 Intended use/ purpose 2.3.2 Population being tested; 2.3.3 Intended user (lay person or professional) | 2.4 Photos of product are provided (YES/NO) | 3.1.1 Disease category (the specific disorder, condition or risk factor of interest that it is intended to detect, define or differentiate) | 3.1.2 Analyte - (Name and type of analyte molecule: Antibody, Antigen, NA, etc.) | 3.1.3 IVD method of analysis | 3.1.4 IVD function (e.g. screening, monitoring, diagnosis or aid to diagnosis); | 3.2.1 Select the specimen type(s) to be used with the product | 4.1 Is Item prequalified by WHO or approved by other UN procurement agencies for long term contract (e.g. WHO Procurement, UNICEF) Describe agency and provide details like year of approval | 4.2 Has the product undergone one of the following: 1) a WHO Emergency Use Evaluation and Listing of IVDs (EUAL), 2) an US FDA Emergency Use of Medical Products and Related Authorities (EUA), 3) an approval process from Stringent Regulatory Authority (SRA) designated by Global Harmonization Task Force (GHFT) Competent Authority. IMDRF | 4.3 Is the product for “Research use only” or “For export only”? YES/NO | 4.4 Is the IVD device considered a humanitarian use device (HUD) and applying for marketing approval through a humanitarian device exemption (HDE)? YES/NO | 4.5 Provide details of any other current regulatory approvals for this product (Do not include ISO 13485 certification details here. This is covered later ) Product name Product code Period of approval: Start (DD/MM/YY) - Expiry (DD/MM/YY) | 4.6.1 Valid EC certificate (= CE mark)? YES/NO | 4.6.2 Describe expiry date of the EC certificate MM/YYYY | 4.6.3 The name and country of EC certifying authority (e.g. notifying body) Describe | 4.6.4 EC Representative's name and country. Describe | 4.7 Valid FDA approval and 510k (or PMA) clearance? YES/NO | 4.8 Is the product approved by any other IMDRF founding members? YES/ NO Nature of certification, (number, name, annexes): Identification of the Notified Body (+ identification number): | 4.9 Is the product approved by National Regulatory Agency? YES/ NO Nature of the approval/certification Name of the National Regulatory Agency: | 3.4 ISO 17050 Declaration of Conformity Signed and dated and with correct scope? YES/NO | 5.1.1 ISO 13485 - QMS converting MDD 93/42/EEC, MDR 2017/745 for Medical Devices in correct device class Valid? YES/NO | 5.1.2 Describe expiry date of the QMS 13485 certificate MM/YYYY | 5.1.3 The name and country of QMS 13485 certifying authority (e.g. notifying body) Describe | 5.2.1 ISO 9001 - QMS Valid? YES/NO | 5.2.2 Describe expiry date of the QMS 9001 certificate MM/YYYY | 5.2.3 The name and country of QMS 9001 certifying authority (e.g. notifying body) Describe | 5.3 FDA CFR 21 (820) | 5.4 Compliance to Other Item-specific Safety Standards Describe Standard | 5.5 Free Sale Certificate, Certificate for Export or similar YES/NO | 6.1.1 ISO 14971:2012 considerations to show analysis has taken place for products sold in the EU minimizing risk | 6.1.2 Performance Evaluation □ Performance □ Specificity □ Radio of False to true results □ Sensitivity □ Limit of Detection □ Support Scientific literature | 6.2.1 Long-term stability studies conditions Describe | 6.2.2 Long-term stability studies duration Describe | 6,2.3 Long-term stability studies periods of testing Describe | 6.2.4 Long-term stability studies packaging material Describe | 6.2.5 Storage conditions and shelf-life supported by data Please specify | 6.2.6 Product suitable for use in which ICH Climatic Zones? (Zone I, II, II, IVa, IVb) | 6.3.1 Special storage conditions. (transport, strage, operation temperature) Describe, if applicable. | 6.3.2 Shelf-life upon manufacture (months) | 6.4 Labels, primary & secondary packaging, Instruction for use etc. where applicable inludede YES/ NO | 6.4.1 Instructions for use are provided in all 3 languages (EN, FR, SP). YES/NO/not applicable | 6.4.2 Packaging requirements are efficient enough to protect the product? YES/NO | 6.4.3 Is the cardboard carton FSC marked? YES/NO | 6.4.4 Photos of primary and secondary LABELS are submitted YES/NO | 6.4.5 Photos of primary and secondary packaging are submitted YES/NO | 7.1 ISO 14001 - Environment Valid YES/NO | 7.2 ISO 50001 - Energy Management Valid YES/NO | 8.1.1 Reusable, re-sterilized Item? YES/NO | 8.1.2 Reusable, re-sterilized Item - Instructions are provided for efficient cleaning, desinfecting and sterilization, and certificate to prove that material endures sterilization process. YES/NO/not applicable | 8.1.3 Sterile or single use Item? YES/NO | 8.1.4 Sterile or single use Item - Method of sterilization Describe | 8.1.5 Sterile or single use Item - ISO 13485 for specific sterilization process: ISO 11135 (ETO) ISO 11137 (gamma radiat) ISO 17665 (steam) ISO 20857 (dry heat) ISO 14937 (other) YES/NO/not applicable | 8.1.6 Sterile or single use Item - Three (3) latest, individual batch sterilization certificates are submitted YES/NO/not applicable | 8.2 Training and Support Needed for device operation YES/NO/not applicable | 8.3 Warranty - sufficient warranty and copy of it is provided YES/NO/not applicable | 8.4 Electrical Equipment - Large variety of voltages, frequencies and plugs quaranteed for various customers globally. YES/NO/not applicable | 8.5 Information of Safe Disposal was provided and approved YES/NO
344 | Blood Tests | Tile, porcelain, with depressions for blood grouping
528 | HIV Tests | HIV 1+2 rapid test (Colloidal Test)
529 | HIV Tests | HIV 1/2 Rapid Antibody Test
530 | HIV Tests | Anti-HIV 1+2 ELISA
531 | Hepatitis Tests | Anti - HCV Rapid Test
532 | Hepatitis Tests | HCV Antibody Test
533 | Hepatitis Tests | Test Hepatitis B surface antigen (HBsAg)
534 | Hepatitis Tests | Test Hepatitis C virus (HCV)
535 | HIV Tests | HPV Test (Professional Use Only)
536 | HIV Tests | HPV Test
537 | HIV Tests | HIV Blot Elisa Test
538 | HIV Tests | Anti-HIV 1+2 ELISA Blot
539 | HIV Tests | HIV Ag/Ab Combination
540 | HIV Tests | HIV (1+2) antibody (colloidal gold)
541 | HIV-Syphilis Tests | HIV1+2/Syphilis
542 | Syphilis Tests | Syphilis Rapid Test
543 | HIV Tests | HIV Self Test
544 | HIV Tests | HIV AntibodyTest
545 | Urine Tests | Urine Test – strips
546 | Blood Tests | Blood Test - Anti A
547 | Blood Tests | Blood Test - Anti B
548 | Blood Tests | Blood Test - Anti A+B
549 | Blood Tests | Test Rhesus, anti D
550 | Pregnancy Tests | Pregnancy Test, Strips

## 5. Specifications MDs and IVDs
Tab 5. Technical Specification Form for the Medical Devices and IVDs
| Name of Company (Please enter in the cell on the right): | to be filled-in
Other comments (if any):
Important guidance to Bidders:
1. Please fill in the blanks under orange column.
2. Technical Specification Form should be submitted in excel form.
3. Complete every single row following Technical Requirements for Medical Devices.
4. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font (red).
5. Bidders are not allowed to copy and paste UNFPA generic specifications or to use general statements like: “offer as per specifications”, “comply”, “offered as requested”, “as per requirement”, “yes, but different pack size instead - pack of 1000“ or similar general statements/information. Similar statements might invalidate the bid. Deviations between this sheet and product literature (e.g, manufacturer brochures/webpage) may invalidate the bid. Bidders should include their corresponding specification to each item they have quoted in each line in the column named “ Bidders Technical Specification of the Offered Product”.
Bid Item No. | Sub-category | Item ID for reference purpose only | UNFPA Item Description (protected copy of specs for UNFPA internal use) | UNFPA Item Description | UNFPA Detailed Technical Specifications (protected copy of specs for UNFPA internal use) | UNFPA Detailed Technical Specifications | Bidders Technical Specification of the Offered Product | Bidders Comments
|  |  | EXAMPLE. DO NOT BID. | EXAMPLE. DO NOT BID. | Product description: Basic, sterile cutting instrument for use in minor surgical incisions, e.g. in cutting umbilical cord. Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation | Product description: Basic, sterile cutting instrument for use in minor surgical incisions, e.g. in cutting umbilical cord. Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation | Product description: Basic, sterile cutting instrument for use in minor surgical incisions, e.g. in cutting umbilical cord. Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard. Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control. Ridges (5) in the handle. Material, blade: Martensitic stainless steel (quenched, magnetic steel) AISI 410. Material, handle: Robust DEHP free ABS. Latex-free. Material, guard/hood: PVC contains DEHP. Latex-free. Hardness, blade: minimum 55 HRC. Sterile and single-use. Disposable. Length: 16.5 cm (blade+handle, excluding the blade guard) Sterilization method: Gamma radiation
[5 empty] | All dimensions and volume related specs will have a 3 - 10% tolerance applied | All dimensions and volume related specs will have a 3 - 10% tolerance applied
101 | Anaesthesia & Resuscitation Equipment | New | Anesthesia machine | Anesthesia machine | GENERAL DESCRIPTION: Anaesthesia unit, closed or semi-closed circuit, AC and battery powered , inlcuding vaporizer, mixer and other accessories. INTENDED USE: Anaesthesia units dispense a mixture of gases and vapours and vary the proportions to control a patient's level of consciousness and/or analgesia during surgical procedures. NOTE : Anaesthesia units primarily perform the following four functions: - Provide oxygen (O₂) to the patient - Blend gas mixtures, in addition to O₂, that include air or nitrous oxide (N₂O) along with an anaesthetic vapor. - Facilitate spontaneous, controlled, or assisted ventilation while using these gas mixtures. - Reduce, if not eliminate, anaesthesia-related risks to the patient and clinical staff. The patient is anesthetized by inspiring a mixture of O₂, the vapor of a volatile liquid halogenated hydrocarbon anaesthetic, and, if necessary, N₂O and other gases. Because normal breathing is routinely depressed by anaesthetic agents and by muscle relaxants administered in conjunction with them, respiratory assistance is also provided via either an automatic ventilator or by manual compression of the reservoir bag. IMPORTANT: Anaesthesia units are highly complex devices requiring specific infrastructure provision to be able to operate. They need a scavenging system, pipped medical air supply, and oxygen and nitrous oxide sources which can be piped or from cylinders. In addition, these machines require trained operators and technicians to carry out preventative maintenance so the units can operate safely. RECOMMENDATION : Prior to procurement it is to be assured that supporting supplies such as (bottled or mural) oxygen and compressed air, nitrious oxide and halothane are systematically available at the hospital site. Wall plug shall be made available at country level. PRODUCT DESCRIPTION : A gas machine for anaesthesia is a device that continuously or intermittently administers and monitors a mixture of gases (oxygen, nitrous oxide, and volatile anaesthetics) to aid in providing a patient with an appropriate level of anaesthesia, and to maintain a patient's ventilation. Anaesthesia machine closed breathing circuit configuration. Suitable for paediatric and adult patients. Stable sturdy construction on 4 antistatic swivel castors, 2 with brakes. Trolley with upper shelf that to mount a monitor in a flexible and easy way, allowing an ergonomic viewing Handles facilitate positioning of the machine. Movable arm holder for supporting patient breathing circuit. If O2 / N2O supplied by bottle, holders for bottles to include secure locking mechanism. MATERIAL: Tubing and mask: DEHP free The unit is provided with a minimum of three (3) gas inlets (O₂, N₂O and Air). Gas inlet connections are compliant with DISS or NIST (needs to be specified when ordering). The unit is provided with a minimum of two (2) gas cylinder yokes (O₂, and N₂O). All pipeline connections have diameter-indexed safety systems (DISSs), or another means of preventing connection of dangerous gases. The unit is equipped with a minimum of four (4) analogue rotameters, two (2) for O₂, one (1) for air and one (1) for N₂O with programmed parameters visualization. The O₂ and N₂O flowmeters have a minimum range of 0.0 – 10 L/min, and a resolution of at least 0.2 L/min. Externally supplied anaesthetic gas, oxygen and air mixture ratios fully controllable. Separate flow control for O2, N20 and anaesthetic agents. Flows and the mixture ratios determined from flowmeter settings are accurate to within ± 10% of set values or better. The unit has the ability to carry out self-diagnosis and integrity testing, including a compliance and leakage test. Mixer secures a minimum of 25% oxygen The vaporizers are provided with an interlock system, preventing the use of more than one vaporizer simultaneously. A built-in rechargeable battery allows for an autonomy of at least one (1) hour。The unit will switch back and forth between battery and mains operation in case of power failures. Recharging resumes automatically when extern power becomes available." "Gas supply gauges required, with scales allowing easy reading. Gas supply input 2 to 6bar Minimum of 2 vaporizers. Nebuliser to deliver particle size of < 3 micron and to be used in both offline and online modes. Soda lime absorber, with 2.5kg reservoir and adjustable pressure limiting valve Non-return and three way valve, with connecting tube Ventilator: Modes Automatic Volumetric (IPPV) and Manual Pneumatically driven, electronically controlled compressor, minute volume 2 to 25L/minute Tidal volume 20 to 1500ml Respiratory rate 5 to 70 cycles/minute The unit supports an I/E ratio 4:1 to 1:10 Inspiration pressure 0 to 80mbar Peak inspiratory flow 0 to 60L/minute Trigger sensitivity 0 to -20mbar Display fit with broncho manometer, range approximately -10 to 100mbar" "Facility to measure and display on screen: a) 3 traces against time: pressure, volume and flow b) 3 two-axis displays: Pressure-Volume, Flow-Volume and Pressure-Flow c) Airway pressure (Peak and Mean). d) Tidal volume (Inspired and Expired). e) Minute volume (Inspired and Expired). f) I:E ratio g) inspiration and expiration times h) Spontaneous Minute Volume i) Respiratory rate (spontaneous and mechanical). j) Total Frequency. k) Oxygen concentration l) End tidal CO2 with capnography m) FiO2 dynamic. n) Intrinsic PEEP and PEEPi Volume. o) Plateau Pressure. p) Resistance and Compliance. Trend display facility for at least the last 24 hours with minimum 5 minutes resolution Automatic compliance and leakage compensation for circuit and tubes." "Front panel shows status, errors and sensors failure (low/high pressure, power failure) Audio-visual alert on low/high pressure, apnoea, power failure Display of operational status, with set and measured values Front panel shows status and errors (low/high pressure, power failure, battery status) Safety features for hypoxic mixtures, oxygen failure (emergency O2 bypass), overpressures Self-diagnosis with each start-up and integrity testing of all system parameters With adjustable patient-circuit support arm Built-in rechargeable battery, autonomy approximately 2 hours Automatic switch to battery in case of power failure, automatic recharge when connected to mains Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Power input to be fitted with compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Resettable overcurrent breaker required on both live and neutral supply lines. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. An internal battery capable of powering the unit for at least 30 minutes. ALARM AND SAFETY FUNCTIONALITIES Alarms are categorized in three categories: caution, advisory and alarm. The unit is equipped with audio and visual alarms for the following parameters: - High airway pressure alarm. - Sub atmospheric pressure alarm. - Respiratory rate alarm. - Minute volume alarm. - Tidal volume alarm. - Expiratory flow alarm. - FiO₂ supply failure alarm. - O₂ supply failure alarm. - Low pressure/apnoea alarm. - External O₂ gas supply failure alarm. - Low battery alarm. - Power failure alarm. - Sensor disconnected alarm. - System or sensor failures alarm. The unit has a provision that prevents an anaesthesia machine from being set to dispense a hypoxic mixture. The N₂O and O₂ flow controls are interlocked so that the proportion of O₂ to N₂O can never fall below a minimum value of 25% to produce a hypoxic breathing mixture. The unit has a provision that protects the patient from inadequate O₂ supply. If the O₂ supply pressure drops below 1.7 - 2.1 Bar (25 to 30 psi), the unit decreases or shuts off the flow of other gases and activates an alarm. The unit is equipped with an Adjustable Pressure Limiting (APL) valve to prevent delivery from too high-pressured gas. The unit should have a provision for emergency O₂ by-pass." "SUPPLIED WITH : 1 x Paediatric reusable breathing circuit (tubes/balloons/valves/masks) 1 x Adult reusable breathing circuits (tubes/balloons/valves/masks) 2 x Soda lime absorber grains, pack/2.5kg 1 x Spare parts/maintenance kit (air filters, tubing, O-rings) 2 x Sets of reusable masks, each of: small, medium and large (textured for easy fit) 1 x Set of spare fuses 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish)" INSTRUCTION FOP USE : 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish) PACKAGING & LABELLING : Unit presentation 1 (one) Anaesthesia machine, w/access Symbols used according ISO 15223 CE mark and notified body number REGULATORY AND COMPLIANCE : CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION : Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS : Must comply with following standards IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 5356-2:2012 Anaesthetic and respiratory equipment -- Conical connectors -- Part 2: Screw-threaded weight-bearing connectors ISO 5358:1992 Anaesthetic machines for use with humans ISO 5360:2012 Anaesthetic vaporizers -- Agent-specific filling systems ISO 5361:2012 Anaesthetic and respiratory equipment -- Tracheal tubes and connectors ISO 5362:2006 Anaesthetic reservoir bags ISO 5364:2008 Anaesthetic and respiratory equipment -- Oropharyngeal airways ISO 5366-1:2000 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 1: Tubes and connectors for use in adults SO 5366-3:2001 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 3: Paediatric tracheostomy tubes ISO 5367:2000 Breathing tubes intended for use with anaesthetic apparatus and ventilators ISO 7396-2:2007 Medical gas pipeline systems -- Part 2: Anaesthetic gas scavenging disposal systems ISO 8835-2:2007: Inhalational anaesthesia systems - Part 2: Anaesthetic breathing systems ISO 8835-3:2007: Inhalational anaesthesia systems - Part 3: Transfer and receiving systems of active anaesthetic gas scavenging systems " "ISO 8835-4:2004: Inhalational anaesthesia systems - Part 4: Anaesthetic vapour delivery devices ISO 8835-5:2004: Inhalational anaesthesia systems - Part 5: Anaesthetic ventilators ISO 8835-7:2011 Inhalational anaesthesia systems -- Part 7: Anaesthetic systems for use in areas with limited logistical supplies of electricity and anaesthetic gases ISO 9170-2:2008 Terminal units for medical gas pipeline systems -- Part 2: Terminal units for anaesthetic gas scavenging systems ISO 9360-1:2000 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 1: HMEs for use with minimum tidal volumes of 250 ml ISO 9360-2:2001 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 2: HMEs for use with tracheostomized patients having minimum tidal volumes of 250 ml ISO 11712:2009 Anaesthetic and respiratory equipment -- Supralaryngeal airways and connectors ISO 15001:2010 Anaesthetic and respiratory equipment -- Compatibility with oxygen ISO 10524-1:2006 Pressure regulators for use with medical gases -- Part 1: Pressure regulators and pressure regulators with flow-metering devices ISO/TS 18835:2004 Inhalational anaesthesia systems -- Draw-over vaporizers and associated equipment ISO 21969:2009 High-pressure flexible connections for use with medical gas systems ISO 23328-1:2003 Breathing system filters for anaesthetic and respiratory use -- Part 1: Salt test method to assess filtration performance ISO 23328-2:2002 Breathing system filters for anaesthetic and respiratory use -- Part 2: Non-filtration aspects ISO 23747:2007 Anaesthetic and respiratory equipment -- Peak expiratory flow meters for the assessment of pulmonary function in spontaneously breathing humans ISO 26782:2009 Anaesthetic and respiratory equipment -- Spirometers intended for the measurement of time forced expired volumes in humans ISO 27427:2010 Anaesthetic and respiratory equipment -- Nebulizing systems and components ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators ISO 80601-2-13:2011 Medical electrical equipment -- Part 2-13: Particular requirements for basic safety and essential performance of an anaesthetic workstation ISO 80601-2-55:2011 Medical electrical equipment -- Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors" ENVIRONMENTAL REQUIREMENTS: Rohs compliance" | GENERAL DESCRIPTION: Anaesthesia unit, closed or semi-closed circuit, AC and battery powered , inlcuding vaporizer, mixer and other accessories. INTENDED USE: Anaesthesia units dispense a mixture of gases and vapours and vary the proportions to control a patient's level of consciousness and/or analgesia during surgical procedures. NOTE : Anaesthesia units primarily perform the following four functions: - Provide oxygen (O₂) to the patient - Blend gas mixtures, in addition to O₂, that include air or nitrous oxide (N₂O) along with an anaesthetic vapor. - Facilitate spontaneous, controlled, or assisted ventilation while using these gas mixtures. - Reduce, if not eliminate, anaesthesia-related risks to the patient and clinical staff. The patient is anesthetized by inspiring a mixture of O₂, the vapor of a volatile liquid halogenated hydrocarbon anaesthetic, and, if necessary, N₂O and other gases. Because normal breathing is routinely depressed by anaesthetic agents and by muscle relaxants administered in conjunction with them, respiratory assistance is also provided via either an automatic ventilator or by manual compression of the reservoir bag. IMPORTANT: Anaesthesia units are highly complex devices requiring specific infrastructure provision to be able to operate. They need a scavenging system, pipped medical air supply, and oxygen and nitrous oxide sources which can be piped or from cylinders. In addition, these machines require trained operators and technicians to carry out preventative maintenance so the units can operate safely. RECOMMENDATION : Prior to procurement it is to be assured that supporting supplies such as (bottled or mural) oxygen and compressed air, nitrious oxide and halothane are systematically available at the hospital site. Wall plug shall be made available at country level. PRODUCT DESCRIPTION : A gas machine for anaesthesia is a device that continuously or intermittently administers and monitors a mixture of gases (oxygen, nitrous oxide, and volatile anaesthetics) to aid in providing a patient with an appropriate level of anaesthesia, and to maintain a patient's ventilation. Anaesthesia machine closed breathing circuit configuration. Suitable for paediatric and adult patients. Stable sturdy construction on 4 antistatic swivel castors, 2 with brakes. Trolley with upper shelf that to mount a monitor in a flexible and easy way, allowing an ergonomic viewing Handles facilitate positioning of the machine. Movable arm holder for supporting patient breathing circuit. If O2 / N2O supplied by bottle, holders for bottles to include secure locking mechanism. MATERIAL: Tubing and mask: DEHP free The unit is provided with a minimum of three (3) gas inlets (O₂, N₂O and Air). Gas inlet connections are compliant with DISS or NIST (needs to be specified when ordering). The unit is provided with a minimum of two (2) gas cylinder yokes (O₂, and N₂O). All pipeline connections have diameter-indexed safety systems (DISSs), or another means of preventing connection of dangerous gases. The unit is equipped with a minimum of four (4) analogue rotameters, two (2) for O₂, one (1) for air and one (1) for N₂O with programmed parameters visualization. The O₂ and N₂O flowmeters have a minimum range of 0.0 – 10 L/min, and a resolution of at least 0.2 L/min. Externally supplied anaesthetic gas, oxygen and air mixture ratios fully controllable. Separate flow control for O2, N20 and anaesthetic agents. Flows and the mixture ratios determined from flowmeter settings are accurate to within ± 10% of set values or better. The unit has the ability to carry out self-diagnosis and integrity testing, including a compliance and leakage test. Mixer secures a minimum of 25% oxygen The vaporizers are provided with an interlock system, preventing the use of more than one vaporizer simultaneously. A built-in rechargeable battery allows for an autonomy of at least one (1) hour。The unit will switch back and forth between battery and mains operation in case of power failures. Recharging resumes automatically when extern power becomes available." "Gas supply gauges required, with scales allowing easy reading. Gas supply input 2 to 6bar Minimum of 2 vaporizers. Nebuliser to deliver particle size of < 3 micron and to be used in both offline and online modes. Soda lime absorber, with 2.5kg reservoir and adjustable pressure limiting valve Non-return and three way valve, with connecting tube Ventilator: Modes Automatic Volumetric (IPPV) and Manual Pneumatically driven, electronically controlled compressor, minute volume 2 to 25L/minute Tidal volume 20 to 1500ml Respiratory rate 5 to 70 cycles/minute The unit supports an I/E ratio 4:1 to 1:10 Inspiration pressure 0 to 80mbar Peak inspiratory flow 0 to 60L/minute Trigger sensitivity 0 to -20mbar Display fit with broncho manometer, range approximately -10 to 100mbar" "Facility to measure and display on screen: a) 3 traces against time: pressure, volume and flow b) 3 two-axis displays: Pressure-Volume, Flow-Volume and Pressure-Flow c) Airway pressure (Peak and Mean). d) Tidal volume (Inspired and Expired). e) Minute volume (Inspired and Expired). f) I:E ratio g) inspiration and expiration times h) Spontaneous Minute Volume i) Respiratory rate (spontaneous and mechanical). j) Total Frequency. k) Oxygen concentration l) End tidal CO2 with capnography m) FiO2 dynamic. n) Intrinsic PEEP and PEEPi Volume. o) Plateau Pressure. p) Resistance and Compliance. Trend display facility for at least the last 24 hours with minimum 5 minutes resolution Automatic compliance and leakage compensation for circuit and tubes." "Front panel shows status, errors and sensors failure (low/high pressure, power failure) Audio-visual alert on low/high pressure, apnoea, power failure Display of operational status, with set and measured values Front panel shows status and errors (low/high pressure, power failure, battery status) Safety features for hypoxic mixtures, oxygen failure (emergency O2 bypass), overpressures Self-diagnosis with each start-up and integrity testing of all system parameters With adjustable patient-circuit support arm Built-in rechargeable battery, autonomy approximately 2 hours Automatic switch to battery in case of power failure, automatic recharge when connected to mains Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Power input to be fitted with compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Resettable overcurrent breaker required on both live and neutral supply lines. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. An internal battery capable of powering the unit for at least 30 minutes. ALARM AND SAFETY FUNCTIONALITIES Alarms are categorized in three categories: caution, advisory and alarm. The unit is equipped with audio and visual alarms for the following parameters: - High airway pressure alarm. - Sub atmospheric pressure alarm. - Respiratory rate alarm. - Minute volume alarm. - Tidal volume alarm. - Expiratory flow alarm. - FiO₂ supply failure alarm. - O₂ supply failure alarm. - Low pressure/apnoea alarm. - External O₂ gas supply failure alarm. - Low battery alarm. - Power failure alarm. - Sensor disconnected alarm. - System or sensor failures alarm. The unit has a provision that prevents an anaesthesia machine from being set to dispense a hypoxic mixture. The N₂O and O₂ flow controls are interlocked so that the proportion of O₂ to N₂O can never fall below a minimum value of 25% to produce a hypoxic breathing mixture. The unit has a provision that protects the patient from inadequate O₂ supply. If the O₂ supply pressure drops below 1.7 - 2.1 Bar (25 to 30 psi), the unit decreases or shuts off the flow of other gases and activates an alarm. The unit is equipped with an Adjustable Pressure Limiting (APL) valve to prevent delivery from too high-pressured gas. The unit should have a provision for emergency O₂ by-pass." "SUPPLIED WITH : 1 x Paediatric reusable breathing circuit (tubes/balloons/valves/masks) 1 x Adult reusable breathing circuits (tubes/balloons/valves/masks) 2 x Soda lime absorber grains, pack/2.5kg 1 x Spare parts/maintenance kit (air filters, tubing, O-rings) 2 x Sets of reusable masks, each of: small, medium and large (textured for easy fit) 1 x Set of spare fuses 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish)" INSTRUCTION FOP USE : 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish) PACKAGING & LABELLING : Unit presentation 1 (one) Anaesthesia machine, w/access Symbols used according ISO 15223 CE mark and notified body number REGULATORY AND COMPLIANCE : CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION : Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS : Must comply with following standards IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 5356-2:2012 Anaesthetic and respiratory equipment -- Conical connectors -- Part 2: Screw-threaded weight-bearing connectors ISO 5358:1992 Anaesthetic machines for use with humans ISO 5360:2012 Anaesthetic vaporizers -- Agent-specific filling systems ISO 5361:2012 Anaesthetic and respiratory equipment -- Tracheal tubes and connectors ISO 5362:2006 Anaesthetic reservoir bags ISO 5364:2008 Anaesthetic and respiratory equipment -- Oropharyngeal airways ISO 5366-1:2000 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 1: Tubes and connectors for use in adults SO 5366-3:2001 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 3: Paediatric tracheostomy tubes ISO 5367:2000 Breathing tubes intended for use with anaesthetic apparatus and ventilators ISO 7396-2:2007 Medical gas pipeline systems -- Part 2: Anaesthetic gas scavenging disposal systems ISO 8835-2:2007: Inhalational anaesthesia systems - Part 2: Anaesthetic breathing systems ISO 8835-3:2007: Inhalational anaesthesia systems - Part 3: Transfer and receiving systems of active anaesthetic gas scavenging systems " "ISO 8835-4:2004: Inhalational anaesthesia systems - Part 4: Anaesthetic vapour delivery devices ISO 8835-5:2004: Inhalational anaesthesia systems - Part 5: Anaesthetic ventilators ISO 8835-7:2011 Inhalational anaesthesia systems -- Part 7: Anaesthetic systems for use in areas with limited logistical supplies of electricity and anaesthetic gases ISO 9170-2:2008 Terminal units for medical gas pipeline systems -- Part 2: Terminal units for anaesthetic gas scavenging systems ISO 9360-1:2000 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 1: HMEs for use with minimum tidal volumes of 250 ml ISO 9360-2:2001 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 2: HMEs for use with tracheostomized patients having minimum tidal volumes of 250 ml ISO 11712:2009 Anaesthetic and respiratory equipment -- Supralaryngeal airways and connectors ISO 15001:2010 Anaesthetic and respiratory equipment -- Compatibility with oxygen ISO 10524-1:2006 Pressure regulators for use with medical gases -- Part 1: Pressure regulators and pressure regulators with flow-metering devices ISO/TS 18835:2004 Inhalational anaesthesia systems -- Draw-over vaporizers and associated equipment ISO 21969:2009 High-pressure flexible connections for use with medical gas systems ISO 23328-1:2003 Breathing system filters for anaesthetic and respiratory use -- Part 1: Salt test method to assess filtration performance ISO 23328-2:2002 Breathing system filters for anaesthetic and respiratory use -- Part 2: Non-filtration aspects ISO 23747:2007 Anaesthetic and respiratory equipment -- Peak expiratory flow meters for the assessment of pulmonary function in spontaneously breathing humans ISO 26782:2009 Anaesthetic and respiratory equipment -- Spirometers intended for the measurement of time forced expired volumes in humans ISO 27427:2010 Anaesthetic and respiratory equipment -- Nebulizing systems and components ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators ISO 80601-2-13:2011 Medical electrical equipment -- Part 2-13: Particular requirements for basic safety and essential performance of an anaesthetic workstation ISO 80601-2-55:2011 Medical electrical equipment -- Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors" ENVIRONMENTAL REQUIREMENTS: Rohs compliance"
102 | Anaesthesia & Resuscitation Equipment | ASPIRATOR | Aspirator, surgical, suction pump, w/accessories | Aspirator, surgical, suction pump, w/accessories | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Suction unit including two collection jars mounted on a trolley and AC powered. IMPORTANT: These pumps have been designed run for 30 minutes continuously, after which the unit should be given a 30-minute period to cool down, during this period they should not be run. INTENDED USE A mobile/portable surgical aspirator/suction unit capable of creating a vacuum pressure up to 600 mmHg. Used for surgical procedures requiring suctioning to remove blood and irrigating fluids that accumulate in the operative field and obstruct the surgeon's view. TECHNICAL SPECIFICATIONS The unit should provide intermittent and continuous suction. Continuous is limited to 30 minutes after which the unit should be given 30 minutes to cool down. The unit should be provided with two (2) reusable collection jars. The unit should have maintenance free operating. The unit is designed in such a way that it requires little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full (overflow control valve). A gauge indicating the negative pressure generated by the unit (preferably in kPa or mmHg). A hydrophobic bacterial filter is included in the airline to the pump. The unit has user-adjustable suction between 0 and at least 600 mmHg (80 kPa). A vacuum control and on/off switch are present on the casing. The flow rate at maximum vacuum is ≥ 40 L/m (2.4 m³/h). A suction tube of at least 1.5 meters is included. The suction tube is autoclavable at 121°C. The unit is equipped with a foot switch for control of the suction. During operating the unit's noise level does not exceed 60.1 dB. The unit should be mounted on a wheeled, stable, portable stand. The stand is equipped with 360°, anti-static, swivel castors. A least two (2) castors should be lockable. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of plastic in accordance with the latest European standards. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). The collection jar is made of shatterproof polycarbonate. COLLECTION JAR: The two collection jars have a combined capacity of 2 litres, are transparent with clear graduation and allow for autoclaving at 121°C. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover (depending on the model). 2 x spare collection bottle including cover, gasket, and overflow valve. 50 x spare inlet anti-bacterial filters. 4 x spare seals for the collection jar. 1 x set of spare fuses (if applicable). 2 x sets of silicone suction tubes, diameter 8/14 mm (inner/outer), length > 2.5 m. 2 x bi-conical connectors fitting the silicone tubes. ESTIMATED LIFE SPAN 7 to 10 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 35°C / 30% – 70% RH. - Storage conditions: -10°C - 55°C / 20% – 80% RH. - Atmospheric pressure: 700 ~ 1060hPa. WEIGHT AND VOLUME - estimated Weight: 24.5 kg. Volume: 290.75 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Instructions for use: To be used in basic health infrastructures for aspiration of body fluids (such as mucus, catarrh and blood) during surgical procedures and/or emergencies. To clean bottles and silicone tubes with neutral detergent after each use, bottles could be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Accessories/ spare parts/consumables: Jar with cover, silicone tube with bi-conical connector, antibacterial filter and fuse. Packaging & Labelling: Unit presentation 1 (one) Pump,suction,surgical,2 bottles,w/access Symbols used according ISO 15223 CE mark and notified body number or similar Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards IEC 60601-1Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests - IEC 60601-1-6: Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - ISO 10079-1 Medical suction equipment -- Part 1: Electrically powered suction equipment. Environmentally requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Suction unit including two collection jars mounted on a trolley and AC powered. IMPORTANT: These pumps have been designed run for 30 minutes continuously, after which the unit should be given a 30-minute period to cool down, during this period they should not be run. INTENDED USE A mobile/portable surgical aspirator/suction unit capable of creating a vacuum pressure up to 600 mmHg. Used for surgical procedures requiring suctioning to remove blood and irrigating fluids that accumulate in the operative field and obstruct the surgeon's view. TECHNICAL SPECIFICATIONS The unit should provide intermittent and continuous suction. Continuous is limited to 30 minutes after which the unit should be given 30 minutes to cool down. The unit should be provided with two (2) reusable collection jars. The unit should have maintenance free operating. The unit is designed in such a way that it requires little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full (overflow control valve). A gauge indicating the negative pressure generated by the unit (preferably in kPa or mmHg). A hydrophobic bacterial filter is included in the airline to the pump. The unit has user-adjustable suction between 0 and at least 600 mmHg (80 kPa). A vacuum control and on/off switch are present on the casing. The flow rate at maximum vacuum is ≥ 40 L/m (2.4 m³/h). A suction tube of at least 1.5 meters is included. The suction tube is autoclavable at 121°C. The unit is equipped with a foot switch for control of the suction. During operating the unit's noise level does not exceed 60.1 dB. The unit should be mounted on a wheeled, stable, portable stand. The stand is equipped with 360°, anti-static, swivel castors. A least two (2) castors should be lockable. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of plastic in accordance with the latest European standards. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). The collection jar is made of shatterproof polycarbonate. COLLECTION JAR: The two collection jars have a combined capacity of 2 litres, are transparent with clear graduation and allow for autoclaving at 121°C. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover (depending on the model). 2 x spare collection bottle including cover, gasket, and overflow valve. 50 x spare inlet anti-bacterial filters. 4 x spare seals for the collection jar. 1 x set of spare fuses (if applicable). 2 x sets of silicone suction tubes, diameter 8/14 mm (inner/outer), length > 2.5 m. 2 x bi-conical connectors fitting the silicone tubes. ESTIMATED LIFE SPAN 7 to 10 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 35°C / 30% – 70% RH. - Storage conditions: -10°C - 55°C / 20% – 80% RH. - Atmospheric pressure: 700 ~ 1060hPa. WEIGHT AND VOLUME - estimated Weight: 24.5 kg. Volume: 290.75 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Instructions for use: To be used in basic health infrastructures for aspiration of body fluids (such as mucus, catarrh and blood) during surgical procedures and/or emergencies. To clean bottles and silicone tubes with neutral detergent after each use, bottles could be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Accessories/ spare parts/consumables: Jar with cover, silicone tube with bi-conical connector, antibacterial filter and fuse. Packaging & Labelling: Unit presentation 1 (one) Pump,suction,surgical,2 bottles,w/access Symbols used according ISO 15223 CE mark and notified body number or similar Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards IEC 60601-1Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests - IEC 60601-1-6: Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - ISO 10079-1 Medical suction equipment -- Part 1: Electrically powered suction equipment. Environmentally requirements: Rohs compliance
103 | Anaesthesia & Resuscitation Equipment | New | Incubator, Infant, automatic | Incubator, Infant, automatic | GENERAL DESCRIPTION Incubator, infant, automatic, with accessories Neonatal incubator, double-wall canopy, mobile. INTENDED USE An infant incubator provides a closed, controlled environment that warms an infant by circulating heated air over the skin. The heat is then absorbed into the body by tissue conduction and blood convection. Ideally, both the skin and core temperatures should be maintained with only minor variations. Incubators feature better temperature control than radiant warmers (overhead heating units), which can overheat or underheat the infant. However, an infant in an incubator is less accessible to the caregiver. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS The unit is equipped with an automatic/servo-controlled air and skin temperature control. Microprocessor-based system for temperature and humidity regulation. Air and skin temperature control modes use thermistor-based automatic servocontrol. Air-mode set temperature range is 23-37°C in 0.1°C increments. Automatic self-test on power up. Performs internal component diagnostics at regular intervals during operation. Includes a visual heater power indicator. Internal backup thermistor. Includes built-in water reservoir for additional air humidity. Maximum internal sound level during normal operations is ≤ 60 dB; during audible alarms ≤ 65 dB. Includes an internal Li MH battery, for alarms and memory back up functionality only. The battery does not support normal operation of the unit. Power requirements either 110 Volt / 60 Hz or 220 Volts – 50 Hz to be specified at the time of ordering. COMPONENTS - Transparent cabinet. Latched door on front panel, at a minimum 30 cm high x 75 cm wide. - Minimum of four latched hand access ports and two iris ports. Minimum of five tube access ports. - Double-wall with air circulation between the hood and the double wall. - One door with air curtain. - Fixed mattress tray that tilts between 0° and at a minimum up to ± 10° Mattress with removable washable and fluid-resistant cover, minimum dimensions 65 cm long x 37 cm wide. - Display shows operational status, with set and measured values and alarms. - Protective bumpers on all four corners. - Accommodates shelves and I/V poles. - Mounted on stationary table, base of which is at least 80 cm high - Mobile version : Four (4) anti-static swivel castors that are at least 10 cm in diameter. A minimum of two (2) the castors are equipped with brakes. - At least four ports for tubes access to the interior of the hood. - At least four ports to access the patient. - At least one door or drawer or accessories base. -Monitor console platform to fit vital monitor or pulse oxymeter preferred. POWER SUPPLY The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO Power requirements either 110 Volt / 60 Hz or 220 Volts – 50 Hz to be specified at the time of ordering." INCUBATOR PERFORMANCE - Starting temperature: approx. 20ºC. - Warm up time: 20 minutes or less. - Servocontrolled mode to adjust patient's skin temperature not lower than 34ºC up to 37ºC. With an amplied range from 37ºC to 38ºC. SERVO AIR MODE SET TEMPERATURE: - Servo controlled mode to adjust air temperature from 23ºC or less to 37ºC or more Within the range of 25 - 37°C, never lower than 20°C, adjustable in 0.1°C increments, and with an accuracy equal to or better than +/-0.8°C. Should include an override up to 39°C. - With an amplied range from 37ºC to 39ºC. - Minimal resolution of 0.1 ºC. - Measurement accuracy: +/- 0.1 degree Celsius. SERVO SKIN MODE SET TEMPERATURE RANGE: Temperature alarms can be set for a minimum range of 34 - 37°C, but never lower than 32°C, adjustable in 0.1 °C increments, with an override up to 38°C and an accuracy equal to or better than +/-0.5°C. - Microcontrolled humidifier with range 40 to 80% - Oxygen input flow rate 5 to15 litres/min or oxygen concentration range 25 to 70%. - Maximum CO2 concentration inside incubator 0.2%. - Air filter: 0.5μm - Air filter capacity: 98% (for 0.5μm). - Noise level in the interior of the hood less to 50 dBA - Maximum internal sound level during normal operations is ≤ 60 dB; during audible alarms ≤ 65 dB. Includes an internal Li MH battery, for alarms and memory back up functionality only. The battery does not support normal operation of the unit." INCUBATOR MONITORING - Integrity testing of all system parameters. - Clear text display of operating and alarm situation, with set and measured values. - Safety availability test: approx. every 5 minutes. - Permanent automatic verification of probes and control heating devices. - Possible technical monitoring via PC (thorough communication plug). - Audible and visual alarms: Power failure, system failure, sensor failure, over temperature failure, air flow failure, heater failure, air temperature failure, skin temperature failure, low-water. ALARM FUNCTIONALITIES Alarms are both audible, as well visual. Alarms are available for: - High/low skin temperature. - High/low air temperature. - Air fan failure. - Temperatures exceeding 39°C in any mode. - Temperature probe failures and/or disconnects. - Power failure alarms. - Alarm if a circuit fault has been detected." SUPPLIED WITH 1 x Spare set of skin probes 2 x IV pole (with medical rail fixation clamp if applicable) 3 x Set of spare air filters 1 x Set of spare fuses 1 x Instructions for assembly, use and maintenance in English, French and Spanish INSTRUCTION FOR USE Automatic incubator for use in neonatal care units. Incubator should be operated by trained staff only. Air filters to be replaced every 6 months. Incubator must be used and maintained according manufacturer's instructions. Recommendation: wall plug shall be made available locally. ACCESSORIES/ SPARE PARTS / CONSUMABLES Spare air filter and fuses, set of skin probes. PACKAGING & LABELLING Unit presentation 1 (one) Pump,suction,surgical,2 bottles,w/access Symbols used according ISO 15223 CE mark and notified body number REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS Must comply with following standards (IEC 60601-1:2005 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-19:2009 (Part 2-19: Particular requirements for the basic safety and essential performance of infant incubators) ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 30% - 90% RH. - Storage conditions: -10°C - 70°C / < 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliance INSTALLATION REQUIREMENTS Assembly and commissioning this product should be carried out by qualified technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. ESTIMATED LIFE SPAN Between 8 – 10 years. WARRANTY Two years from shipping date. | GENERAL DESCRIPTION Incubator, infant, automatic, with accessories Neonatal incubator, double-wall canopy, mobile. INTENDED USE An infant incubator provides a closed, controlled environment that warms an infant by circulating heated air over the skin. The heat is then absorbed into the body by tissue conduction and blood convection. Ideally, both the skin and core temperatures should be maintained with only minor variations. Incubators feature better temperature control than radiant warmers (overhead heating units), which can overheat or underheat the infant. However, an infant in an incubator is less accessible to the caregiver. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS The unit is equipped with an automatic/servo-controlled air and skin temperature control. Microprocessor-based system for temperature and humidity regulation. Air and skin temperature control modes use thermistor-based automatic servocontrol. Air-mode set temperature range is 23-37°C in 0.1°C increments. Automatic self-test on power up. Performs internal component diagnostics at regular intervals during operation. Includes a visual heater power indicator. Internal backup thermistor. Includes built-in water reservoir for additional air humidity. Maximum internal sound level during normal operations is ≤ 60 dB; during audible alarms ≤ 65 dB. Includes an internal Li MH battery, for alarms and memory back up functionality only. The battery does not support normal operation of the unit. Power requirements either 110 Volt / 60 Hz or 220 Volts – 50 Hz to be specified at the time of ordering. COMPONENTS - Transparent cabinet. Latched door on front panel, at a minimum 30 cm high x 75 cm wide. - Minimum of four latched hand access ports and two iris ports. Minimum of five tube access ports. - Double-wall with air circulation between the hood and the double wall. - One door with air curtain. - Fixed mattress tray that tilts between 0° and at a minimum up to ± 10° Mattress with removable washable and fluid-resistant cover, minimum dimensions 65 cm long x 37 cm wide. - Display shows operational status, with set and measured values and alarms. - Protective bumpers on all four corners. - Accommodates shelves and I/V poles. - Mounted on stationary table, base of which is at least 80 cm high - Mobile version : Four (4) anti-static swivel castors that are at least 10 cm in diameter. A minimum of two (2) the castors are equipped with brakes. - At least four ports for tubes access to the interior of the hood. - At least four ports to access the patient. - At least one door or drawer or accessories base. -Monitor console platform to fit vital monitor or pulse oxymeter preferred. POWER SUPPLY The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO Power requirements either 110 Volt / 60 Hz or 220 Volts – 50 Hz to be specified at the time of ordering." INCUBATOR PERFORMANCE - Starting temperature: approx. 20ºC. - Warm up time: 20 minutes or less. - Servocontrolled mode to adjust patient's skin temperature not lower than 34ºC up to 37ºC. With an amplied range from 37ºC to 38ºC. SERVO AIR MODE SET TEMPERATURE: - Servo controlled mode to adjust air temperature from 23ºC or less to 37ºC or more Within the range of 25 - 37°C, never lower than 20°C, adjustable in 0.1°C increments, and with an accuracy equal to or better than +/-0.8°C. Should include an override up to 39°C. - With an amplied range from 37ºC to 39ºC. - Minimal resolution of 0.1 ºC. - Measurement accuracy: +/- 0.1 degree Celsius. SERVO SKIN MODE SET TEMPERATURE RANGE: Temperature alarms can be set for a minimum range of 34 - 37°C, but never lower than 32°C, adjustable in 0.1 °C increments, with an override up to 38°C and an accuracy equal to or better than +/-0.5°C. - Microcontrolled humidifier with range 40 to 80% - Oxygen input flow rate 5 to15 litres/min or oxygen concentration range 25 to 70%. - Maximum CO2 concentration inside incubator 0.2%. - Air filter: 0.5μm - Air filter capacity: 98% (for 0.5μm). - Noise level in the interior of the hood less to 50 dBA - Maximum internal sound level during normal operations is ≤ 60 dB; during audible alarms ≤ 65 dB. Includes an internal Li MH battery, for alarms and memory back up functionality only. The battery does not support normal operation of the unit." INCUBATOR MONITORING - Integrity testing of all system parameters. - Clear text display of operating and alarm situation, with set and measured values. - Safety availability test: approx. every 5 minutes. - Permanent automatic verification of probes and control heating devices. - Possible technical monitoring via PC (thorough communication plug). - Audible and visual alarms: Power failure, system failure, sensor failure, over temperature failure, air flow failure, heater failure, air temperature failure, skin temperature failure, low-water. ALARM FUNCTIONALITIES Alarms are both audible, as well visual. Alarms are available for: - High/low skin temperature. - High/low air temperature. - Air fan failure. - Temperatures exceeding 39°C in any mode. - Temperature probe failures and/or disconnects. - Power failure alarms. - Alarm if a circuit fault has been detected." SUPPLIED WITH 1 x Spare set of skin probes 2 x IV pole (with medical rail fixation clamp if applicable) 3 x Set of spare air filters 1 x Set of spare fuses 1 x Instructions for assembly, use and maintenance in English, French and Spanish INSTRUCTION FOR USE Automatic incubator for use in neonatal care units. Incubator should be operated by trained staff only. Air filters to be replaced every 6 months. Incubator must be used and maintained according manufacturer's instructions. Recommendation: wall plug shall be made available locally. ACCESSORIES/ SPARE PARTS / CONSUMABLES Spare air filter and fuses, set of skin probes. PACKAGING & LABELLING Incubator in 2 packages: Main body and cabinet Symbols=ISO 13485 & ISO 15223. CE marked and notified body number REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS Must comply with following standards (IEC 60601-1:2005 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-19:2009 (Part 2-19: Particular requirements for the basic safety and essential performance of infant incubators) ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 30% - 90% RH. - Storage conditions: -10°C - 70°C / < 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliance INSTALLATION REQUIREMENTS Assembly and commissioning this product should be carried out by qualified technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. ESTIMATED LIFE SPAN Between 8 – 10 years. WARRANTY Two years from shipping date.
104 | Anaesthesia & Resuscitation Equipment | INFUSIONPUMP | Infusion pump, with accessories | Infusion pump, with accessories | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Volumetric single channel infusion pump, AC and battery powered including accessories. INTENDED USE Infusion pumps are used to accurately deliver liquids through intravenous (IV) or epidural routes for therapeutic and/or diagnostic purposes. They are used in hospitals, in alternative care settings (e.g., homes, long-term care facilities, physicians' offices, outpatient infusion centres), and, occasionally, in emergency medical service vehicles. In general, infusion pumps are used when the solution to be administered must be delivered with greater accuracy than can be provided through a manually adjusted gravity administration set. TECHNICAL SPECIFICATIONS Automatic calculation of third parameter when user enters other two parameters (volume, time, and flow rate). Minimum guaranteed flow rate range of 1-1500 mL/hr in either 0.1 or 1 mL/hr increments. Keep Vein Open (KVO) rate of 1-5 mL/hr. The accuracy of the flow rate should be 5% or better. Robust design allow use in demanding environments, resistant against hospital-grade cleaning solutions, fluid proof. Capable of being mounted on mobile pole/(roll) stand, bed rail and wall-mounted rail. Open system, compatible with wide range brands of giving sets. Unit is to be calibrated to the characteristics of a specific brand infusion set through DIP switches prior to using the unit. Designed for frequent and easy dismount and disinfection with hospital-grade products Built in battery depending on the model either a lithium ion or a lithium metal hydride. Battery life lasts at least 4 hours at 25mL/hr flow rate. Automatic switch from mains to battery during power failure. Auto-off when not in use. Power requirements: 240 Volts – 50 Hz (110 Volts – 60 Hz available on request, indicate when ordering). DISPLAY FEATURES. An integrated display indicating following parameters/information: - Alarms. - Pumping status. - Volume infused. - Volume limit/Volume To Be Infused (VTBI). ALARMS AND SAFETY FEATURES Alarms are audible and visual. Ability to silence audio alarms for maximum of 2 minutes. The following alarms are included: - Air-in-line alarm. - Down-stream occlusion alarm. - Open door alarm. - Infusion complete notification. - Low/depleted battery alarm. - Incorrectly loaded set alarm. - The unit should be protected against uncontrolled gravity flow, a so-called free-flow protection. - The unit is equipped with a "control lock-out" feature, preventing tampering by patients or visitors with the controls of the unit. (depending on the supplied model). Power supply and/or internal re-chargeable battery: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO SUPPLIED WITH Instructions for assembly, use and maintenance in English. 1 x Plastic protective dustcover. 1 x Start-up set of 30 giving sets. 1 x Spare battery pack. 1 x Mounting bracket required for fixation to mobile pole/(roll) stand, bed rail and wall-mounted rail. 1 x pole/stand, option to roll. 1 x Set of spare fuses, if required. ESTIMATED LIFE SPAN 5 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 30% - 75% RH. - Storage conditions: -20°C - 55°C / 20% - 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. - Ingress protection rating: IPX3. WEIGHT AND VOLUME Weight: 12.896 kg. Volume: 53.33 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Accessories/Spare parts/Consumables: Battery pack, fuses. Instructions for use: For controlled administration of intravenous fluids to patients. Should be operated by qualified staff only. Infusion pump must be used and maintained properly according manufacturer's instructions. Recommendation: universal plug adapter shall be available at country level. Packaging & Labelling: Unit presentation: 1 (one) infusion pump, w/access Symbols used according ISO 15223 CE mark and notified body number or similar Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8 Medical electrical equipment - Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests, and guidance for alarm systems in medical electrical equipment and medical electrical systems IEC 60601-2-24 Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers Environmental requirements: Rohs compliance. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Volumetric single channel infusion pump, AC and battery powered including accessories. INTENDED USE Infusion pumps are used to accurately deliver liquids through intravenous (IV) or epidural routes for therapeutic and/or diagnostic purposes. They are used in hospitals, in alternative care settings (e.g., homes, long-term care facilities, physicians' offices, outpatient infusion centres), and, occasionally, in emergency medical service vehicles. In general, infusion pumps are used when the solution to be administered must be delivered with greater accuracy than can be provided through a manually adjusted gravity administration set. TECHNICAL SPECIFICATIONS Automatic calculation of third parameter when user enters other two parameters (volume, time, and flow rate). Minimum guaranteed flow rate range of 1-1500 mL/hr in either 0.1 or 1 mL/hr increments. Keep Vein Open (KVO) rate of 1-5 mL/hr. The accuracy of the flow rate should be 5% or better. Robust design allow use in demanding environments, resistant against hospital-grade cleaning solutions, fluid proof. Capable of being mounted on mobile pole/(roll) stand, bed rail and wall-mounted rail. Open system, compatible with wide range brands of giving sets. Unit is to be calibrated to the characteristics of a specific brand infusion set through DIP switches prior to using the unit. Designed for frequent and easy dismount and disinfection with hospital-grade products Built in battery depending on the model either a lithium ion or a lithium metal hydride. Battery life lasts at least 4 hours at 25mL/hr flow rate. Automatic switch from mains to battery during power failure. Auto-off when not in use. Power requirements: 240 Volts – 50 Hz (110 Volts – 60 Hz available on request, indicate when ordering). DISPLAY FEATURES. An integrated display indicating following parameters/information: - Alarms. - Pumping status. - Volume infused. - Volume limit/Volume To Be Infused (VTBI). ALARMS AND SAFETY FEATURES Alarms are audible and visual. Ability to silence audio alarms for maximum of 2 minutes. The following alarms are included: - Air-in-line alarm. - Down-stream occlusion alarm. - Open door alarm. - Infusion complete notification. - Low/depleted battery alarm. - Incorrectly loaded set alarm. - The unit should be protected against uncontrolled gravity flow, a so-called free-flow protection. - The unit is equipped with a "control lock-out" feature, preventing tampering by patients or visitors with the controls of the unit. (depending on the supplied model). Power supply and/or internal re-chargeable battery: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO SUPPLIED WITH Instructions for assembly, use and maintenance in English. 1 x Plastic protective dustcover. 1 x Start-up set of 30 giving sets. 1 x Spare battery pack. 1 x Mounting bracket required for fixation to mobile pole/(roll) stand, bed rail and wall-mounted rail. 1 x pole/stand, option to roll. 1 x Set of spare fuses, if required. ESTIMATED LIFE SPAN 5 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 30% - 75% RH. - Storage conditions: -20°C - 55°C / 20% - 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. - Ingress protection rating: IPX3. WEIGHT AND VOLUME Weight: 12.896 kg. Volume: 53.33 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Accessories/Spare parts/Consumables: Battery pack, fuses. Instructions for use: For controlled administration of intravenous fluids to patients. Should be operated by qualified staff only. Infusion pump must be used and maintained properly according manufacturer's instructions. Recommendation: universal plug adapter shall be available at country level. Packaging & Labelling: Unit presentation: 1 (one) infusion pump, w/access Symbols used according ISO 15223 CE mark and notified body number or similar Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8 Medical electrical equipment - Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests, and guidance for alarm systems in medical electrical equipment and medical electrical systems IEC 60601-2-24 Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers Environmental requirements: Rohs compliance.
105 | Anaesthesia & Resuscitation Equipment | OXIMETER | Pulse oximeter, portable and accessories | Pulse oximeter, portable and accessories | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Handheld, portable, battery powered pulse oximeter, continuously displaying patient oxygen saturation (SpO2) and pulse rate in real time using an external probe on the skin. Product description: Monitors arterial blood oxygen saturation (SpO2) and pulse rate (bmp) in real time using an external probe on the skin. Suitable for neonates, infants and adults TECHNICAL SPECIFICATIONS: Design tested by manufacturer in accordance with ISO 80601-2-61. Measurement: SpO2 and pulse rate monitor, with plethysmography waveform, for adults, children and neonates. SpO2 detection includes the range 70 - 100 %. Also capable of measuring lower saturations. SpO2 resolution 1 %. SpO2 accuracy (in the range 70 - 100%) within ± 3%, for all patients and perfusion / movement conditions. Pulse Rate detection includes the range 30 - 240 bpm. Pulse Rate resolution 1 bpm. Pulse Rate accuracy within ± 2 bpm or ± 2%. Environmental conditions: - Operation temperature; 0-40°C - Storage temperature: 0-50°C - Relative humidity: 10 to 90% Display: Data update period for valid data displayed ≤ 15s. Display with main parameters: SpO2, pulse rate, plethysmographic waveform, signal quality, alarm messages, battery state indication. Features: Audible and visual alarms for low/high saturation and pulse rate. Audible and visual alarms for sensor error or disconnected, system errors, low battery. Alarm override and temporary silencing function. Capable of working with adult, paediatric and neonatal reusable probes. Enclosure protection IPX1 or better. Suitable for cleaning and disinfection with hospital-grade cleaning products. Device and probe weight < 200 g. ALARM FUNCTIONALITIES Alarms are both audible, as well visual. Alarms are available for: - High/low Spo2,pulse rate - Power failure alarms. - Alarm if a circuit fault has been detected. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO Electrical characteristics: Operated by replaceable battery power supply, either rechargeable or single use. External AC battery charger, plug style and voltage as per local supply, as specified by purchaser. Running time on battery only > 12 hours. ITEMS SUPPLIED WITH: 2 x each size of reusable probes, adult, paediatric and neonatal sizes, i.e. 6 probes in total; probe cable length > 1 m (including extender cable for each if necessary) 1 x spare set rechargeable batteries 1 x carry case 1 x battery charger 1 x user manual for assembly, use and maintenance, as standard in English, French and Spanish language WEIGHT AND VOLUME (packaged) - estimated Weight: 1.881 kg Volume: 0.029 m3 WARRANTY: 24 months Instructions for use: For monitoring of percentage arterial oxygen saturation (SpO2, %) and pulse rate (HR, bpm) of all patient categories, bedside or during transportation. Non-invasive monitor using a sensor placed on a finger, a toe or on the ear lobe of the patient. Pulse oximeter must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Packaging & Labelling: Unit presentation: 1 (one) Pulse oximeter, portable, w/access Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 80601-2-61 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment ISO/IEEE 11073-10404 Health informatics -- Personal health device communication -- Part 10404: Device specialization -- Pulse oximeter Environmental requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Handheld, portable, battery powered pulse oximeter, continuously displaying patient oxygen saturation (SpO2) and pulse rate in real time using an external probe on the skin. Product description: Monitors arterial blood oxygen saturation (SpO2) and pulse rate (bmp) in real time using an external probe on the skin. Suitable for neonates, infants and adults TECHNICAL SPECIFICATIONS: Design tested by manufacturer in accordance with ISO 80601-2-61. Measurement: SpO2 and pulse rate monitor, with plethysmography waveform, for adults, children and neonates. SpO2 detection includes the range 70 - 100 %. Also capable of measuring lower saturations. SpO2 resolution 1 %. SpO2 accuracy (in the range 70 - 100%) within ± 3%, for all patients and perfusion / movement conditions. Pulse Rate detection includes the range 30 - 240 bpm. Pulse Rate resolution 1 bpm. Pulse Rate accuracy within ± 2 bpm or ± 2%. Environmental conditions: - Operation temperature; 0-40°C - Storage temperature: 0-50°C - Relative humidity: 10 to 90% Display: Data update period for valid data displayed ≤ 15s. Display with main parameters: SpO2, pulse rate, plethysmographic waveform, signal quality, alarm messages, battery state indication. Features: Audible and visual alarms for low/high saturation and pulse rate. Audible and visual alarms for sensor error or disconnected, system errors, low battery. Alarm override and temporary silencing function. Capable of working with adult, paediatric and neonatal reusable probes. Enclosure protection IPX1 or better. Suitable for cleaning and disinfection with hospital-grade cleaning products. Device and probe weight < 200 g. ALARM FUNCTIONALITIES Alarms are both audible, as well visual. Alarms are available for: - High/low Spo2,pulse rate - Power failure alarms. - Alarm if a circuit fault has been detected. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO Electrical characteristics: Operated by replaceable battery power supply, either rechargeable or single use. External AC battery charger, plug style and voltage as per local supply, as specified by purchaser. Running time on battery only > 12 hours. ITEMS SUPPLIED WITH: 2 x each size of reusable probes, adult, paediatric and neonatal sizes, i.e. 6 probes in total; probe cable length > 1 m (including extender cable for each if necessary) 1 x spare set rechargeable batteries 1 x carry case 1 x battery charger 1 x user manual for assembly, use and maintenance, as standard in English, French and Spanish language WEIGHT AND VOLUME (packaged) - estimated Weight: 1.881 kg Volume: 0.029 m3 WARRANTY: 24 months Instructions for use: For monitoring of percentage arterial oxygen saturation (SpO2, %) and pulse rate (HR, bpm) of all patient categories, bedside or during transportation. Non-invasive monitor using a sensor placed on a finger, a toe or on the ear lobe of the patient. Pulse oximeter must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Packaging & Labelling: Unit presentation: 1 (one) Pulse oximeter, portable, w/access Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 80601-2-61 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment ISO/IEEE 11073-10404 Health informatics -- Personal health device communication -- Part 10404: Device specialization -- Pulse oximeter Environmental requirements: Rohs compliance
106 | Anaesthesia & Resuscitation Equipment | OXYCONCENTRATOR | Oxygen concentrator, set | Oxygen concentrator, set | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: A mains electricity powered device that concentrates oxygen from ambient air and delivers the concentrated oxygen in a controlled manner to a patient requiring oxygen therapy, up to 10 L/min. IMPORTANT INFORMATION Oxygen concentrators are not suitable to be used as an oxygen source to other medical equipment providing respiratory support, such as ventilators, CPAP devices, High Flow Nasal Oxygen, etc. TECHNICAL SPECIFICATIONS Device concentrates oxygen from ambient air On 4 antistatic swivel castors, 2 with brakes Integrated handle allows for easy moving and positioning Device weight: 28 kg Provides continuous flow of concentrated oxygen 93% ± 3% Oxygen sensing device is integrated and measures concentration at flow meter entrance. Oxygen outlet durable, with approx.6mm barbed fitting or equivalent Flowmeter range: 1 to 10 LPM Flowmeter continuously adjustable with markings at 0.5 L intervals Maximum output flow: 10 LPM Outlet pressure: 40-70 kPa // 6-10 psi Continuous monitoring, with visual and audible alert for: Low/ high output pressure, no flow, low oxygen concentration < 82%, power failure Sound level produced: less than 52 dB (A) Power requirements: 220V-240V, 50Hz-60Hz or 110V 50/60Hz User interface easy to operate with numbers and displays that are clearly visible Digital or analogue meter to display cumulative hours of device operation Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. SUPPLIED WITH 1 x Power cord 2 x Adult cannula, with 2 m kink-resistant tubing with standard connectors. 4 x Infant cannula, with 2 m kink-resistant tubing with standard connectors. 4 x New-born cannula, with 2 m kink-resistant tubing with standard connectors. 4 x Humidifiers. 4 x Tubing adapter kit. 6 x Spare course filters 3 x Spare pre-filters 3 x Spare inlet-filters 3 x Spare bacterial-filters 1 x Instruction for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages(English, French, Spanish) Instructions for use: Oxygen concentrator produces oxygen from ambient air. For delivery of oxygen therapy to newborns and infants, the use of a flow splitter is recommended. It allows providing oxygen to 2, 3 or 4 patients simultaneously. The filters should be washed or changed at recommended intervals that vary according intensity of use and/or environment (dust): (i) external coarse filter (to be washed daily to weekly), (ii) pre-filter (to be changed monthly), (iii) inlet-filter (to be changed 6-monthly or when top of filter is discoloured), and (iv) bacterial filter (to be changed annually). The oxygen concentrator must be operated by adequately trained staff only. It is recommended to follow manufacturer's operating instructions at all times. Recommendation: wall plug shall be made available locally. ENVIRONMENTAL CONDITIONS Operating conditions: 10°C to 40°C / RH: 15 to 85% Storage conditions: -40°C to 55°C / RH: ≤ 95% Operating altitude: 0 to 2286m above sea level WEIGHT AND VOLUME - estimated Weight: 31 kg (packaged) Volume: 0.140 m3 (packaged) ESTIMATED LEAD TIME 4-6 weeks to FCA port/airport WARRANTY 24 months Packaging & Labelling: Unit presentation: 1 (one) Oxygen concentrator/SET Labelling on primary packaging must include: - Name and/or trademark of the manufacturer - Manufacturer's product reference - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging Extra information required: Number of units per box Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 8359 Oxygen concentrators for medical use - Safety requirements Environmental requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: A mains electricity powered device that concentrates oxygen from ambient air and delivers the concentrated oxygen in a controlled manner to a patient requiring oxygen therapy, up to 10 L/min. IMPORTANT INFORMATION Oxygen concentrators are not suitable to be used as an oxygen source to other medical equipment providing respiratory support, such as ventilators, CPAP devices, High Flow Nasal Oxygen, etc. TECHNICAL SPECIFICATIONS Device concentrates oxygen from ambient air On 4 antistatic swivel castors, 2 with brakes Integrated handle allows for easy moving and positioning Device weight: 28 kg Provides continuous flow of concentrated oxygen 93% ± 3% Oxygen sensing device is integrated and measures concentration at flow meter entrance. Oxygen outlet durable, with approx.6mm barbed fitting or equivalent Flowmeter range: 1 to 10 LPM Flowmeter continuously adjustable with markings at 0.5 L intervals Maximum output flow: 10 LPM Outlet pressure: 40-70 kPa // 6-10 psi Continuous monitoring, with visual and audible alert for: Low/ high output pressure, no flow, low oxygen concentration < 82%, power failure Sound level produced: less than 52 dB (A) Power requirements: 220V-240V, 50Hz-60Hz or 110V 50/60Hz User interface easy to operate with numbers and displays that are clearly visible Digital or analogue meter to display cumulative hours of device operation Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. SUPPLIED WITH 1 x Power cord 2 x Adult cannula, with 2 m kink-resistant tubing with standard connectors. 4 x Infant cannula, with 2 m kink-resistant tubing with standard connectors. 4 x New-born cannula, with 2 m kink-resistant tubing with standard connectors. 4 x Humidifiers. 4 x Tubing adapter kit. 6 x Spare course filters 3 x Spare pre-filters 3 x Spare inlet-filters 3 x Spare bacterial-filters 1 x Instruction for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages(English, French, Spanish) Instructions for use: Oxygen concentrator produces oxygen from ambient air. For delivery of oxygen therapy to newborns and infants, the use of a flow splitter is recommended. It allows providing oxygen to 2, 3 or 4 patients simultaneously. The filters should be washed or changed at recommended intervals that vary according intensity of use and/or environment (dust): (i) external coarse filter (to be washed daily to weekly), (ii) pre-filter (to be changed monthly), (iii) inlet-filter (to be changed 6-monthly or when top of filter is discoloured), and (iv) bacterial filter (to be changed annually). The oxygen concentrator must be operated by adequately trained staff only. It is recommended to follow manufacturer's operating instructions at all times. Recommendation: wall plug shall be made available locally. ENVIRONMENTAL CONDITIONS Operating conditions: 10°C to 40°C / RH: 15 to 85% Storage conditions: -40°C to 55°C / RH: ≤ 95% Operating altitude: 0 to 2286m above sea level WEIGHT AND VOLUME - estimated Weight: 31 kg (packaged) Volume: 0.140 m3 (packaged) ESTIMATED LEAD TIME 4-6 weeks to FCA port/airport WARRANTY 24 months Packaging & Labelling: Unit presentation: 1 (one) Oxygen concentrator/SET Labelling on primary packaging must include: - Name and/or trademark of the manufacturer - Manufacturer's product reference - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging Extra information required: Number of units per box Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 8359 Oxygen concentrators for medical use - Safety requirements Environmental requirements: Rohs compliance
107 | Anaesthesia & Resuscitation Equipment | RESUSCITATBAGCHIL | Resuscitator, hand-operating, child/newborn | Resuscitator, hand-operating, child/newborn | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Manual operated resuscitator for chilren between 7 and 30 kg, including three different sizes face masks, an oxygen resevoir bag with required tubing to connect to oxygen supply, and three different sizes of airway guedels (Oropharyngeal airway). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating and reusable The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is 500ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. An oxygen reservoir bag should be included with a capacity between 600-2 ,600 ml. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. The resuscitator is equipped with a pressure release (limiting) valve which safeguards against generating ventilation pressures over 45 cmH₂O, while still allowing for airway pressures up to at least 30 cmH₂O. The pressure release valve is equipped with a feature allowing it to be overridden by the user when higher pressures need to be generated in the airways. When the pressure release valve is not overridden by the user, it should either automatically return to its default position whereby it protects against excess pressures or be clearly visible so there is no doubt as to its closed position, even during an emergency. To ensure users are fully aware that the limiting valve is in closed position. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4. The masks made are of silicone rubber or any material fulfilling ISO 10993-1; ISO 10993-5; ISO 10993-10 or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set. 1 x Masks, transparent, for infant, face shape type, size 2. 1 x Masks, transparent, for child, face shape type, size 3 1 x Masks, transparent, for, large child/teenager, face shape type, size 4. 1 x Oxygen reservoir bag with a capacity between: 600-2,600 ml. 1 x Required tubing and other accessories to connect to oxygen supply. 1 x Airway Guedel, size 0, diameter 50mm. 1 x Airway Guedel, size 1, diameter 60mm. 1 x Airway Guedel, size 2, diameter 70mm. LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. Accessories/Spare parts/Consumables: Masks, airways EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME Weight: 0.7 kg Volume: 4.056 dm3 ESTIMATED DELIVERY LEAD TIME 90 days Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: children/ infant (with a body weight over 30 kg). Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10651-4: Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Manual operated resuscitator for children between 7 and 30 kg, including three different sizes face masks, an oxygen resevoir bag with required tubing to connect to oxygen supply, and three different sizes of reusable and autoclavable airway Guedel (Oropharyngeal airway). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating and reusable The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is 500ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. An oxygen reservoir bag should be included with a capacity between 600-2 ,600 ml. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. The resuscitator is equipped with a pressure release (limiting) valve which safeguards against generating ventilation pressures over 45 cmH₂O, while still allowing for airway pressures up to at least 30 cmH₂O. The pressure release valve is equipped with a feature allowing it to be overridden by the user when higher pressures need to be generated in the airways. When the pressure release valve is not overridden by the user, it should either automatically return to its default position whereby it protects against excess pressures or be clearly visible so there is no doubt as to its closed position, even during an emergency. To ensure users are fully aware that the limiting valve is in closed position. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4. The masks made are of silicone rubber or any material fulfilling ISO 10993-1; ISO 10993-5; ISO 10993-10 or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set. 1 x Masks, transparent, for infant, face shape type, size 2. 1 x Masks, transparent, for child, face shape type, size 3 1 x Masks, transparent, for, large child/teenager, face shape type, size 4. 1 x Oxygen reservoir bag with a capacity between: 600-2,600 ml. 1 x Required tubing and other accessories to connect to oxygen supply. 1 x Airway Guedel, size 0, diameter 50mm. 1 x Airway Guedel, size 1, diameter 60mm. 1 x Airway Guedel, size 2, diameter 70mm. LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. Accessories/Spare parts/Consumables: Masks, airways EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME Weight: 0.7 kg Volume: 4.056 dm3 ESTIMATED DELIVERY LEAD TIME 90 days Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: children/ infant (with a body weight over 30 kg). Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10651-4: Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements.
108 | Anaesthesia & Resuscitation Equipment | RESUSCITATNEONATE | Resuscitator, hand-operating, neonate, set | Resuscitator, hand-operating, neonate, set | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION: Manual operated resuscitator for Resuscitator intended for full-term babies, preterm and low-weight infants up to 7 kg, including two different sizes face masks, and an oxygen resevoir bag with required tubing to connect to oxygen supply. INTENDED USE Resuscitators are used for ventilation assistance in critical situations of respiratory distress or failure. It is considered essential equipment for health structures and emergency rescue team (to be available in emergency room, delivery room, operating theatre, intensive care unit, ambulance, etc.). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating and reusable The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is between 250 - 280ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. The resuscitator is equipped with a pressure release (limiting) valve which safeguards against generating ventilation pressures over 45 cmH₂O, while still allowing for airway pressures up to at least 30 cmH₂O. The pressure release valve is equipped with a feature allowing it to be overridden by the user when higher pressures need to be generated in the airways. When the pressure release valve is not overridden by the user, it should either automatically return to its default position whereby it protects against excess pressures or be clearly visible so there is no doubt as to its closed position, even during an emergency. To ensure users are fully aware that the limiting valve is in closed position. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4. The masks made are of silicone rubber or any material fulfilling ISO 10993-1; ISO 10993-5; or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. Material: Latex free SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set 1 x Mask, transparent: size 0 for preterm and low-weight baby, round type, diameter 35–50 mm 1 x Mask transparent: size 1 for term baby, round type, diameter 50–65 mm 1 x Oxygen reservoir bag with an appropriate capacity 1 x Required tubing and other accessories to connect to oxygen supply LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME - estimated Weight: 0.5 kg Volume: 4.056 dm3 ESTIMATED DELIVERY LEAD TIME 90 days Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: preterm, newborn, small infant (body weight 0-7 kg). Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10651-4: Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION: Manual operated resuscitator for Resuscitator intended for full-term babies, preterm and low-weight infants up to 7 kg, including two different sizes face masks, and an oxygen resevoir bag with required tubing to connect to oxygen supply. INTENDED USE Resuscitators are used for ventilation assistance in critical situations of respiratory distress or failure. It is considered essential equipment for health structures and emergency rescue team (to be available in emergency room, delivery room, operating theatre, intensive care unit, ambulance, etc.). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating and reusable The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is between 250 - 280ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. The resuscitator is equipped with a pressure release (limiting) valve which safeguards against generating ventilation pressures over 45 cmH₂O, while still allowing for airway pressures up to at least 30 cmH₂O. The pressure release valve is equipped with a feature allowing it to be overridden by the user when higher pressures need to be generated in the airways. When the pressure release valve is not overridden by the user, it should either automatically return to its default position whereby it protects against excess pressures or be clearly visible so there is no doubt as to its closed position, even during an emergency. To ensure users are fully aware that the limiting valve is in closed position. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4. The masks made are of silicone rubber or any material fulfilling ISO 10993-1; ISO 10993-5; or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. Material: Latex free SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set 1 x Mask, transparent: size 0 for preterm and low-weight baby, round type, diameter 35–50 mm 1 x Mask transparent: size 1 for term baby, round type, diameter 50–65 mm 1 x Oxygen reservoir bag with an appropriate capacity 1 x Required tubing and other accessories to connect to oxygen supply LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME - estimated Weight: 0.5 kg Volume: 4.056 dm3 ESTIMATED DELIVERY LEAD TIME 90 days Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: preterm, newborn, small infant (body weight 0-7 kg). Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10651-4: Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements.
109 | Anaesthesia & Resuscitation Equipment | RESUSCITATORADULT | Resuscitator, hand-operated, adult, set | Resuscitator, hand-operated, adult, set | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Manual operated resuscitator for adult patients from 30 kg and upwards, including three different sizes face masks, an oxygen resevoir bag with required tubing to connect to oxygen supply, and three different sizes of airway guedels (Oropharyngeal airway). INTENDED USE Resuscitators are used for ventilation assistance in critical situations of respiratory distress or failure. It is considered essential equipment for health structures and emergency rescue team (to be available in emergency room, delivery room, operating theatre, intensive care unit, ambulance, etc.). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating, reusable. The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is 1,600ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. An oxygen reservoir bag should be included with a capacity between 2 ,500-2,600ml. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. All parts manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4 or equivalent. The masks made are of silicone rubber or any material fulfilling ISO 10993-1:2018; ISO 10993-5:2009; ISO 10993-10:2013 or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set. 1 x Masks, transparent, for large child/teenager, face shape type, size 4. 1 x Mask, transparent, for adult, face shape type, size 5. 1 x Mask, transparent, for large adult, face shape type, size 6. 1 x Oxygen reservoir bag with a capacity between: 2,500-2,600ml. 1 x Required tubing and other accessories to connect to oxygen supply. 1 x Airway Guedel, size 2, diameter 70 mm. 1 x Airway Guedel, size 3, diameter 80 mm. 1 x Airway Guedel, size 4, diameter 90 mm. LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME - estimated Weight: 1.00 kg Volume: 6.86 dm³ ESTIMATED DELIVERY LEAD TIME 90 days Accessories/Spare parts/Consumables: Masks, airways Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: children (with a body weight over 30 kg) and adults. Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1 ISO 10651-4 Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Manual operated resuscitator for adult patients from 30 kg and upwards, including three different sizes face masks, an oxygen resevoir bag with required tubing to connect to oxygen supply, and three different sizes of airway reusable and autoclavable Guedel (Oropharyngeal airway). INTENDED USE Resuscitators are used for ventilation assistance in critical situations of respiratory distress or failure. It is considered essential equipment for health structures and emergency rescue team (to be available in emergency room, delivery room, operating theatre, intensive care unit, ambulance, etc.). TECHNICAL SPECIFICATIONS Hand-operated in a horizontal position. Portable, self-inflating, reusable. The resuscitator is autoclavable between 121°C and 134°C. The capacity of the self-inflating ventilation bag is 1,600ml. Ventilation can be done with ambient air or with oxygen. The resuscitator has an inlet valve with a nipple for oxygen tubing. The resuscitator is equipped with a non-rebreather valve. An oxygen reservoir bag should be included with a capacity between 2 ,500-2,600ml. Resuscitator can be fully disassembled, is easy to clean, disinfect, and reassemble. All parts manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. The resuscitator, all its components and accessories are free of natural rubber latex. The (intake, non-rebreather) valves and nipple are made from polycarbonate/polysulfone or any other material fulfilling ISO 10651-4 or equivalent. The masks made are of silicone rubber or any material fulfilling ISO 10993-1:2018; ISO 10993-5:2009; ISO 10993-10:2013 or equivalents. The self-inflating ventilation bag is made of transparent or translucent silicone rubber or any material fulfilling ISO 10651-4 or equivalent. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective pouch which can contain the full set. 1 x Masks, transparent, for large child/teenager, face shape type, size 4. 1 x Mask, transparent, for adult, face shape type, size 5. 1 x Mask, transparent, for large adult, face shape type, size 6. 1 x Oxygen reservoir bag with a capacity between: 2,500-2,600ml. 1 x Required tubing and other accessories to connect to oxygen supply. 1 x Airway Guedel, size 2, diameter 70 mm. 1 x Airway Guedel, size 3, diameter 80 mm. 1 x Airway Guedel, size 4, diameter 90 mm. LIFE SPAN The life span of the device is minimum 3 years or 20 autoclave sessions, whichever comes first. EXPIRY DATE This product has a minimum total shelf life of 5 years. WARRANTY Two years from shipping date ENVIRONMENTAL CONDITIONS - Operating conditions: -18°C - 45°C / 15% – 90% RH - Storage conditions: -30°C - 50°C / 40% – 90% RH WEIGHT AND VOLUME - estimated Weight: 1.00 kg Volume: 6.86 dm³ ESTIMATED DELIVERY LEAD TIME 90 days Accessories/Spare parts/Consumables: Masks, airways Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. Resuscitator is used to ventilate patients: children (with a body weight over 30 kg) and adults. Resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations. Resuscitator should only be operated by a person who has received adequate training in resuscitation technique. The resuscitator and the mask must be cleaned and disinfected after each use. Recommendation: After dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual. Packaging & Labelling: Unit presentation: 1 (one) Resuscitator in a plastic bag Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1 ISO 10651-4 Lung ventilators - Part 4: Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). ISO 14971 Medical devices – Application of risk management to medical devices ISO 5356-1 Anaesthetic and respiratory equipment – Conical connectors – Part 1: Cones and sockets. ISO 15223-1 Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements.
110 | Anaesthesia & Resuscitation Equipment | RESUSCITATTABLE | Table, resuscitation, newborn, w/accessories | Table, resuscitation, newborn, w/accessories | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION: Neonatal mobile resuscitation table, including radiant warmer, overhead light and APGAR timer and X-ray slot including accessories. INTENDED USE Newborn resuscitation and support of transition of infants at birth. Additionally, used to provide thermal support for newborns in the delivery suite and for critically ill infants who require constant nursing intervention. TECHNICAL SPECIFICATIONS The unit is supported by a sturdy and stable wheeled base. Bases is provided with a minimum four (4) anti-static swivel castors and at least two (2) of the castors are provided with brakes. The base is equipped with a side handle/rail for easy positioning. Integrated support for 2 x 10 L oxygen cylinders. Side rail and IV pole for mounting accessories and equipment. Provided with a minimum of one (1) lower drawer for storage. An X-ray cassette compartment is integrated under the infant bed. Fixed-height hood above the table integrates heating element and overhead light. The vertical column integrates controls and a display which visualises working parameters and alarms. Equipped with a Treatment /CPR timer which allows pre-sets between 1 to 59 minutes in increments of 1 min. Equipped with an APGAR timer which allows for up and down counting, and provides warning signals after 1, 5 and 10 min. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). TABLE FEATURES: Surface 58 W x 71 L cm (minimum guaranteed dimensions), minimum surface area 4,118 cm². The side panels made of transparent acrylic resin are collapsible/foldable for complete accessibility. Is equipped with a tiltable cradle, up to at least 10°. Mattress covering the entire table surface, thickness no less than 5 cm. Mattress cover flame-retardant, waterproof, and able to withstand hospital-grade disinfection products. Warm-up time < 30 min from 25°C. OVERHEAD HEATER FEATURES: Overhead radiant heater at a fixed height of 1.80m. Radiant heating is achieved by quartz/ceramic elements, placed in a parabolic reflector. Radiator is equipped with a safety grid. The unit support three (3) modes: Pre-heating, manual and servo mode. The unit allows for instant switching been the different operating modes. The minimum skin temperature setting range is 32°C to 37°C. The temperature control accuracy is: ± 0.5°C and the setting adjustment resolution is 0.1°C. In manual and pre-heating modes heating settings allow for a range of 0 – 100%. LED spotlight(s) for examination are integrated in the overhead radiant heater, the angle of the overhead light is adjustable. DISPLAY FEATURES: The display panel is integrated in the column for visualising working parameters and alarms The display indicates: - The current active mode: pre-heating, manual or servo. - The current heating power in percentages. - The current pre-set temperature. - The current actual skin temperature. - The current air temperature. The unit is equipped with a self-check feature. ALARM AND SAFETY FEATURES: Alarms are audible as well as visual. Alarms indicating: - Sensor disconnection or malfunction. - Power failure. - Skin temperature, after stabilizing, varying beyond 0.5° to 1°C from set-point. - 10 to 15 minutes after manual mode has been engaged an audible alarm will sound. In addition, the unit will either reduce or cut the power to the heater. In skin mode, when the temperature exceeds 39°C, the unit will cut the power to the heater. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Mattress. 1 x Installed and 2 x spare reusable skin temperature sensors, with connection cable and plug. 1 x Spare quartz/ceramic heating element. 1 x Set of spare fuses, if applicable. OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - Oxygen regulator including a pressure gauge, flow meter, humidifier bottle and all required tubing. - Integrated weighing scale. - Electrically operated, slow suction unit including a vacuum regulator, vacuum gauge and autoclavable collection jar. - Electrical height adjustment. - T-piece resuscitator. - Integrated phototherapy unit. Accessories/Spare parts/Consumable: Reusable skin temperature probes, halogen bulbs and fuses. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years. ENVIRONMENTAL CONDITIONS - Operating conditions: 18°C - 30°C / 30% – 75% RH. - Storage conditions: -40°C - -55°C / ≤ 95% RH. - Atmospheric pressure: 500 ~ 1060 hPa. WEIGHT AND VOLUME - estimated Weight: 131 kg. Volume: 1,299.600 dm³ ESTIIMATED DELIVERY LEAD TIME 120 days. Instructions for use: For neonatal resuscitation, general baby care (specifically additional controlled warming) and roaming in wards. New born resuscitation table with radiant warmer should be operated by qualified staff only, and must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available at locally. Packaging & Labelling: Unit presentation: 1 (one) Table, resuscitation, newborn, w/access Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-21 Medical electrical equipment -- Part 2-21: Particular requirements for the basic safety and essential performance of infant radiant warmers Environmental requirements: RoSH Compliant | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION: Neonatal mobile resuscitation table, including radiant warmer, overhead light and APGAR timer and X-ray slot including accessories. INTENDED USE Newborn resuscitation and support of transition of infants at birth. Additionally, used to provide thermal support for newborns in the delivery suite and for critically ill infants who require constant nursing intervention. TECHNICAL SPECIFICATIONS The unit is supported by a sturdy and stable wheeled base. Bases is provided with a minimum four (4) anti-static swivel castors and at least two (2) of the castors are provided with brakes. The base is equipped with a side handle/rail for easy positioning. Integrated support for 2 x 10 L oxygen cylinders. Side rail and IV pole for mounting accessories and equipment. Provided with a minimum of one (1) lower drawer for storage. An X-ray cassette compartment is integrated under the infant bed. Fixed-height hood above the table integrates heating element and overhead light. The vertical column integrates controls and a display which visualises working parameters and alarms. Equipped with a Treatment /CPR timer which allows pre-sets between 1 to 59 minutes in increments of 1 min. Equipped with an APGAR timer which allows for up and down counting, and provides warning signals after 1, 5 and 10 min. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). TABLE FEATURES: Surface 58 W x 71 L cm (minimum guaranteed dimensions), minimum surface area 4,118 cm². The side panels made of transparent acrylic resin are collapsible/foldable for complete accessibility. Is equipped with a tiltable cradle, up to at least 10°. Mattress covering the entire table surface, thickness no less than 5 cm. Mattress cover flame-retardant, waterproof, and able to withstand hospital-grade disinfection products. Warm-up time < 30 min from 25°C. OVERHEAD HEATER FEATURES: Overhead radiant heater at a fixed height of 1.80m. Radiant heating is achieved by quartz/ceramic elements, placed in a parabolic reflector. Radiator is equipped with a safety grid. The unit support three (3) modes: Pre-heating, manual and servo mode. The unit allows for instant switching been the different operating modes. The minimum skin temperature setting range is 32°C to 37°C. The temperature control accuracy is: ± 0.5°C and the setting adjustment resolution is 0.1°C. In manual and pre-heating modes heating settings allow for a range of 0 – 100%. LED spotlight(s) for examination are integrated in the overhead radiant heater, the angle of the overhead light is adjustable. DISPLAY FEATURES: The display panel is integrated in the column for visualising working parameters and alarms The display indicates: - The current active mode: pre-heating, manual or servo. - The current heating power in percentages. - The current pre-set temperature. - The current actual skin temperature. - The current air temperature. The unit is equipped with a self-check feature. ALARM AND SAFETY FEATURES: Alarms are audible as well as visual. Alarms indicating: - Sensor disconnection or malfunction. - Power failure. - Skin temperature, after stabilizing, varying beyond 0.5° to 1°C from set-point. - 10 to 15 minutes after manual mode has been engaged an audible alarm will sound. In addition, the unit will either reduce or cut the power to the heater. In skin mode, when the temperature exceeds 39°C, the unit will cut the power to the heater. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Mattress. 1 x Installed and 2 x spare reusable skin temperature sensors, with connection cable and plug. 1 x Spare quartz/ceramic heating element. 1 x Set of spare fuses, if applicable. OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - Oxygen regulator including a pressure gauge, flow meter, humidifier bottle and all required tubing. - Integrated weighing scale. - Electrically operated, slow suction unit including a vacuum regulator, vacuum gauge and autoclavable collection jar. - Electrical height adjustment. - T-piece resuscitator. - Integrated phototherapy unit. Accessories/Spare parts/Consumable: Reusable skin temperature probes, halogen bulbs and fuses. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years. ENVIRONMENTAL CONDITIONS - Operating conditions: 18°C - 30°C / 30% – 75% RH. - Storage conditions: -40°C - -55°C / ≤ 95% RH. - Atmospheric pressure: 500 ~ 1060 hPa. WEIGHT AND VOLUME - estimated Weight: 131 kg. Volume: 1,299.600 dm³ ESTIIMATED DELIVERY LEAD TIME 120 days. Instructions for use: For neonatal resuscitation, general baby care (specifically additional controlled warming) and roaming in wards. New born resuscitation table with radiant warmer should be operated by qualified staff only, and must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available at locally. Packaging & Labelling: Unit presentation: 1 (one) Table, resuscitation, newborn, w/access Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-21 Medical electrical equipment -- Part 2-21: Particular requirements for the basic safety and essential performance of infant radiant warmers Environmental requirements: RoSH Compliant
111 | Anaesthesia & Resuscitation Equipment | SUCTIONPUMPELECTR | Pump, suction, electric, aspirator, portable, w/accessories | Pump, suction, electric, aspirator, portable, w/accessories | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Portable suction unit for airways emergencies, portable, electrical with the ability to operate from mains as well as battery power, including all accessories. INTENDED USE Portable suction machine for airway aspiration, minor surgery, emergency, patient transfer, suitable for use in hospitals and at home. TECHNICAL SPECIFICATIONS The unit provides intermediate and continuous suction. Continuous is limited to 30 minutes after which the unit should be given 30 minutes to cool down. The unit should be provided with one (1) reusable collection jar. The unit should have maintenance free operating. The unit is designed in such a way that it requires little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full (overflow control valve). A gauge indicating the negative pressure generated by the unit. A hydrophobic bacterial filter is included in the airline to the pump. The user adjustable suction ranges between 0 and at least 525 mmHg (70 kPa). The flow rate at maximum vacuum is ≥ 33 L/m (2 m³/h). The unit is provided with a suction tube of at least 1.5 meters. The suction tube is made of medical grade silicone which is autoclavable at 121°C. During operating the unit's noise level does not exceed 65.5 dB. The weight of the unit does not exceed 5.5 kg. The unit should be provided with a carrying handle and storage provisions for all accessories, allowing to carry the unit with one hand. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of stainless steel or material which is non-corrosive or has a corrosion resistant coating. The unit is equipped with a lead-acid battery. The battery should be able to power the unit for up to 60 min. A visual alarm should be incorporated in the unit signalling depletion of the batteries. The unit should be provided with a battery level indicator and a charging indicator. The unit should be able to simultaneously operate and charge from 12 Volt as well as the main power grid. Power requirements: 230 Volts - 50Hz (A 110 V / 60Hz unit is available on demand). THE COLLECTION JAR The transparent, shatterproof, polycarbonate collection jar has a capacity of 1 litre and has a clear graduation. It is capable of withstanding autoclaving at 121°C. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Power consumption, approximately 100W SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish: 1 x spare collection bottle including cover, gasket, and overflow valve. 10 x spare inlet anti-bacterial filters. 2 x spare seal for the collection jar. 1 x set of spare fuses (if applicable). 2 x sets of silicone suction tubes, diameter 8/14 mm (inner/outer), length > 2.5 m. 2 x bi-conical connectors fitting the silicone tubes. ESTIMATED LIFE SPAN More than 1000 hours of operation (or 3 years). WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 35°C / 20% – 90% RH. - Storage conditions: -10°C - 55°C / 20 % – 80% RH. - Atmospheric pressure: 800 ~ 1060hPa. - Ingress protection rating: IP21. WEIGHT AND VOLUME - estimated Weight: 8.000 kg. Volume: 100.00 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Instructions for use: To be used in basic health infrastructures for aspiration of body fluids (such as mucus, catarrh and blood) during medical interventions like resuscitation or minor surgery. To clean bottle and silicone tube with neutral detergent after each use, bottle could be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Packaging & Labelling: Unit presentation: 1 (one) Pump, suction, electric, portable, with accessories Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards: IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10079-1 Medical suction equipment -- Part 1: Electrically powered suction equipment -- Safety requirements More details Environmental requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Portable suction unit for airways emergencies, portable, electrical with the ability to operate from mains as well as battery power, including all accessories. INTENDED USE Portable suction machine for airway aspiration, minor surgery, emergency, patient transfer, suitable for use in hospitals and at home. TECHNICAL SPECIFICATIONS The unit provides intermediate and continuous suction. Continuous is limited to 30 minutes after which the unit should be given 30 minutes to cool down. The unit should be provided with one (1) reusable collection jar. The unit should have maintenance free operating. The unit is designed in such a way that it requires little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full (overflow control valve). A gauge indicating the negative pressure generated by the unit. A hydrophobic bacterial filter is included in the airline to the pump. The user adjustable suction ranges between 0 and at least 525 mmHg (70 kPa). The flow rate at maximum vacuum is ≥ 33 L/m (2 m³/h). The unit is provided with a suction tube of at least 1.5 meters. The suction tube is made of medical grade silicone which is autoclavable at 121°C. During operating the unit's noise level does not exceed 65.5 dB. The weight of the unit does not exceed 5.5 kg. The unit should be provided with a carrying handle and storage provisions for all accessories, allowing to carry the unit with one hand. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of stainless steel or material which is non-corrosive or has a corrosion resistant coating. The unit is equipped with a lead-acid battery. The battery should be able to power the unit for up to 60 min. A visual alarm should be incorporated in the unit signalling depletion of the batteries. The unit should be provided with a battery level indicator and a charging indicator. The unit should be able to simultaneously operate and charge from 12 Volt as well as the main power grid. Power requirements: 230 Volts - 50Hz (A 110 V / 60Hz unit is available on demand). THE COLLECTION JAR The transparent, shatterproof, polycarbonate collection jar has a capacity of 1 litre and has a clear graduation. It is capable of withstanding autoclaving at 121°C. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Power consumption, approximately 100W SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish: 1 x spare collection bottle including cover, gasket, and overflow valve. 10 x spare inlet anti-bacterial filters. 2 x spare seal for the collection jar. 1 x set of spare fuses (if applicable). 2 x sets of silicone suction tubes, diameter 8/14 mm (inner/outer), length > 2.5 m. 2 x bi-conical connectors fitting the silicone tubes. ESTIMATED LIFE SPAN More than 1000 hours of operation (or 3 years). WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 35°C / 20% – 90% RH. - Storage conditions: -10°C - 55°C / 20 % – 80% RH. - Atmospheric pressure: 800 ~ 1060hPa. - Ingress protection rating: IP21. WEIGHT AND VOLUME - estimated Weight: 8.000 kg. Volume: 100.00 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Instructions for use: To be used in basic health infrastructures for aspiration of body fluids (such as mucus, catarrh and blood) during medical interventions like resuscitation or minor surgery. To clean bottle and silicone tube with neutral detergent after each use, bottle could be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: wall plug shall be made available locally. Packaging & Labelling: Unit presentation: 1 (one) Pump, suction, electric, portable, with accessories Symbols used according ISO 15223 CE mark and notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards: IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10079-1 Medical suction equipment -- Part 1: Electrically powered suction equipment -- Safety requirements More details Environmental requirements: Rohs compliance
112 | Anaesthesia & Resuscitation Equipment | SUCTIONPUMPFOOT | Pump, suction, foot-operated, aspirator | Pump, suction, foot-operated, aspirator | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Suction pump, manually operated by foot for pharyngeal and tracheal suction. INTENDED USE Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theatre, delivery room, intensive care unit, etc. High performance suction pump, hand or foot-operated for pharyngeal and tracheal suction. Supplied with instruction manual and diagrams covering the function of the pump, how to use it, dismantle and assemble it, to clean, disinfect and sterilize it, its maintenance and spare-parts. The suction pump should be operated only by a person who has received adequate trainingin pharyngeal and tracheal suction techniques. TECHNICAL SPECIFICATIONS The unit provides continuous suction either through: - A see-saw movement of a pedal providing suction with both up and downwards motion. Or, - Through a spring-loaded pedal which returns to its original position. The unit should have maintenance free operating requiring little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full. The user adjustable suction ranges between 0 and at least 525 mmHg (70 kPa). The flow rate at maximum vacuum is 40 L/min (2.4 m³/h). The unit is provided with a suction tube of at least 1.3 meters. The suction tube is made of medical grade silicone which is autoclavable at 121°C. The unit should be mounted on a stable platform and easy to carry. The base is to have an anti-slip provision, preventing the pump from sliding away when in use. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of stainless steel or material which is non-corrosive or has a corrosion resistant coating. Designed for frequent and easy dismount and disinfection with hospital-grade products. THE COLLECTION JAR(S) The unit is equipped with one or two collection jars with a total capacity of no less than 1L. The collection jar(s) are made of shatterproof polycarbonate. The collection jars are transparent with clear graduation and are fully autoclavable at 121°C. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish: 1 x Aspirating tube, 10 mm (internal diameter), at least 135 cm long. 1 x Spare aspirating tube, 10 mm (internal diameter), at least 135 cm long. 1 x Conical nozzle-connector, adjustable distal tip diameter 7 and 10 mm, fits narrow and wide suction catheters. 1 x Spare conical nozzle-connector, adjustable distal tip diameter 7 and10 mm, fits narrow and wide suction catheters. ESTIMATED LIFE SPAN 5 - 7 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 0°C - 50°C / 20 % – 80% RH. - Storage conditions: -10°C - 55°C / 20 % – 90% RH. WEIGHT AND VOLUME - estimated Weight: 1.550 kg. Volume: 10.00 dm³. ESTIMATED DELIVERY LEAD TIME 90 days. Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. High performance suction pump, hand or foot-operated for pharyngeal and tracheal suction. The suction pump and the aspirating tube must be cleaned and disinfected after each use. All parts can be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: Use suction tubes (sterile and disposable) for pharyngeal and tracheal suction. MAINTENANCE/USER REQUIREMENTS Use suction tubes (sterile and disposable) for pharyngeal and tracheal suction. These suction tubes fit with the aspirating tube of the suction pumpwith combination suction tip (narrow nozzle). The aspirating tube of the suction pump, foot-operated, must be supplied with a combination suction tip. This combination suction tip is made of two nozzles. The narrow nozzle can be used directly or can be connected to a suction cannula (see suction tubes, sterile, disposable). If large amounts of liquids or solid particles have to be aspirated,e.g. vomit, the narrow nozzle can easily be removed and the large nozzle(with an aperture of 10mm) can be used directly. Safety procedure: The suction pump and the aspirating tube must be cleaned and disinfected after each use. All parts can be sterilized in a steam sterilizer. Important: After dismantling and cleaning, the pump must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to closely follow manufacturer's instruction manual. Packaging & Labelling: Unit presentation: 1 (one) Pump, suction, hand or foot-operated in a plastic bag + box with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10079-2 Medical suction equipment -- Part 2: Manually powered suction equipment Environmental requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Suction pump, manually operated by foot for pharyngeal and tracheal suction. INTENDED USE Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theatre, delivery room, intensive care unit, etc. High performance suction pump, hand or foot-operated for pharyngeal and tracheal suction. Supplied with instruction manual and diagrams covering the function of the pump, how to use it, dismantle and assemble it, to clean, disinfect and sterilize it, its maintenance and spare-parts. The suction pump should be operated only by a person who has received adequate trainingin pharyngeal and tracheal suction techniques. TECHNICAL SPECIFICATIONS The unit provides continuous suction either through: - A see-saw movement of a pedal providing suction with both up and downwards motion. Or, - Through a spring-loaded pedal which returns to its original position. The unit should have maintenance free operating requiring little or no oil. The pump is equipped with a provision preventing fluid flowing to the pump if the collection jar is full. The user adjustable suction ranges between 0 and at least 525 mmHg (70 kPa). The flow rate at maximum vacuum is 40 L/min (2.4 m³/h). The unit is provided with a suction tube of at least 1.3 meters. The suction tube is made of medical grade silicone which is autoclavable at 121°C. The unit should be mounted on a stable platform and easy to carry. The base is to have an anti-slip provision, preventing the pump from sliding away when in use. The design of the device should be such that it is easy to clean/disinfect by avoiding difficult-to-reach areas and having smooth surfaces. The device should have a sturdy, shock resistant casing. The body is to be constructed of stainless steel or material which is non-corrosive or has a corrosion resistant coating. Designed for frequent and easy dismount and disinfection with hospital-grade products. THE COLLECTION JAR(S) The unit is equipped with one or two collection jars with a total capacity of no less than 1L. The collection jar(s) are made of shatterproof polycarbonate. The collection jars are transparent with clear graduation and are fully autoclavable at 121°C. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish: 1 x Aspirating tube, 10 mm (internal diameter), at least 135 cm long. 1 x Spare aspirating tube, 10 mm (internal diameter), at least 135 cm long. 1 x Conical nozzle-connector, adjustable distal tip diameter 7 and 10 mm, fits narrow and wide suction catheters. 1 x Spare conical nozzle-connector, adjustable distal tip diameter 7 and10 mm, fits narrow and wide suction catheters. ESTIMATED LIFE SPAN 5 - 7 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 0°C - 50°C / 20 % – 80% RH. - Storage conditions: -10°C - 55°C / 20 % – 90% RH. WEIGHT AND VOLUME - estimated Weight: 1.550 kg. Volume: 10.00 dm³. ESTIMATED DELIVERY LEAD TIME 90 days. Instructions for use: Basic hospital equipment for health structures and emergency situations in wards, emergency room, operating theater, delivery room, intensive care unit, ambulance, etc. High performance suction pump, hand or foot-operated for pharyngeal and tracheal suction. The suction pump and the aspirating tube must be cleaned and disinfected after each use. All parts can be sterilized in a steam sterilizer. Surgical suction pump must be used and maintained properly according manufacturer's instructions. Recommendation: Use suction tubes (sterile and disposable) for pharyngeal and tracheal suction. MAINTENANCE/USER REQUIREMENTS Use suction tubes (sterile and disposable) for pharyngeal and tracheal suction. These suction tubes fit with the aspirating tube of the suction pumpwith combination suction tip (narrow nozzle). The aspirating tube of the suction pump, foot-operated, must be supplied with a combination suction tip. This combination suction tip is made of two nozzles. The narrow nozzle can be used directly or can be connected to a suction cannula (see suction tubes, sterile, disposable). If large amounts of liquids or solid particles have to be aspirated,e.g. vomit, the narrow nozzle can easily be removed and the large nozzle(with an aperture of 10mm) can be used directly. Safety procedure: The suction pump and the aspirating tube must be cleaned and disinfected after each use. All parts can be sterilized in a steam sterilizer. Important: After dismantling and cleaning, the pump must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to closely follow manufacturer's instruction manual. Packaging & Labelling: Unit presentation: 1 (one) Pump, suction, hand or foot-operated in a plastic bag + box with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards ISO 10079-2 Medical suction equipment -- Part 2: Manually powered suction equipment Environmental requirements: Rohs compliance
113 | Anaesthesia & Resuscitation Equipment | SYRINGEPUMP | Syringe pump | Syringe pump | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Syringe pump, single channel, AC and battery powered, with accessories Volume controlled portable syringe pump for precise administration of fluids. INTENDED USE Syringe infusion pumps are used to administer intravenous (IV) fluids such as antibiotics, regional anaesthetics, antiarrhythmic medications, and chemotherapeutic agents. Syringe pumps ensure highly accurate volume delivery and consistent flow for small volumes (≤60 mL) of potent pharmacologic agents that are typically delivered at flow settings between 0.5 and 10 mL/hr. TECHNICAL SPECIFICATIONS Open system, compatible with all standard brands of syringes. Compatible with 10, 20, 30 and 50mL capacity syringes at a minimum. Continuous delivery, linear motor and plunger driven. The unit is equipped with occlusion detection. Self-test is performed each time the device is switched on. User programmable for syringe size, infusion volume, time, and flow rate. Automatic calculation of third parameter when user enters in other two (volume, time, and flow rate). Minimum guaranteed flow rate of 0.1-1300mL/hr, depending on syringe size. Minimum Keep Vein Open (KVO) rate of 0.1-1mL/hr but never greater than programmed flow rate. Volume delivered with an accuracy of at least 3%. Maximum pressure of at least 17.4 PSI / 120 kPa. User adjustable high pressure/occlusion settings. Display includes start/stop, volume limit, flow rate and volume so far delivered. System reports with audio-visual alert on operational status such as, but not limited to ready, end-of-injection. System reports with audio-visual alert on malfunctions such as, but not limited to syringe position, occlusion, low/high flow. System reports with audio-visual alert on low battery status. Ability to silence audio alarms for maximum of 2 minutes. Built in battery with a capacity to run the unit for 7 hours at 5mL/hr flow rate. Automatic switch from mains to battery in case of power failure. Automatic battery charge when mains connection is re-established. Capable of being mounted on mobile pole/(roll) stand, bed rail and wall-mounted rails. Designed for frequent and easy dismount and disinfection with hospital-grade products. Power requirements: 240 Volts - 50/60 Hz (110 Volts can be supplied if specifically requested). SUPPLIED WITH Instructions for assembly, use and maintenance in English. 1 x Spare battery pack. 1 x Mounting bracket for fixation to standard bed/wall rail and mobile pole/stand. 1 x Set of spare fuses, if applicable. ESTIMATED LIFE SPAN 5 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 90% RH. - Storage conditions: 0°C - 50°C / 90% RH. - Atmospheric pressure: 860 - 1060 hPa. WEIGHT AND VOLUME - estimated Weight: 4.559 kg. Volume: 35.000 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Packaging & Labelling: Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-24 Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers. - EN 62304 Medical device software — Software life cycle processes. Environmental requirements: Rohs compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Syringe pump, single channel, AC and battery powered, with accessories Volume controlled portable syringe pump for precise administration of fluids. INTENDED USE Syringe infusion pumps are used to administer intravenous (IV) fluids such as antibiotics, regional anaesthetics, antiarrhythmic medications, and chemotherapeutic agents. Syringe pumps ensure highly accurate volume delivery and consistent flow for small volumes (≤60 mL) of potent pharmacologic agents that are typically delivered at flow settings between 0.5 and 10 mL/hr. TECHNICAL SPECIFICATIONS Open system, compatible with all standard brands of syringes. Compatible with 10, 20, 30 and 50mL capacity syringes at a minimum. Continuous delivery, linear motor and plunger driven. The unit is equipped with occlusion detection. Self-test is performed each time the device is switched on. User programmable for syringe size, infusion volume, time, and flow rate. Automatic calculation of third parameter when user enters in other two (volume, time, and flow rate). Minimum guaranteed flow rate of 0.1-1300mL/hr, depending on syringe size. Minimum Keep Vein Open (KVO) rate of 0.1-1mL/hr but never greater than programmed flow rate. Volume delivered with an accuracy of at least 3%. Maximum pressure of at least 17.4 PSI / 120 kPa. User adjustable high pressure/occlusion settings. Display includes start/stop, volume limit, flow rate and volume so far delivered. System reports with audio-visual alert on operational status such as, but not limited to ready, end-of-injection. System reports with audio-visual alert on malfunctions such as, but not limited to syringe position, occlusion, low/high flow. System reports with audio-visual alert on low battery status. Ability to silence audio alarms for maximum of 2 minutes. Built in battery with a capacity to run the unit for 7 hours at 5mL/hr flow rate. Automatic switch from mains to battery in case of power failure. Automatic battery charge when mains connection is re-established. Capable of being mounted on mobile pole/(roll) stand, bed rail and wall-mounted rails. Designed for frequent and easy dismount and disinfection with hospital-grade products. Power requirements: 240 Volts - 50/60 Hz (110 Volts can be supplied if specifically requested). SUPPLIED WITH Instructions for assembly, use and maintenance in English. 1 x Spare battery pack. 1 x Mounting bracket for fixation to standard bed/wall rail and mobile pole/stand. 1 x Set of spare fuses, if applicable. ESTIMATED LIFE SPAN 5 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 90% RH. - Storage conditions: 0°C - 50°C / 90% RH. - Atmospheric pressure: 860 - 1060 hPa. WEIGHT AND VOLUME - estimated Weight: 4.559 kg. Volume: 35.000 dm³. ESTIMATED DELIVERY LEAD TIME 120 days. Packaging & Labelling: Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards: Must comply with following standards - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-24 Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers. - EN 62304 Medical device software — Software life cycle processes. Environmental requirements: Rohs compliance
114 | Medical Electrical Equipment | CRYOSURGICALUNIT | Cryosurgical Unit | Cryosurgical Unit | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General description: An assembly of devices designed to apply cold from a gaseous or liquid refrigerant (cryogen) [e.g., liquid nitrogen (LN2), nitrous oxide (N2O), carbon dioxide (CO2)] to a target tissue for its destruction and removal. The system typically includes a mechanical regulator to control the flow of cryogen, contained in an attached cylinder, and the probe(s) to apply the cold. The system is typically used across clinical specialties (e.g., general surgery, dermatology, oral surgery, gynaecology, urology, ENT, proctology, oncology) to remove malignant or abnormal benign tissues. Technical specification: Handheld compressed gas (nitrous oxide or carbon dioxide) operated cryosurgical unit capable of achieving temperatures at the cryotip below -79°C (-110.3oF) with CO2 or -89°C (-128.2oF) with N2O. The unit shall have a trigger mechanism to control the freeze/thaw cycle (active defrost preferred but not essential), removable circular, closed design cryotips (19 +/- 2) mm diameter with flat surfaces or with a cone extrusion not exceeding 5 mm, insulated cryoshaft of length 170 mm tod 200 mm, hose assembly (high pressure) with cylinder connector, pressure gauge and relief valve, and exhaust port to which a hose can be connected to safely vent the exhaust gas. Battery included (if required). Compressed gas in cylinders (nitrous oxide or carbon dioxide). Accessories: Mandatory: Cryoprobes according to the specific use Optional: integral timer and/or temperature indicator Spare parts: Cryotips, cryoshafts, O-ring and sealing washers, hose assembly Other components: Compressed gas in cylinders (nitrous oxide or carbon dioxide) Requirements for Commissioning: Local installation, safety and operation checks by technician or engineer. Local clinical staff to affirm completion of acceptance testing. Training of users: Training of users in operation and basic maintenance shall be provided on set up and shall be subject to local requirements Other requirements: Manufacturer's documentation on installation, operation, trouble-shooting, maintenance, advanced maintenance tasks User, technical and maintenance manuals to be supplied in English or specified language. List to be provided of equipment and procedures required for local and routine maintenance Packaging and labelling: Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Regulations: US regulations 21 CFR part 820 21CFR section 882.4250 device, surgical, cryogenic JP regulations MHLW Ordinance No.169 11067000 General-purpose cryosurgical unit | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General description: An assembly of devices designed to apply cold from a gaseous or liquid refrigerant (cryogen) [e.g., liquid nitrogen (LN2), nitrous oxide (N2O), carbon dioxide (CO2)] to a target tissue for its destruction and removal. The system typically includes a mechanical regulator to control the flow of cryogen, contained in an attached cylinder, and the probe(s) to apply the cold. The system is typically used across clinical specialties (e.g., general surgery, dermatology, oral surgery, gynaecology, urology, ENT, proctology, oncology) to remove malignant or abnormal benign tissues. Technical specification: Handheld compressed gas (nitrous oxide or carbon dioxide) operated cryosurgical unit capable of achieving temperatures at the cryotip below -79°C (-110.3oF) with CO2 or -89°C (-128.2oF) with N2O. The unit shall have a trigger mechanism to control the freeze/thaw cycle (active defrost preferred but not essential), removable circular, closed design cryotips (19 +/- 2) mm diameter with flat surfaces or with a cone extrusion not exceeding 5 mm, insulated cryoshaft of length 170 mm tod 200 mm, hose assembly (high pressure) with cylinder connector, pressure gauge and relief valve, and exhaust port to which a hose can be connected to safely vent the exhaust gas. Battery included (if required). Compressed gas in cylinders (nitrous oxide or carbon dioxide). Accessories: Mandatory: Cryoprobes according to the specific use Optional: integral timer and/or temperature indicator Spare parts: Cryotips, cryoshafts, O-ring and sealing washers, hose assembly Other components: Compressed gas in cylinders (nitrous oxide or carbon dioxide) Requirements for Commissioning: Local installation, safety and operation checks by technician or engineer. Local clinical staff to affirm completion of acceptance testing. Training of users: Training of users in operation and basic maintenance shall be provided on set up and shall be subject to local requirements Other requirements: Manufacturer's documentation on installation, operation, trouble-shooting, maintenance, advanced maintenance tasks User, technical and maintenance manuals to be supplied in English or specified language. List to be provided of equipment and procedures required for local and routine maintenance Packaging and labelling: Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Regulations: US regulations 21 CFR part 820 21CFR section 882.4250 device, surgical, cryogenic JP regulations MHLW Ordinance No.169 11067000 General-purpose cryosurgical unit
115 | Medical Electrical Equipment | New | Electrosurgical unit, w/accessories | Electrosurgical unit, w/accessories | GENERAL DESCRIPTION Electrosurgical unit, for general surgery including a trolley. INTENDED USE Electro surgical generator for use in general surgery for the cutting of soft tissues and coagulation. Electrosurgical units are used for surgical cutting and for controlling bleeding by causing coagulation (haemostasis) at the surgical site. They deliver high-frequency electrical current through an active electrode tip, causing vaporization, desiccation, or charring by resistive heating in the target tissue. Electrosurgery is useful for nearly all surgical procedures, both open and laparoscopic. The haemostatic effect of electrosurgery makes it useful for procedures on organs with oozing capillary beds, such as the liver, spleen, thyroid, and lungs, as well as for open-heart procedures in which large amounts of anticoagulants are used." "PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Electrosurgical unit, for general surgery, mono- and bi-polar, including a trolley and accessories. Microprocessor controlled generator provides mono-polar and bi-polar output High frequency Suitable for monopolar cutting and coagulation Suitable for bipolar electrosurgery Monopolar maximum power up to 200W (cutting), up to 100W (coagulation) Bipolar maximum power up to 45 W (coagulation) Modes: cutting (pure cut, blend and haemostasis) and coagulating (soft contact, force, spray and bi-polar) Front panel allows control of: power cutting, power coagulation, on/off Double function foot switch (mono and bi-polar), with yellow pedal cutting and blue pedal coagulation Hand switch mode when button-activated probes are connected Return circuit sensing monitors and deactivates generator in case patient plate fails Display reports: output power, system errors and electrode failure Automatic limitation of output voltage Automatic shut off/generator deactivation on grounding pad/return electrode connection failure. Audible and visual indicators of activation and alarms. Self-test mode. Reduced emission of interferences and limitation of HF leakage current Alarm in case of HF overpower Visual and audible activation indicators required, Visual and audible cable disconnection alarm required Protected against defibrillation shocks Autoclavable handle and electrodes that can withstands more than 250 sterilization cycles at 134ºC. Power requirements: 100 - 240 Volts - 50/60 Hz (indicate exact requirements at the time of ordering). Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. SUPPLIED WITH 1 x Plastic protective dustcover. 1 x Foot switch, two pedals, yellow and blue, with connecting cable 2 x Patient plate (grounding pad/ return electrodes), reusable, with 3m connecting cable (adult & child) 2 x Mono-polar electrode handle, reusable, foot switch controlled, with connecting cable 2 x Mono-polar electrode handle, reusable, finger switch controlled, with connecting cable 1 x Set different mono-polar reusable electrodes (needle, blade, ball and loop) 2 x Bi-polar forceps, reusable, foot switch controlled, with connecting cable (short, straight, tip-angled) 2 x Bi-polar forceps, reusable, foot switch controlled, with connecting cable (long, straight, tip-angled) 1 x Trolley on 4 antistatic swivel castors, 2 with brakes, fit with a drawer and storage for foot pedal/switch. 1 x Instructions for use: Manual with instructions for assembly, use and maintenance in English, French and Spanish" "INSTRUCTION FOR USE To be used for cutting and coagulation of soft tissues and blood vessels during surgical interventions. Electrosurgical unit must be used and maintained properly according to manufacturer's instructions. Instructions for assembly, use and maintenance in English, French and Spanish." "ACCESSORIES / SPARE PARTS / CONSUMABLES Two sets of spare fuses (if replaceable fuses used) 2 reusable, sterilizable, monopolar patient plates with minimum 2m long connection cable Two each of reusable, sterilizable, monopolar and bipolar probe Disposable plates connection cable for plates..Removable Electrodes" PACKAGING & LABELLING Unit presentation: 1 (one) Electrosurgical unit,w/accessories Symbols used according ISO 15223 CE mark and reference number of notifying body. WEIGHT AND VOLUME - estimated Weight: 28 kg. Volume: 160.00 dm³." REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS Must comply with the following standards ISO 13485:2003 IEC 60601-1:2005 IEC 60601-1-1:2000 EC 60601-1-2:2007 IEC 60601-2-2:2009" ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 30% – 75% RH. - Storage conditions: -10°C - 60°C / 30% – 90% RH. - Atmospheric pressure: 700 ~ 7060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliant WARRANTY Depending on the model two years from shipping date ESTIMATED LIFE SPAN 5 years | GENERAL DESCRIPTION Electrosurgical unit, for general surgery including a trolley. INTENDED USE Electro surgical generator for use in general surgery for the cutting of soft tissues and coagulation. Electrosurgical units are used for surgical cutting and for controlling bleeding by causing coagulation (haemostasis) at the surgical site. They deliver high-frequency electrical current through an active electrode tip, causing vaporization, desiccation, or charring by resistive heating in the target tissue. Electrosurgery is useful for nearly all surgical procedures, both open and laparoscopic. The haemostatic effect of electrosurgery makes it useful for procedures on organs with oozing capillary beds, such as the liver, spleen, thyroid, and lungs, as well as for open-heart procedures in which large amounts of anticoagulants are used." "PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Electrosurgical unit, for general surgery, mono- and bi-polar, including a trolley and accessories. Microprocessor controlled generator provides mono-polar and bi-polar output High frequency Suitable for monopolar cutting and coagulation Suitable for bipolar electrosurgery Monopolar maximum power up to 200W (cutting), up to 100W (coagulation) Bipolar maximum power up to 45 W (coagulation) Modes: cutting (pure cut, blend and haemostasis) and coagulating (soft contact, force, spray and bi-polar) Front panel allows control of: power cutting, power coagulation, on/off Double function foot switch (mono and bi-polar), with yellow pedal cutting and blue pedal coagulation Hand switch mode when button-activated probes are connected Return circuit sensing monitors and deactivates generator in case patient plate fails Display reports: output power, system errors and electrode failure Automatic limitation of output voltage Automatic shut off/generator deactivation on grounding pad/return electrode connection failure. Audible and visual indicators of activation and alarms. Self-test mode. Reduced emission of interferences and limitation of HF leakage current Alarm in case of HF overpower Visual and audible activation indicators required, Visual and audible cable disconnection alarm required Protected against defibrillation shocks Autoclavable handle and electrodes that can withstands more than 250 sterilization cycles at 134ºC. Power requirements: 100 - 240 Volts - 50/60 Hz (indicate exact requirements at the time of ordering). Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. SUPPLIED WITH 1 x Plastic protective dustcover. 1 x Foot switch, two pedals, yellow and blue, with connecting cable 2 x Patient plate (grounding pad/ return electrodes), reusable, with 3m connecting cable (adult & child) 2 x Mono-polar electrode handle, reusable, foot switch controlled, with connecting cable 2 x Mono-polar electrode handle, reusable, finger switch controlled, with connecting cable 1 x Set different mono-polar reusable electrodes (needle, blade, ball and loop) 2 x Bi-polar forceps, reusable, foot switch controlled, with connecting cable (short, straight, tip-angled) 2 x Bi-polar forceps, reusable, foot switch controlled, with connecting cable (long, straight, tip-angled) 1 x Trolley on 4 antistatic swivel castors, 2 with brakes, fit with a drawer and storage for foot pedal/switch. 1 x Instructions for use: Manual with instructions for assembly, use and maintenance in English, French and Spanish" "INSTRUCTION FOR USE To be used for cutting and coagulation of soft tissues and blood vessels during surgical interventions. Electrosurgical unit must be used and maintained properly according to manufacturer's instructions. Instructions for assembly, use and maintenance in English, French and Spanish." "ACCESSORIES / SPARE PARTS / CONSUMABLES Two sets of spare fuses (if replaceable fuses used) 2 reusable, sterilizable, monopolar patient plates with minimum 2m long connection cable Two each of reusable, sterilizable, monopolar and bipolar probe Disposable plates connection cable for plates..Removable Electrodes" PACKAGING & LABELLING Unit presentation: 1 (one) Electrosurgical unit,w/accessories Symbols used according ISO 15223 CE mark and reference number of notifying body. WEIGHT AND VOLUME - estimated Weight: 28 kg. Volume: 160.00 dm³." REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS Must comply with the following standards ISO 13485:2003 IEC 60601-1:2005 IEC 60601-1-1:2000 EC 60601-1-2:2007 IEC 60601-2-2:2009" ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 30% – 75% RH. - Storage conditions: -10°C - 60°C / 30% – 90% RH. - Atmospheric pressure: 700 ~ 7060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliant WARRANTY Depending on the model two years from shipping date ESTIMATED LIFE SPAN 5 years
116 | Medical Electrical Equipment | EXAMINATIONLIGHT | Light, examination, mobile | Light, examination, mobile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Mobile examination light, with an LED light source and AC powered. Mobile examination LED light with articulating arm. INTENDED USE Examination lights directly illuminate specific areas of a patient's body during diagnostic procedures and general examination. These lights are intended to emit their useful radiation mainly in the visible spectrum. Examination lights may also be used during minor surgical procedures (e.g., minor cosmetic surgery). TECHNICAL SPECIFICATIONS Heavy base with low centre of gravity. Single light head with LED lights in cluster. Articulating arm between light and base stand. Base with 5 anti-static swivel castors in a star formation. At least two castors have been equipped with brakes. Minimum 14 cm fixed focus field of view. Minimum illumination of 60,000 lux at 0.5m. Minimum colour temperature between 4,000°K. Minimum colour rendering index (CRI) of 95. Minimum LED life more than 40,000 hours. Floor to light height adjustable to include at least range of 1.1 m to 1.75 m. On/off switch. Power requirements: 100 - 240 Volts - 50/60 Hz. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. OPTIONAL ACCESSORIES AVAILABLE - An integrated rechargeable battery with a minimum of 2 hours battery life can be supplied upon request at extra cost. ESTIMATED LIFE SPAN 10 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 75% RH. - Storage conditions: -10°C - 50°C / 75% RH. - Atmospheric pressure: 700 ~ 1060 hPa. - Ingress protection rating: IP20. WEIGHT AND VOLUME - estimated Weight: 13.4 kg Volume 80.000 dm³ Instructions for use: Medical Light to be proposed as basic equipment in health structures, it can be used for medical and gynaecological examination and minor operation. One of the advantages of a halogen examination lamp is its compactness. It can be placed in front of the operator without obstructing sight. Packaging & Labelling: One (1) examination light wrapped in a plastic film, with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13485 - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-41 Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. Environmental requirements: LED technology Low consumption Rosh compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Mobile examination light, with an LED light source and AC powered. Mobile examination LED light with articulating arm. INTENDED USE Examination lights directly illuminate specific areas of a patient's body during diagnostic procedures and general examination. These lights are intended to emit their useful radiation mainly in the visible spectrum. Examination lights may also be used during minor surgical procedures (e.g., minor cosmetic surgery). TECHNICAL SPECIFICATIONS Heavy base with low centre of gravity. Single light head with LED lights in cluster. Articulating arm between light and base stand. Base with 5 anti-static swivel castors in a star formation. At least two castors have been equipped with brakes. Minimum 14 cm fixed focus field of view. Minimum illumination of 60,000 lux at 0.5m. Minimum colour temperature between 4,000°K. Minimum colour rendering index (CRI) of 95. Minimum LED life more than 40,000 hours. Floor to light height adjustable to include at least range of 1.1 m to 1.75 m. On/off switch. Power requirements: 100 - 240 Volts - 50/60 Hz. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. OPTIONAL ACCESSORIES AVAILABLE - An integrated rechargeable battery with a minimum of 2 hours battery life can be supplied upon request at extra cost. ESTIMATED LIFE SPAN 10 years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 75% RH. - Storage conditions: -10°C - 50°C / 75% RH. - Atmospheric pressure: 700 ~ 1060 hPa. - Ingress protection rating: IP20. WEIGHT AND VOLUME - estimated Weight: 13.4 kg Volume 80.000 dm³ Instructions for use: Medical Light to be proposed as basic equipment in health structures, it can be used for medical and gynaecological examination and minor operation. One of the advantages of a halogen examination lamp is its compactness. It can be placed in front of the operator without obstructing sight. Packaging & Labelling: One (1) examination light wrapped in a plastic film, with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13485 - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-41 Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. Environmental requirements: LED technology Low consumption Rosh compliance
117 | Medical Electrical Equipment | FLASHLIGHT | Torch, flashlight, pre-focused, angled, medical | Torch, flashlight, pre-focused, angled, medical | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Torch/flashlight, pre-focused for medical use to look down the throat. Material: Heavy-walled plastic case or metal case. Right-angled head type Splash waterproof, sealed with o-ring. Lamp head approx. 40mm. Sturdy long-lasting mechanical switch. Light bulb 2.5 V Xenon. Batteries required, 2 x standard alkaline, dry cell batteries (type AA, 1.5V). Supplied with: - one spare light bulb (2.5V Xenon). - 2 x standard alkaline, dry cell batteries (type AA, 1.5V), separately. Instructions for use: Flashlight for medical use to look down the throat. For providing additional lighting during diagnosis examinations. Safety process: This item must be cleaned and disinfected after each use. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) flashlight in a box or case, with Manufacturer's instruction for use, spare parts and accessories (when applicable). Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: IEC 60601-1 IEC 60601-1-1 IEC 60601-1-2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Torch/flashlight, pre-focused for medical use to look down the throat. Material: Heavy-walled plastic case or metal case. Right-angled head type Splash waterproof, sealed with o-ring. Lamp head approx. 40mm. Sturdy long-lasting mechanical switch. Light bulb 2.5 V Xenon. Batteries required, 2 x standard alkaline, dry cell batteries (type AA, 1.5V). Supplied with: - one spare light bulb (2.5V Xenon). - 2 x standard alkaline, dry cell batteries (type AA, 1.5V), separately. Instructions for use: Flashlight for medical use to look down the throat. For providing additional lighting during diagnosis examinations. Safety process: This item must be cleaned and disinfected after each use. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) flashlight in a box or case, with Manufacturer's instruction for use, spare parts and accessories (when applicable). Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: IEC 60601-1 IEC 60601-1-1 IEC 60601-1-2
118 | Medical Electrical Equipment | LIGHTOPERATING | Light, operating, mobile | Light, operating, mobile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Surgical light, LED, mobile, AC and battery powered. Mobile surgical light, LED, with articulating arm and back up battery. INTENDED USE Surgical lights illuminate the surgical site for optimal visualization of small, low-contrast objects at varying depths in incisions and body cavities. Because hands, heads, and instruments can intrude over the surgical field, these lights are designed to reduce shadows and minimize colour distortion. They are also intended to operate for extended periods of time without radiating excessive heat, which could cause discomfort or dry tissues at the surgical site. TECHNICAL SPECIFICATIONS Single light head with LED lights in cluster. Adjustable field of view including at least 0.17 - 0.30m range (depending on the model supplied). Focal length of at least 0.7 m. Maximum illumination of at least 120,000 lux at 1 m. Colour temperature between 3000°K and 5000°K. Colour rendering index (CRI) of 95. Sterilizable and removable light handle for positioning. Minimum LED life of 50,000 hours. Minimum of 100° horizontal turning. Articulating arm between light and base stand. Base with at least four (4) anti-static swivel castors in a star formation. A minimum of two (2) castors are equipped with brakes. Heavy base with low centre of gravity, indicate weight of base relative to total device weight. Floor to light height adjustable to include at least range of 1.1 m to 2.12 m. On/off switch, battery status indicator (depending on model) and controls to adjust light intensity. Dimming function with a minimum of 5 settings or fully adjustable. Built in backup rechargeable battery with minimum 3 hours of battery life, battery type lead-acid. Automatic switch from mains to batteries in case of power failure. Power requirements: 100 - 240 Volts - 50/60 Hz. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 2 x Spare sterilizable handle. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 30% – 75% RH. - Storage conditions: -10°C - 50°C / 20% – 90% RH. - Atmospheric pressure: 700 ~ 7060 hPa. WEIGHT AND VOLUME - estimated Weight: 88.00 kg. Volume: 587.88 dm³. Instructions for use: For basic health infrastructures; in support of surgical interventions in operating theatres, delivery rooms, dressing rooms and emergency departments. After each intervention the handle should be disinfected, cleaned and sterilized by steam sterilizer. Never touch a halogen bulb with bare hands. Must be used and maintained properly, according manufacturer's instructions. Packaging & Labelling: One (1) operation light ,with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13845 - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-2-41 Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. Environmental requirements: LED technology Low consumption Rosh compliance | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Surgical light, LED, mobile, AC and battery powered. Mobile surgical light, LED, with articulating arm and back up battery. INTENDED USE Surgical lights illuminate the surgical site for optimal visualization of small, low-contrast objects at varying depths in incisions and body cavities. Because hands, heads, and instruments can intrude over the surgical field, these lights are designed to reduce shadows and minimize colour distortion. They are also intended to operate for extended periods of time without radiating excessive heat, which could cause discomfort or dry tissues at the surgical site. TECHNICAL SPECIFICATIONS Single light head with LED lights in cluster. Adjustable field of view including at least 0.17 - 0.30m range (depending on the model supplied). Focal length of at least 0.7 m. Maximum illumination of at least 120,000 lux at 1 m. Colour temperature between 3000°K and 5000°K. Colour rendering index (CRI) of 95. Sterilizable and removable light handle for positioning. Minimum LED life of 50,000 hours. Minimum of 100° horizontal turning. Articulating arm between light and base stand. Base with at least four (4) anti-static swivel castors in a star formation. A minimum of two (2) castors are equipped with brakes. Heavy base with low centre of gravity, indicate weight of base relative to total device weight. Floor to light height adjustable to include at least range of 1.1 m to 2.12 m. On/off switch, battery status indicator (depending on model) and controls to adjust light intensity. Dimming function with a minimum of 5 settings or fully adjustable. Built in backup rechargeable battery with minimum 3 hours of battery life, battery type lead-acid. Automatic switch from mains to batteries in case of power failure. Power requirements: 100 - 240 Volts - 50/60 Hz. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 2 x Spare sterilizable handle. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years from shipping date. ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 30°C / 30% – 75% RH. - Storage conditions: -10°C - 50°C / 20% – 90% RH. - Atmospheric pressure: 700 ~ 7060 hPa. WEIGHT AND VOLUME - estimated Weight: 88.00 kg. Volume: 587.88 dm³. Instructions for use: For basic health infrastructures; in support of surgical interventions in operating theatres, delivery rooms, dressing rooms and emergency departments. After each intervention the handle should be disinfected, cleaned and sterilized by steam sterilizer. Never touch a halogen bulb with bare hands. Must be used and maintained properly, according manufacturer's instructions. Packaging & Labelling: One (1) operation light ,with manufacturer's instruction for use, spare parts and accessories. Symbols used according ISO 15223 Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Labelling displays, at least: Product name, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13845 - IEC 60601-1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-2-41 Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. Environmental requirements: LED technology Low consumption Rosh compliance
119 | Medical Electrical Equipment | New | Monitor, patient, portable,w/access | Monitor, patient, portable,w/access | GENERAL DESCRIPTION Patient vital signs monitor, 6 parameters, suitable for adults, paediatric/infants and newborn patients. Bedside patient monitor for monitoring of vital signs of all patient categories. INTENDED USE Vital signs monitors are used to measure basic physiologic parameters and track the status of low-acuity patients (patients found in medical-surgical areas, emergency departments, and ambulatory care centres) so that clinicians can be informed of changes in a patient's condition. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Measures and stores patient status for the following: EGC, with ST Segment Analysis Heart rate (HR) Respiration rate (RR) Invasive blood pressure (IBP) Non-invasive blood pressure (NIBP) Arrhythmia Pulse rate Temperature Cardiac output (Central) apnea SpO2 With colour display monitor. At least 3 channel Digital display of all active parameters and trace display of at least three selectable parameters Unwanted parameters can be deselected from display MEASUREMENT, RANGES ECG leads I, II, III HR at least 30 to 250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm RR at least 0 to 100 bpm, minimum gradation 1 bpm Blood pressure monitoring at least 30 to 300 mmHg, minimum gradation 1 mmHg Temperature range at least 30 to 40 deg C, minimum gradation 0.1 deg C SpO2 at least 70 to 99 %, with accuracy better than ± 3% and minimum gradation 1%" "NIBP oscillometric step deflation, manual/automatic, initial inflation pressure user selectable. Sensitivity (amplitude) of all signals user adjustable Standardising marker, 1mV Temperature probe to be reusable, external skin contact type Data interface (for ECG) RS232, BNC, USB or equivalent Viewing monitor: minimum 10-inch Adjustable signal amplitude and sensitivity. Multi-waveform and parameters visualization, up to 7 wave forms simultaneously. Defibrillator sync and protection. Patient information and trend internal database. Ability to save a minimum of 72 events. Trend storage ≥ 24 hours. Suitable for standard bed/wall rail and pole stand mount. Robust design for use in demanding environment. Designed for frequent and easy dismount and disinfection with hospital-grade products. Automatic self-test and continuous system monitoring. Built-in rechargeable battery, autonomy of at least 2 hours. Automatic switch to battery in case of power failure, automatic recharge on connection to mains. Main cable at least 3 meter long. Electrical protection provided by fuses in both live and neutral supply lines. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). ELECTROCARDIOGRAM (ECG) AND HEART RATE (HR) The ECG is derived from a minimum of 5-leads I, II, III, aVR, aVL, aVF, V. The unit allows for a simultaneously display of a minimum of 2 ECG traces. The accuracy of the HR is equal or better than ± 1 bpm or 1%. The resolution of the HR is equal or better than: 1 bpm. The unit provides for S-T analyses on ECGs. The unit provide for arrhythmia analysis on the ECGs. The unit is equipped with a standardising marker at 1 mV. The adult HR measurement range is equal to or better than: 15 – 300 bpm. The paediatrics/neonates HR measurement range is equal to or better than: 30 – 300 bpm. RESPIRATION RATE (RR) The adult RR measurement range is equal to or better than: 6 – 120 bpm. The paediatric/neonatal RR measurement range is equal to or better than: 6 – 150 bpm. 68.1 Select your answer from the below drop-down menu. 68.2 Indicate here the actual specification of the offered unit. The resolution of the RR is equal to or better then: 1 bpm. NON-INVASIVE BLOOD PRESSURE (NIBP) The NIBP is measured through the oscillometric step deflation method. The unit supports manual and automatic modes. The unit allows for adjustable inflation pressure. The NIBP adult range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 270 mmHg. The NIBP paediatric NIBP measurement range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 200 mmHg. The BP neonatal measurement range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 135 mmHg. The resolution of the NIBP is equal to or better than: 1 mmHg or better. The minimum average error for the NIBP is: ±5 mmHg (0.7 kPa), Standard deviation: ≥ 8 mmHg (1.1 kPa). OXYGEN SATURATION (SPO₂) The SpO₂ measurement range is equal to or better than: 1 – 100%. The resolution of the SpO₂ is equal to or better than: 1% or better. The accuracy of the SpO₂ is equal to or better than: ± 3% between 70% - 100%. The HR detection range derived from the SpO₂ is equal to or better than: 25 – 250 bpm. The accuracy of the HR detection is equal to or better than: ± 2 bpm or ± 2%, whichever is greater (static). TEMPERATURE (T) The temperature measurement range is equal to or better than: 0 - 50°C. The temperature resolution is equal to or better than: 0.1°C or better. The temperature accuracy is equal to or better than: ± 0.1°C. ALARM FUNCTIONIONALITIES The unit provide audio-visual alarms for all monitored parameters The unit allows for user pre-set of high and low alarms for all monitored parameters for each different patient category The alarm has override and temporary silence functions. The unit generates alarms for leads-off or sensor disconnect or sensor failure. The unit is provided with an apnoea alarm. The unit is equipped with an alarm for AC status and low battery. Memory storage: minimum 72 hours Event management save feature Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO." SUPPLIED WITH 1 x Plastic protective dustcover. 1 x wall mount bracket. 1 x spare rechargeable battery pack. 1 x Set of spare fuses NIBP accessories: 3 x NIBP hose (1 x neonate, 1 x infant, 1 x adult) 3 x Blood pressure cuff (1 x infant, 1 x child, 1 x adult) ECG accessories: 2 x Patient cable extremities (1x neonate/paediatric, 1 x adult) 2 x Set of electrodes (1x neonate/paediatric, 1 x adult) Temperature accessories: 2 x Skin temperature probes (incl. connection cable) Pulse oximetry (SpO2) sensors with cable and plug 2 x Adult size, reusable clip-on type 2 x Infant size, reusable clip-on type 3 x New-born size, reusable clip-on type 10 x New-born size, single-use wrap-around type 1 x Instructions for use ; instruction manual including instructions for assembly, use and maintenance, and list of accessories and spare parts, in 3 languages (English, French and Spanish)" INSTRUCTION FOR USE For monitoring vital signs of all patient categories, at bedside Must be used and maintained properly, according to manufacturer's instructions.." ACCESSORIES/ SPARE PARTS / CONSUMABLES NIBP accessories, ECG accessories , SpO2 accessories, temperature accessories, spare fuses (if non-resettable fuses used)" PACKAGING & LABELLING Unit presentation 1 (one) Monitor unit, with accessories and Instructions for use. Symbols used according ISO 15223 CE mark and reference number of notifying body. WEIGHT AND VOLUME Weight: 10.7 kg (without battery). Volume: 92.63 dm³ - estimated REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS ISO 13845:2003 IEC 60601-1:2005; IEC 60601-1-1:2000 ; IEC 60601-1-2:2007; IEC 60601-2-49:2011; IEC 60601-1-8 :2012" "ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 15% – 85% RH. - Storage conditions: -20°C - 55°C / 15% – 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliant Solar Refrigerator/freezer WARRANTY Two years from shipping date. | GENERAL DESCRIPTION Patient vital signs monitor, 6 parameters, suitable for adults, paediatric/infants and newborn patients. Bedside patient monitor for monitoring of vital signs of all patient categories. INTENDED USE Vital signs monitors are used to measure basic physiologic parameters and track the status of low-acuity patients (patients found in medical-surgical areas, emergency departments, and ambulatory care centres) so that clinicians can be informed of changes in a patient's condition. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Measures and stores patient status for the following: EGC, with ST Segment Analysis Heart rate (HR) Respiration rate (RR) Invasive blood pressure (IBP) Non-invasive blood pressure (NIBP) Arrhythmia Pulse rate Temperature Cardiac output (Central) apnea SpO2 With colour display monitor. At least 3 channel Digital display of all active parameters and trace display of at least three selectable parameters Unwanted parameters can be deselected from display MEASUREMENT, RANGES ECG leads I, II, III HR at least 30 to 250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm RR at least 0 to 100 bpm, minimum gradation 1 bpm Blood pressure monitoring at least 30 to 300 mmHg, minimum gradation 1 mmHg Temperature range at least 30 to 40 deg C, minimum gradation 0.1 deg C SpO2 at least 70 to 99 %, with accuracy better than ± 3% and minimum gradation 1%" "NIBP oscillometric step deflation, manual/automatic, initial inflation pressure user selectable. Sensitivity (amplitude) of all signals user adjustable Standardising marker, 1mV Temperature probe to be reusable, external skin contact type Data interface (for ECG) RS232, BNC, USB or equivalent Viewing monitor: minimum 10-inch Adjustable signal amplitude and sensitivity. Multi-waveform and parameters visualization, up to 7 wave forms simultaneously. Defibrillator sync and protection. Patient information and trend internal database. Ability to save a minimum of 72 events. Trend storage ≥ 24 hours. Suitable for standard bed/wall rail and pole stand mount. Robust design for use in demanding environment. Designed for frequent and easy dismount and disinfection with hospital-grade products. Automatic self-test and continuous system monitoring. Built-in rechargeable battery, autonomy of at least 2 hours. Automatic switch to battery in case of power failure, automatic recharge on connection to mains. Main cable at least 3 meter long. Electrical protection provided by fuses in both live and neutral supply lines. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit). ELECTROCARDIOGRAM (ECG) AND HEART RATE (HR) The ECG is derived from a minimum of 5-leads I, II, III, aVR, aVL, aVF, V. The unit allows for a simultaneously display of a minimum of 2 ECG traces. The accuracy of the HR is equal or better than ± 1 bpm or 1%. The resolution of the HR is equal or better than: 1 bpm. The unit provides for S-T analyses on ECGs. The unit provide for arrhythmia analysis on the ECGs. The unit is equipped with a standardising marker at 1 mV. The adult HR measurement range is equal to or better than: 15 – 300 bpm. The paediatrics/neonates HR measurement range is equal to or better than: 30 – 300 bpm. RESPIRATION RATE (RR) The adult RR measurement range is equal to or better than: 6 – 120 bpm. The paediatric/neonatal RR measurement range is equal to or better than: 6 – 150 bpm. 68.1 Select your answer from the below drop-down menu. 68.2 Indicate here the actual specification of the offered unit. The resolution of the RR is equal to or better then: 1 bpm. NON-INVASIVE BLOOD PRESSURE (NIBP) The NIBP is measured through the oscillometric step deflation method. The unit supports manual and automatic modes. The unit allows for adjustable inflation pressure. The NIBP adult range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 270 mmHg. The NIBP paediatric NIBP measurement range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 200 mmHg. The BP neonatal measurement range in minimum diastolic and maximum systolic values should be equal or better than: 10 – 135 mmHg. The resolution of the NIBP is equal to or better than: 1 mmHg or better. The minimum average error for the NIBP is: ±5 mmHg (0.7 kPa), Standard deviation: ≥ 8 mmHg (1.1 kPa). OXYGEN SATURATION (SPO₂) The SpO₂ measurement range is equal to or better than: 1 – 100%. The resolution of the SpO₂ is equal to or better than: 1% or better. The accuracy of the SpO₂ is equal to or better than: ± 3% between 70% - 100%. The HR detection range derived from the SpO₂ is equal to or better than: 25 – 250 bpm. The accuracy of the HR detection is equal to or better than: ± 2 bpm or ± 2%, whichever is greater (static). TEMPERATURE (T) The temperature measurement range is equal to or better than: 0 - 50°C. The temperature resolution is equal to or better than: 0.1°C or better. The temperature accuracy is equal to or better than: ± 0.1°C. ALARM FUNCTIONIONALITIES The unit provide audio-visual alarms for all monitored parameters The unit allows for user pre-set of high and low alarms for all monitored parameters for each different patient category The alarm has override and temporary silence functions. The unit generates alarms for leads-off or sensor disconnect or sensor failure. The unit is provided with an apnoea alarm. The unit is equipped with an alarm for AC status and low battery. Memory storage: minimum 72 hours Event management save feature Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO." SUPPLIED WITH 1 x Plastic protective dustcover. 1 x wall mount bracket. 1 x spare rechargeable battery pack. 1 x Set of spare fuses NIBP accessories: 3 x NIBP hose (1 x neonate, 1 x infant, 1 x adult) 3 x Blood pressure cuff (1 x infant, 1 x child, 1 x adult) ECG accessories: 2 x Patient cable extremities (1x neonate/paediatric, 1 x adult) 2 x Set of electrodes (1x neonate/paediatric, 1 x adult) Temperature accessories: 2 x Skin temperature probes (incl. connection cable) Pulse oximetry (SpO2) sensors with cable and plug 2 x Adult size, reusable clip-on type 2 x Infant size, reusable clip-on type 3 x New-born size, reusable clip-on type 10 x New-born size, single-use wrap-around type 1 x Instructions for use ; instruction manual including instructions for assembly, use and maintenance, and list of accessories and spare parts, in 3 languages (English, French and Spanish)" INSTRUCTION FOR USE For monitoring vital signs of all patient categories, at bedside Must be used and maintained properly, according to manufacturer's instructions.." ACCESSORIES/ SPARE PARTS / CONSUMABLES NIBP accessories, ECG accessories , SpO2 accessories, temperature accessories, spare fuses (if non-resettable fuses used)" PACKAGING & LABELLING Unit presentation 1 (one) Monitor unit, with accessories and Instructions for use. Symbols used according ISO 15223 CE mark and reference number of notifying body. WEIGHT AND VOLUME Weight: 10.7 kg (without battery). Volume: 92.63 dm³ - estimated REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device SAFETY & PRODUCT STANDARDS ISO 13845:2003 IEC 60601-1:2005; IEC 60601-1-1:2000 ; IEC 60601-1-2:2007; IEC 60601-2-49:2011; IEC 60601-1-8 :2012" "ENVIRONMENTAL CONDITIONS - Operating conditions: 10°C - 40°C / 15% – 85% RH. - Storage conditions: -20°C - 55°C / 15% – 90% RH. - Atmospheric pressure: 700 ~ 1060 hPa. ENVIRONMENTAL REQUIREMENTS Rohs compliant Solar Refrigerator/freezer WARRANTY Two years from shipping date.
120 | Medical Electrical Equipment | REFRIGERATORC | Refrigerator, compression | Refrigerator, compression | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Refrigerator/freezer for vaccine storage and laboratory use. Compression cycle with one or two compressors Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. gross capacity: 110-120 l (or similar size) Vaccine storage capacity up to 17 litres Stainless steel structure CFC, HCFC and ammonia free. Thermostat Internal circulating fan to maintain temperature Door lock with key(s) Alarm for over/under temperature Temperature range 2°C to + 8°C Accuracy, whatever the load: +/-1C Alarm for open door Easily adjustable shelves of stainless steel and lacquered wire or glass Insulation material: polyurethane, CFC-free Level indicator Adjustable levelling feet. Temperature monitoring External digital temperature display Suitable for tropical use Electrical safety rating 4 and 5 only (see WHO equipment performance specifications E3) Meets WHO/UNICEF standard: E3/RF.1 Power consumption: approximately 250 w Supplied with: 1 x Instructions for use : instruction manual including Instructions for use, preventive maintenance and troubleshooting in 3 languages (English, French and Spanish) and spare list. Instructions for use: Used for vaccines storage and laboratory. Must be used and maintained properly, according manufacturer's instructions. Packaging & Labelling: Primary packaging: Unit presentation: 1(one) Refrigerator, 110L Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not Applicable Safety & product Standards: Must comply with the following standards: ISO 9001 IEC 61010-1 IEC 61236-1 ISO8187/EN28187 Environmental requirements: Rohs compliance Solar Refrigerator/freezer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Refrigerator/freezer for vaccine storage and laboratory use. Compression cycle with one or two compressors Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. gross capacity: 110-120 l (or similar size) Vaccine storage capacity up to 17 litres Stainless steel structure CFC, HCFC and ammonia free. Thermostat Internal circulating fan to maintain temperature Door lock with key(s) Alarm for over/under temperature Temperature range 2°C to + 8°C Accuracy, whatever the load: +/-1C Alarm for open door Easily adjustable shelves of stainless steel and lacquered wire or glass Insulation material: polyurethane, CFC-free Level indicator Adjustable levelling feet. Temperature monitoring External digital temperature display Suitable for tropical use Electrical safety rating 4 and 5 only (see WHO equipment performance specifications E3) Meets WHO/UNICEF standard: E3/RF.1 Power consumption: approximately 250 w Supplied with: 1 x Instructions for use : instruction manual including Instructions for use, preventive maintenance and troubleshooting in 3 languages (English, French and Spanish) and spare list. Instructions for use: Used for vaccines storage and laboratory. Must be used and maintained properly, according manufacturer's instructions. Packaging & Labelling: Primary packaging: Unit presentation: 1(one) Refrigerator, 110L Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not Applicable Safety & product Standards: Must comply with the following standards: ISO 9001 IEC 61010-1 IEC 61236-1 ISO8187/EN28187 Environmental requirements: Rohs compliance Solar Refrigerator/freezer
121 | Medical Electrical Equipment | New | Warmer, baby, electric | Warmer, baby, electric | GENERAL DESCRIPTION Radiant warmer, newborn, with accessories. INTENDED USE Infant radiant warmers are commonly used to provide thermal support for newborns in the delivery suite, for critically ill infants who require constant nursing intervention, and for infants undergoing treatment that prolongs exposure to a cool environment. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS The Warmer is used to assist premature babies maintaining the appropriate body temperature. System: control unit, a heating pad, a mattress and an inflatable frame Self-diagnosis with each start-up Integrity testing of all system parameters every 5 minute Flexible waterproof low voltage heating pad slides under mattress Electronic servo temperature control Pre-set temperature range of mattress 35 to 38C, increments 0.1C Accuracy temperature control +/-0.1C Automatic warm-up to 37C, time to stabilize approximately 20 minutes (starting at 20C) Large colour display shows operational status, with set and measured values System reports, with audio-visual alerts on high/low temperature, malfunctions (temperature sensor, connecting wires), water level, power failure, built-in back-up battery Strong soft plastic mattress fills with water, approximately 4.5L Reusable inflatable frame keeps baby in place Cotton cover frame can be tightly secured around the water mattress and heating pad Fit with connector for standard YSI skin temperature sensor Power requirements 110/230V, 50/60Hz, or battery 12 or 24V Power consumption, approximately 20W (depending on ambient temperature) Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO." SUPPLIED WITH 1 x Control unit with fixture / handle for attachment to bassinet or cradle 1 x Heating pad, fit with connection cable, length approx. 1.3 m 1 x Water mattress, 2 x Inflatable frame/mattress cover (‘nest’) 1 x Anti-algae, bottle of 100ml, pack of 10; alternatively 1 x Clean water tablets, bottle of 10 tablets 1 x Power cord main supply. Length approx. 1 m 1 x Instructions for Use: Manual with instructions for assembling, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish)" INSTRUCTION FOR USE Temperature controlled open warmer system to help newborns maintain their body temperature. The item must be operated by trained staff only. It is recommended to strictly follow manufacturer's instructions. If the voltage of the power supply is expected to fluctuate, the use of a voltage stabiliser / surge protector is recommended. Control unit can connect to external battery 12 or 24V, To prevent the formation of algae and bacteria inside the water mattress, decontaminating solution / tablet should be added as per the intsructions supplied, or each time the water is changed or topped-up. It keeps the closed water mattress free from algae and bacteria for 12 months." ACCESSORIES/ SPARE PARTS / CONSUMABLES 1 power cord main supply. 1 heating pad, fit with connection cable 2 x Inflatable frame/mattress cover 1 x Water mattress, 1 x Anti-algae, bottle of 100ml, pack of 10 / or 1 clean water tablet, bottle of 10 tablets" PACKAGING & LABELLING One (1) Warmer system,newborn, basic,set and instructions for use CE mark and reference number of notifying body. REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device ALARM AND SAFETY FEATURES : Alarms are audible as well as visual. Alarms indicating: - Sensor disconnection or malfunction. - Power failure. - Skin temperature, after stabilizing, varying beyond 0.5° to 1°C from set-point. - 10 to 15 minutes after manual mode has been engaged an audible alarm will sound. In addition, the unit will either reduce or cut the power to the heater. In skin mode, when the temperature exceeds 39°C, the unit will cut the power to the heater." SAFETY & PRODUCT STANDARDS Must comply with the following standards: ISO 13845:2003 IEC 60601-1:2005; IEC 60601-1-1:2000; IEC 60601-1-2:2007; IEC 60601-2-21:2009" ENVIRONMENTAL REQUIREMENTS Rohs compliant ENVIRONMENTAL CONDITIONS - Operating conditions: 18°C - 30 °C / 30% – 75% RH. - Storage conditions: -40°C - 55 °C / < 93% RH. - Atmospheric pressure: 700 ~ 1060hPa. WEIGHT AND VOLUME - estimated Weight: 160.00 kg. Volume: 1,311.00 dm³. INSTALLATION REQUIREMENTS Assembly and commissioning this product should be carried out by qualified technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years from shipping date." | GENERAL DESCRIPTION Radiant warmer, newborn, with accessories. INTENDED USE Infant radiant warmers are commonly used to provide thermal support for newborns in the delivery suite, for critically ill infants who require constant nursing intervention, and for infants undergoing treatment that prolongs exposure to a cool environment. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS The Warmer is used to assist premature babies maintaining the appropriate body temperature. System: control unit, a heating pad, a mattress and an inflatable frame Self-diagnosis with each start-up Integrity testing of all system parameters every 5 minute Flexible waterproof low voltage heating pad slides under mattress Electronic servo temperature control Pre-set temperature range of mattress 35 to 38C, increments 0.1C Accuracy temperature control +/-0.1C Automatic warm-up to 37C, time to stabilize approximately 20 minutes (starting at 20C) Large colour display shows operational status, with set and measured values System reports, with audio-visual alerts on high/low temperature, malfunctions (temperature sensor, connecting wires), water level, power failure, built-in back-up battery Strong soft plastic mattress fills with water, approximately 4.5L Reusable inflatable frame keeps baby in place Cotton cover frame can be tightly secured around the water mattress and heating pad Fit with connector for standard YSI skin temperature sensor Power requirements 110/230V, 50/60Hz, or battery 12 or 24V Power consumption, approximately 20W (depending on ambient temperature) Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO." SUPPLIED WITH 1 x Control unit with fixture / handle for attachment to bassinet or cradle 1 x Heating pad, fit with connection cable, length approx. 1.3 m 1 x Water mattress, 2 x Inflatable frame/mattress cover (‘nest’) 1 x Anti-algae, bottle of 100ml, pack of 10; alternatively 1 x Clean water tablets, bottle of 10 tablets 1 x Power cord main supply. Length approx. 1 m 1 x Instructions for Use: Manual with instructions for assembling, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish)" INSTRUCTION FOR USE Temperature controlled open warmer system to help newborns maintain their body temperature. The item must be operated by trained staff only. It is recommended to strictly follow manufacturer's instructions. If the voltage of the power supply is expected to fluctuate, the use of a voltage stabiliser / surge protector is recommended. Control unit can connect to external battery 12 or 24V, To prevent the formation of algae and bacteria inside the water mattress, decontaminating solution / tablet should be added as per the intsructions supplied, or each time the water is changed or topped-up. It keeps the closed water mattress free from algae and bacteria for 12 months." ACCESSORIES/ SPARE PARTS / CONSUMABLES 1 power cord main supply. 1 heating pad, fit with connection cable 2 x Inflatable frame/mattress cover 1 x Water mattress, 1 x Anti-algae, bottle of 100ml, pack of 10 / or 1 clean water tablet, bottle of 10 tablets" PACKAGING & LABELLING One (1) Warmer system,newborn, basic,set and instructions for use CE mark and reference number of notifying body. REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIb device ALARM AND SAFETY FEATURES : Alarms are audible as well as visual. Alarms indicating: - Sensor disconnection or malfunction. - Power failure. - Skin temperature, after stabilizing, varying beyond 0.5° to 1°C from set-point. - 10 to 15 minutes after manual mode has been engaged an audible alarm will sound. In addition, the unit will either reduce or cut the power to the heater. In skin mode, when the temperature exceeds 39°C, the unit will cut the power to the heater." SAFETY & PRODUCT STANDARDS Must comply with the following standards: ISO 13845:2003 IEC 60601-1:2005; IEC 60601-1-1:2000; IEC 60601-1-2:2007; IEC 60601-2-21:2009" ENVIRONMENTAL REQUIREMENTS Rohs compliant ENVIRONMENTAL CONDITIONS - Operating conditions: 18°C - 30 °C / 30% – 75% RH. - Storage conditions: -40°C - 55 °C / < 93% RH. - Atmospheric pressure: 700 ~ 1060hPa. WEIGHT AND VOLUME - estimated Weight: 160.00 kg. Volume: 1,311.00 dm³. INSTALLATION REQUIREMENTS Assembly and commissioning this product should be carried out by qualified technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. ESTIMATED LIFE SPAN Eight years. WARRANTY Two years from shipping date."
122 | Medical Sterilization Equipment | INDICATOR.TST_1 | Indicator, TST control, spot, adhesive, P1 | Indicator, TST control, spot, adhesive, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Indicator TST (Temperature, Steam, Time) control spot for steam sterilization process, type (class) 6 indicator, pack of 300. INTENDED USE A sterilization indicator designed to be placed within a sterilizing chamber, and to respond with a characteristic chemical and/or physical change to one or more of the physical conditions within the chamber, to validate the sterilization process. This is a single-use device. TECHNICAL SPECIFICATIONS Steam cycle verification indicator; TST indicators (Time, Steam and Temperature). Model: Type (class) 6 Emulating Indicator (cycle verification). The indicator shall be marked with the type (class) of process it is intended for. The indicator shall be marked with "STEAM". The control spot should irreversibly change colour after exposure to saturated steam with a temperature of 121° Celsius for a minimum period of 15 minutes. The indicator shall indicate the time required (15 minutes) for discoloration at 121° Celsius. The indicator shall have an adhesive back. Suitable for use in portable steam sterilizers of the pressure cooker type, 24 and 39 Litres. SUPPLIED WITH Instructions for use in English, French and Spanish. 1 x record sheet. SHELF LIFE 4 years ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 30°C / 30% - 60% RH. Operating conditions: 0 - 30°C / 30% - 60% RH. Store in a cool dark place, away from direct sunlight. Check before use. Do not use TST spot if the colour is not YELLOW. Monitor shelf life (batch (lot) number and expired date) Important: It is recommended to follow Manufacturer's instructions for use. WEIGHT AND VOLUME (Packaged) - estimated Weight: 45 gr. Volume: 0.1 dm³. Instructions for use: The use of TST control spot gives an immediate indication to the person in charge on successful sterilisation cycle or not. TST control spot must be used systematically for each steam sterilization cycle. TST control spot must be attached to the lid of drum (or other steam sterilizing container), TST control spot must be checked when the sterilizing cycle is finished. TST control spots are delivered with 1 record sheet to retain 300 TST control spots and to be used to record the results of all sterilizing cycles, pass or fail, remedial action taken to remedy failed cycles. If the TST control spot has changed from YELLOW to BLUE, items in the drum (or other steam sterilizing container) can be used. If the TST control spot has NOT changed to BLUE, items in the drum (or other steam sterilizing container) MUST NOT BE USED, they must be re-sterilized, using a new TST control spot. Packaging & Labelling: Primary packaging: Unit of use. One (1) plastic pack of 300 TST indicator spots. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). CE mark Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit. Fifty (50) packs in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008 ISO 11140-1:2005 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Indicator TST (Temperature, Steam, Time) control spot for steam sterilization process, type (class) 6 indicator, pack of 300. INTENDED USE A sterilization indicator designed to be placed within a sterilizing chamber, and to respond with a characteristic chemical and/or physical change to one or more of the physical conditions within the chamber, to validate the sterilization process. This is a single-use device. TECHNICAL SPECIFICATIONS Steam cycle verification indicator; TST indicators (Time, Steam and Temperature). Model: Type (class) 6 Emulating Indicator (cycle verification). The indicator shall be marked with the type (class) of process it is intended for. The indicator shall be marked with "STEAM". The control spot should irreversibly change colour after exposure to saturated steam with a temperature of 121° Celsius for a minimum period of 15 minutes. The indicator shall indicate the time required (15 minutes) for discoloration at 121° Celsius. The indicator shall have an adhesive back. Suitable for use in portable steam sterilizers of the pressure cooker type, 24 and 39 Litres. SUPPLIED WITH Instructions for use in English, French and Spanish. 1 x record sheet. SHELF LIFE 4 years ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 30°C / 30% - 60% RH. Operating conditions: 0 - 30°C / 30% - 60% RH. Store in a cool dark place, away from direct sunlight. Check before use. Do not use TST spot if the colour is not YELLOW. Monitor shelf life (batch (lot) number and expired date) Important: It is recommended to follow Manufacturer's instructions for use. WEIGHT AND VOLUME (Packaged) - estimated Weight: 45 gr. Volume: 0.1 dm³. Instructions for use: The use of TST control spot gives an immediate indication to the person in charge on successful sterilisation cycle or not. TST control spot must be used systematically for each steam sterilization cycle. TST control spot must be attached to the lid of drum (or other steam sterilizing container), TST control spot must be checked when the sterilizing cycle is finished. TST control spots are delivered with 1 record sheet to retain 300 TST control spots and to be used to record the results of all sterilizing cycles, pass or fail, remedial action taken to remedy failed cycles. If the TST control spot has changed from YELLOW to BLUE, items in the drum (or other steam sterilizing container) can be used. If the TST control spot has NOT changed to BLUE, items in the drum (or other steam sterilizing container) MUST NOT BE USED, they must be re-sterilized, using a new TST control spot. Packaging & Labelling: Primary packaging: Unit of use. One (1) plastic pack of 300 TST indicator spots. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). CE mark Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit. Fifty (50) packs in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008 ISO 11140-1:2005
123 | Medical Sterilization Equipment | INDICATOR.TST_1050 | Indicator, TST control, spot, adhesive, P1050 | Indicator, TST control, spot, adhesive, P1050 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Indicator TST (Temperature, Steam, Time) control spot for steam sterilization process, type (class) 6 indicator, pack of 300. INTENDED USE A sterilization indicator designed to be placed within a sterilizing chamber, and to respond with a characteristic chemical and/or physical change to one or more of the physical conditions within the chamber, to validate the sterilization process. This is a single-use device. TECHNICAL SPECIFICATIONS Steam cycle verification indicator; TST indicators (Time, Steam and Temperature). Model: Type (class) 6 Emulating Indicator (cycle verification). The indicator shall be marked with the type (class) of process it is intended for. The indicator shall be marked with "STEAM". The control spot should irreversibly change colour after exposure to saturated steam with a temperature of 121° Celsius for a minimum period of 15 minutes. The indicator shall indicate the time required (15 minutes) for discoloration at 121° Celsius. The indicator shall have an adhesive back. Suitable for use in portable steam sterilizers of the pressure cooker type, 24 and 39 Litres. SUPPLIED WITH Instructions for use in English, French and Spanish. 1 x record sheet. SHELF LIFE 4 years ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 30°C / 30% - 60% RH. Operating conditions: 0 - 30°C / 30% - 60% RH. Store in a cool dark place, away from direct sunlight. Check before use. Do not use TST spot if the colour is not YELLOW. Monitor shelf life (batch (lot) number and expired date) Important: It is recommended to follow Manufacturer's instructions for use. WEIGHT AND VOLUME (Packaged) - estimated Weight: 45 gr. Volume: 0.1 dm³. Instructions for use: The use of TST control spot gives an immediate indication to the person in charge on successful sterilisation cycle or not. TST control spot must be used systematically for each steam sterilization cycle. TST control spot must be attached to the lid of drum (or other steam sterilizing container), TST control spot must be checked when the sterilizing cycle is finished. TST control spots are delivered with 1 record sheet to retain 300 TST control spots and to be used to record the results of all sterilizing cycles, pass or fail, remedial action taken to remedy failed cycles. If the TST control spot has changed from YELLOW to BLUE, items in the drum (or other steam sterilizing container) can be used. If the TST control spot has NOT changed to BLUE, items in the drum (or other steam sterilizing container) MUST NOT BE USED, they must be re-sterilized, using a new TST control spot. Packaging & Labelling: Primary packaging: Unit of use. One (1) plastic pack of 300 TST indicator spots. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). CE mark Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit. Fifty (50) packs in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008 ISO 11140-1:2005 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Indicator TST (Temperature, Steam, Time) control spot for steam sterilization process, type (class) 6 indicator, pack of 300. INTENDED USE A sterilization indicator designed to be placed within a sterilizing chamber, and to respond with a characteristic chemical and/or physical change to one or more of the physical conditions within the chamber, to validate the sterilization process. This is a single-use device. TECHNICAL SPECIFICATIONS Steam cycle verification indicator; TST indicators (Time, Steam and Temperature). Model: Type (class) 6 Emulating Indicator (cycle verification). The indicator shall be marked with the type (class) of process it is intended for. The indicator shall be marked with "STEAM". The control spot should irreversibly change colour after exposure to saturated steam with a temperature of 121° Celsius for a minimum period of 15 minutes. The indicator shall indicate the time required (15 minutes) for discoloration at 121° Celsius. The indicator shall have an adhesive back. Suitable for use in portable steam sterilizers of the pressure cooker type, 24 and 39 Litres. SUPPLIED WITH Instructions for use in English, French and Spanish. 1 x record sheet. SHELF LIFE 4 years ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 30°C / 30% - 60% RH. Operating conditions: 0 - 30°C / 30% - 60% RH. Store in a cool dark place, away from direct sunlight. Check before use. Do not use TST spot if the colour is not YELLOW. Monitor shelf life (batch (lot) number and expired date) Important: It is recommended to follow Manufacturer's instructions for use. WEIGHT AND VOLUME (Packaged) - estimated Weight: 45 gr. Volume: 0.1 dm³. Instructions for use: The use of TST control spot gives an immediate indication to the person in charge on successful sterilisation cycle or not. TST control spot must be used systematically for each steam sterilization cycle. TST control spot must be attached to the lid of drum (or other steam sterilizing container), TST control spot must be checked when the sterilizing cycle is finished. TST control spots are delivered with 1 record sheet to retain 300 TST control spots and to be used to record the results of all sterilizing cycles, pass or fail, remedial action taken to remedy failed cycles. If the TST control spot has changed from YELLOW to BLUE, items in the drum (or other steam sterilizing container) can be used. If the TST control spot has NOT changed to BLUE, items in the drum (or other steam sterilizing container) MUST NOT BE USED, they must be re-sterilized, using a new TST control spot. Packaging & Labelling: Primary packaging: Unit of use. One (1) plastic pack of 300 TST indicator spots. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). CE mark Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit. Fifty (50) packs in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008 ISO 11140-1:2005
124 | Medical Sterilization Equipment | STERILIZEBASKET | Basket, sterilizing, approx. 120 x 250 x 60 mm | Basket, sterilizing, approx. 120 x 250 x 60 mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Stainless steel wired basket used as support for sterilizing surgical instruments. Material : austenitic stainless steel (non-quenched, non-magnetic steel).AISI 316 L stainless steel Rectangular basket, rounded not soldered corners. Close mesh netting (4.5x4.5mm). Tilting handles. Dimensions: approx. 120 x 250 x 60 mm. Instructions for use: Stainless steel wired basket used as support for sterilizing surgical instruments. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) sterilization basket in a protective plastic bag with manufacturer's instruction for user in English. Labelling: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels — Part _: List of stainless steels ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Stainless steel wired basket used as support for sterilizing surgical instruments. Material : austenitic stainless steel (non-quenched, non-magnetic steel).AISI 316 L stainless steel Rectangular basket, rounded not soldered corners. Close mesh netting (4.5x4.5mm). Tilting handles. Dimensions: approx. 120 x 250 x 60 mm. Instructions for use: Stainless steel wired basket used as support for sterilizing surgical instruments. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) sterilization basket in a protective plastic bag with manufacturer's instruction for user in English. Labelling: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels — Part _: List of stainless steels ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: Not use PVC polymer
125 | Medical Sterilization Equipment | STERILIZEBASKETL | Basket, instruments, for sterilization, wired, 400 x 200 x 90 | Basket, instruments, for sterilization, wired, 400 x 200 x 90 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Stainless steel wired basket used as support for sterilizing surgical instruments. Material : austenitic stainless steel (non-quenched, non-magnetic steel).AISI 316 L stainless steel Rectangular basket, rounded not soldered corners. Close mesh netting (4.5x4.5mm). Tilting handles. Dimensions: approx. 400 x 200 x 90 mm. Instructions for use: Stainless steel wired basket used as support for sterilizing surgical instruments. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) sterilization basket in a protective plastic bag with manufacturer's instruction for user in English. Labelling : Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels — Part _1: List of stainless steels ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Stainless steel wired basket used as support for sterilizing surgical instruments. Material : austenitic stainless steel (non-quenched, non-magnetic steel).AISI 316 L stainless steel Rectangular basket, rounded not soldered corners. Close mesh netting (4.5x4.5mm). Tilting handles. Dimensions: approx. 400 x 200 x 90 mm. Instructions for use: Stainless steel wired basket used as support for sterilizing surgical instruments. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) sterilization basket in a protective plastic bag with manufacturer's instruction for user in English. Labelling : Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels — Part _1: List of stainless steels ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: Not use PVC polymer
126 | Medical Sterilization Equipment | STERILIZEDRUM165 | Drum, sterilizing, diameter 165mm | Drum, sterilizing, diameter 165mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Drum, sterilizing, 165 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.) Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 150 to 165 mm. Height: approx. 100 to 120 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.)in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities(i.e: dressing, injection etc.). Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e protected from dust, humidity, etc) It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Drum, sterilizing, 165 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.) Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 150 to 165 mm. Height: approx. 100 to 120 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.)in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities(i.e: dressing, injection etc.). Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e protected from dust, humidity, etc) It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices
127 | Medical Sterilization Equipment | STERILIZEDRUM260 | Drum, sterilizing, diameter 260mm | Drum, sterilizing, diameter 260mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 260 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 240 to 260 mm. Height: approx. 160 to 170 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments and others reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e. protected from dust, humidity etc.). It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 260 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 240 to 260 mm. Height: approx. 160 to 170 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments and others reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e. protected from dust, humidity etc.). It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices
128 | Medical Sterilization Equipment | STERILIZEDRUM290 | Drum, sterilizing, diameter 290mm | Drum, sterilizing, diameter 290mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 290 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 280 to 290 mm Height: approx. 160 to 180 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments and others reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e protected from dust, humidity, etc) It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark NOTE: after 27 May 2024 only certificate/Declaration of Conformity to MDR is acceptable (certificate to MDD is no longer acceptable) Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 290 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 280 to 290 mm Height: approx. 160 to 180 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments and others reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimize sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk of oxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e protected from dust, humidity, etc) It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark NOTE: after 27 May 2024 only certificate/Declaration of Conformity to MDR is acceptable (certificate to MDD is no longer acceptable) Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices
129 | Medical Sterilization Equipment | STERILIZEDRUM340 | Drum, sterilizing, diameter 340mm | Drum, sterilizing, diameter 340mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 340 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 340 to 350 mm. Height: approx. 180 to 200 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments, surgical gowns and towels and other reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimise sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk o foxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e:protected from dust, humidity etc.). It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Drum, sterilizing, 340 mm diameter. Intended use: Cylindrical container used to sterilise dressing materials (gauze compress or cotton etc.) in a steam sterilizer (autoclave), and to keep them as "sterile" dressing materials for medical activities (i.e. dressing, injection etc.). Technical Specifications: Drum, sterilizer. Closing lid with a clip lock. Carrying handle. Air vents system to allow steam to circulate freely during the sterilisation cycle. Lateral air vents system is preferable to top and bottom air vents. External diameter: approx. 340 to 350 mm. Height: approx. 180 to 200 mm. Thickness: approx. 0.6 to 0.7 mm. Material: Austenitic stainless steel, smooth surface. Austenitic stainless steel composition: approx. 8-10 % nickel, 18-20 % chromium. Reusable. Supplied with instructions for use. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 40% RH to 70% RH. Instructions for use: Cylindrical container used to sterilise dressing materials, surgical instruments, surgical gowns and towels and other reusable medical devices, in a steam sterilizer (autoclave), and to keep them as sterile products for medical activities. Open air vents system for sterilisation process. When the sterilisation cycle is completed, close air vents system immediately when the drum is removed from the autoclave. Do not stuff the drum full with compress or cotton, as the pressurised steam will not be able to penetrate into the centre and therefore not sterilise. The size has been chosen as being the most commonly used for the purpose of use. The sizes have been chosen taking into consideration different types of steam sterilizers (capacity and internal dimensions), in which this container will be used, to optimise sterilisation activities. Safety process: Check the drum (empty) before sterilisation process: This item must be cleaned, and disinfected. Make sure that the lid and air vents system close correctly. Never use chlorinated products on stainless steel, as there is a risk o foxidation. Check the drum (full) after sterilization cycle: Make sure that the lid and air vents system have been closed correctly. The date of the sterilisation must be written on the drum (use a sticker). The material (sterile) can be kept in the drum for up to one month, provided the drum is kept closed and in a good condition of storage (i.e:protected from dust, humidity etc.). It is recommended to stock sterile goods in a special room and / or cabinet and to have one person in charge. Packaging & Labelling: One (1) drum in a plastic bag + box with manufacturer's instruction for use in 3 languages ( English, French & Spanish) Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: EN 10088-1, Stainless steels ISO 14971 Application of risk management to medical devices
130 | Medical Sterilization Equipment | STERILIZER24L | Sterilizer, steam, pressure-type, 24 litres | Sterilizer, steam, pressure-type, 24 litres | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General description: Steam sterilizer, pressure type, capacity 24 litres. INTENDED USE A device designed for total elimination and/or inactivation of microorganisms from medical devices and related products using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It includes a treatment chamber with an inner container, it may be intended to sterilize wrapped and/or unwrapped devices. The device comes without a heating source and can be heated through a kerosine stove, electrical heater or even an open fire. TECHNICAL SPECIFICATIONS Minimum unit gross capacity: 24 litres, indicate if different. Sterilization capacity: 16.6 litres, indicate if different. Aluminium vessel suitable for sterilisation under superheated steam. Operational pressure: 17 - 21 PSI / 1.17 - 1.45 bar, , indicate if different. Operational temperature: 250°F / 121°C, indicate if different. Made of cast aluminium alloy which resist corrosion and aging. No separate seal required; uses metal-to-metal seal for the lid. Provided with a safety mechanism which prevents opening the unit it is still pressurized. Equipped with clamping locks made from Bakelite or an equivalent material. The removable cover is provided with a handle from Bakelite or an equivalent material. Contains a control valve with an extended exhaust tube to allow air trapped at the bottom to escape. Contains an excess-pressure relief valve. Contains a safety overpressure plug. Comes with an inner aluminium container with handles and a rack to be placed under the container. Scored water level mark inside chamber. Dial type geared steam gauge graduated in kg / cm², and/or PSI and degrees Fahrenheit and/or Celsius, and with colour-coding showing sterilizing zone (green) and caution zone (red). SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish 1 x Spare over pressure rubber plug. 1 x Document listing of accessories and spare parts. ESTIMATED LIFE SPAN Seven years WARRANTY Two years. ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Operating altitude: Atmospheric pressure: 1 – 70 kPa WEIGHT AND VOLUME (Packaged) - estimated Weight: 24.2 Kg. Volume: 153.5 dm³. Instructions for use: Comprehensive though basic sterilising device to be used in health centres and/or emergency situations. Requires a powerful heating source (kerosene stove, charcoal fire or electric plate 15000 W min.). Steam sterilizer should only be operated by staff who received adequate training on the device and on steam sterilisation processes. Always place the steam sterilizer on a stable surface (i.e: use a strong iron tripod with retaining lugs supporting the sterilizer). Never heat the steam sterilizer unless there is water in it. It is recommended to use low mineral containing water (e.g. rain water deionised or distilled water). The steam sterilizers must be cleaned and maintain properly according manufacturer's instructions. Packaging & Labelling: One (1) steam sterilizer wrapped in a plastic film with manufacturer's instruction for use, spare parts and accessories. CE mark and reference number of notifying body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 13060 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General description: Steam sterilizer, pressure type, capacity 24 litres. INTENDED USE A device designed for total elimination and/or inactivation of microorganisms from medical devices and related products using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It includes a treatment chamber with an inner container, it may be intended to sterilize wrapped and/or unwrapped devices. The device comes without a heating source and can be heated through a kerosine stove, electrical heater or even an open fire. TECHNICAL SPECIFICATIONS Minimum unit gross capacity: 24 litres, indicate if different. Sterilization capacity: 16.6 litres, indicate if different. Aluminium vessel suitable for sterilisation under superheated steam. Operational pressure: 17 - 21 PSI / 1.17 - 1.45 bar, , indicate if different. Operational temperature: 250°F / 121°C, indicate if different. Made of cast aluminium alloy which resist corrosion and aging. No separate seal required; uses metal-to-metal seal for the lid. Provided with a safety mechanism which prevents opening the unit it is still pressurized. Equipped with clamping locks made from Bakelite or an equivalent material. The removable cover is provided with a handle from Bakelite or an equivalent material. Contains a control valve with an extended exhaust tube to allow air trapped at the bottom to escape. Contains an excess-pressure relief valve. Contains a safety overpressure plug. Comes with an inner aluminium container with handles and a rack to be placed under the container. Scored water level mark inside chamber. Dial type geared steam gauge graduated in kg / cm², and/or PSI and degrees Fahrenheit and/or Celsius, and with colour-coding showing sterilizing zone (green) and caution zone (red). SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish 1 x Spare over pressure rubber plug. 1 x Document listing of accessories and spare parts. ESTIMATED LIFE SPAN Seven years WARRANTY Two years. ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Operating altitude: Atmospheric pressure: 1 – 70 kPa WEIGHT AND VOLUME (Packaged) - estimated Weight: 24.2 Kg. Volume: 153.5 dm³. Instructions for use: Comprehensive though basic sterilising device to be used in health centres and/or emergency situations. Requires a powerful heating source (kerosene stove, charcoal fire or electric plate 15000 W min.). Steam sterilizer should only be operated by staff who received adequate training on the device and on steam sterilisation processes. Always place the steam sterilizer on a stable surface (i.e: use a strong iron tripod with retaining lugs supporting the sterilizer). Never heat the steam sterilizer unless there is water in it. It is recommended to use low mineral containing water (e.g. rain water deionised or distilled water). The steam sterilizers must be cleaned and maintain properly according manufacturer's instructions. Packaging & Labelling: One (1) steam sterilizer wrapped in a plastic film with manufacturer's instruction for use, spare parts and accessories. CE mark and reference number of notifying body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 13060
131 | Medical Sterilization Equipment | STERILIZER39L | Sterilizer, steam, pressure-type, 39 litres | Sterilizer, steam, pressure-type, 39 litres | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Steam sterilizer, pressure type, capacity 39 litres. INTENDED USE A device designed for total elimination and/or inactivation of microorganisms from medical devices and related products using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It includes a treatment chamber with an inner container, it may be intended to sterilize wrapped and/or unwrapped devices. The device comes without a heating source and can be heated through a kerosine stove, electrical heater or even an open fire. TECHNICAL SPECIFICATIONS Minimum unit gross capacity: 39 litres, indicate if different. Sterilization capacity: 26.4 litres, indicate if different. Aluminium vessel suitable for sterilisation under superheated steam. Operational pressure: 17 - 21 PSI / 1.17 - 1.45 bar, , indicate if different. Operational temperature: 250°F / 121°C, indicate if different. Made of cast aluminium alloy which resist corrosion and aging. No separate seal required; uses metal-to-metal seal for the lid. Provided with a safety mechanism which prevents opening the unit it is still pressurized. Equipped with clamping locks made from Bakelite or an equivalent material. The removable cover is provided with a handle from Bakelite or an equivalent material. Contains a control valve with an extended exhaust tube to allow air trapped at the bottom to escape. Contains an excess-pressure relief valve. Contains a safety overpressure plug. Comes with an inner aluminium container with handles and a rack to be placed under the container. Scored water level mark inside chamber. Dial type geared steam gauge graduated in kg / cm², and/or PSI and degrees Fahrenheit and/or Celsius, and with colour-coding showing sterilizing zone (green) and caution zone (red). SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish 1 x Spare over pressure rubber plug. 1 x Document listing of accessories and spare parts ESTIMATED LIFE SPAN Seven years WARRANTY Two years. ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Atmospheric pressure: 1 – 70 kPa WEIGHT AND VOLUME (Packaged) - estimated Weight: 24.2 Kg. Volume: 153.5 dm³. Instructions for use: Comprehensive though basic sterilising device to be used in health centres and/or emergency situations. Requires a powerful heating source (kerosene stove, charcoal fire or electric plate 15000 W min.). Steam sterilizer should only be operated by staff who received adequate training on the device and on steam sterilisation processes. Always place the steam sterilizer on a stable surface (i.e: use a strong iron tripod with retaining lugs supporting the sterilizer). Never heat the steam sterilizer unless there is water in it. It is recommended to use low mineral containing water (e.g. rain water deionised or distilled water). The steam sterilizers must be cleaned and maintain properly according manufacturer's instructions. Packaging & Labelling: One (1) steam sterilizer wrapped in a plastic film with manufacturer's instruction for use, spare parts and accessories. CE mark and reference number of notifying body. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 13060 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. GENERAL DESCRIPTION Steam sterilizer, pressure type, capacity 39 litres. INTENDED USE A device designed for total elimination and/or inactivation of microorganisms from medical devices and related products using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It includes a treatment chamber with an inner container, it may be intended to sterilize wrapped and/or unwrapped devices. The device comes without a heating source and can be heated through a kerosine stove, electrical heater or even an open fire. TECHNICAL SPECIFICATIONS Minimum unit gross capacity: 39 litres, indicate if different. Sterilization capacity: 26.4 litres, indicate if different. Aluminium vessel suitable for sterilisation under superheated steam. Operational pressure: 17 - 21 PSI / 1.17 - 1.45 bar, , indicate if different. Operational temperature: 250°F / 121°C, indicate if different. Made of cast aluminium alloy which resist corrosion and aging. No separate seal required; uses metal-to-metal seal for the lid. Provided with a safety mechanism which prevents opening the unit it is still pressurized. Equipped with clamping locks made from Bakelite or an equivalent material. The removable cover is provided with a handle from Bakelite or an equivalent material. Contains a control valve with an extended exhaust tube to allow air trapped at the bottom to escape. Contains an excess-pressure relief valve. Contains a safety overpressure plug. Comes with an inner aluminium container with handles and a rack to be placed under the container. Scored water level mark inside chamber. Dial type geared steam gauge graduated in kg / cm², and/or PSI and degrees Fahrenheit and/or Celsius, and with colour-coding showing sterilizing zone (green) and caution zone (red). SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish 1 x Spare over pressure rubber plug. 1 x Document listing of accessories and spare parts ESTIMATED LIFE SPAN Seven years WARRANTY Two years. ENVIRONMENTAL CONDITIONS Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Atmospheric pressure: 1 – 70 kPa WEIGHT AND VOLUME (Packaged) - estimated Weight: 24.2 Kg. Volume: 153.5 dm³. Instructions for use: Comprehensive though basic sterilising device to be used in health centres and/or emergency situations. Requires a powerful heating source (kerosene stove, charcoal fire or electric plate 15000 W min.). Steam sterilizer should only be operated by staff who received adequate training on the device and on steam sterilisation processes. Always place the steam sterilizer on a stable surface (i.e: use a strong iron tripod with retaining lugs supporting the sterilizer). Never heat the steam sterilizer unless there is water in it. It is recommended to use low mineral containing water (e.g. rain water deionised or distilled water). The steam sterilizers must be cleaned and maintain properly according manufacturer's instructions. Packaging & Labelling: One (1) steam sterilizer wrapped in a plastic film with manufacturer's instruction for use, spare parts and accessories. CE mark and reference number of notifying body. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 13060
132 | Medical Sterilization Equipment | STOVEKEROSENE1.5L | Stove, kerosene, single-burner, pressure-type | Stove, kerosene, single-burner, pressure-type | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Stove, kerosene, single-burner, pressure Kerosene stove designed to provide safe and sufficient heat for medical sterilization applications. Single burner kerosene stove for use with steam sterilizers. Technical specifications: Comprises of a fuel tank, vapour burner and pot holder Single burner with top plate Diameter flame holder: 5-7 cm Hand-operated plunger pump allows pressurisation of fuel tank and control of flame intensity Pressure, max: 1.3 bar Fit with overpressure safety valve Capacity kerosene fuel tank: 1.2 to 2.5 L Fuel consumption: 0.4-0.6 L/h Height: 20-29 cm Material: Heat and rustproof, steel or brass Items supplied with: 10x Cleaning needle 1 x Nipple key 1 x Plastic funnel with filter 1 x Spanner for burner 2 x Spare nipple 2 x Spare packing for burner 2 x Spare leather cup for pump Must be supplied with clear instructions for use / diagrams for assembly in 3 languages (English, French and Spanish) and list of accessories / parts. Instructions for use: For heating steam sterilizers and for cooking purposes in emergency settings. Use only kerosene. Fill the tank 2/3 full, never fill it completely. Stove kerosene must be cleaned and maintained properly. Clean the nozzle of the burner regularly with the cleaning needle. Grease the leather washer of the pump barrel regularly(once a week). Operate the stove on a level ground. To increase stability when using large steam sterilizers and cooking pots, it is recommended to use a three or four leg separate support frame. Important: It is recommended to follow manufacturer's instructions for use. Packaging & Labelling: Primary packaging: Unit of use One (1) kerosene stove in a plastic bag + box with manufacturer's instruction for use, spare parts and accessories. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit to be the same as primary packaging. Extra information required: Number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Stove, kerosene, single-burner, pressure Kerosene stove designed to provide safe and sufficient heat for medical sterilization applications. Single burner kerosene stove for use with steam sterilizers. Technical specifications: Comprises of a fuel tank, vapour burner and pot holder Single burner with top plate Diameter flame holder: 5-7 cm Hand-operated plunger pump allows pressurisation of fuel tank and control of flame intensity Pressure, max: 1.3 bar Fit with overpressure safety valve Capacity kerosene fuel tank: 1.2 to 2.5 L Fuel consumption: 0.4-0.6 L/h Height: 20-29 cm Material: Heat and rustproof, steel or brass Items supplied with: 10x Cleaning needle 1 x Nipple key 1 x Plastic funnel with filter 1 x Spanner for burner 2 x Spare nipple 2 x Spare packing for burner 2 x Spare leather cup for pump Must be supplied with clear instructions for use / diagrams for assembly in 3 languages (English, French and Spanish) and list of accessories / parts. Instructions for use: For heating steam sterilizers and for cooking purposes in emergency settings. Use only kerosene. Fill the tank 2/3 full, never fill it completely. Stove kerosene must be cleaned and maintained properly. Clean the nozzle of the burner regularly with the cleaning needle. Grease the leather washer of the pump barrel regularly(once a week). Operate the stove on a level ground. To increase stability when using large steam sterilizers and cooking pots, it is recommended to use a three or four leg separate support frame. Important: It is recommended to follow manufacturer's instructions for use. Packaging & Labelling: Primary packaging: Unit of use One (1) kerosene stove in a plastic bag + box with manufacturer's instruction for use, spare parts and accessories. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit to be the same as primary packaging. Extra information required: Number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 9001:2008
133 | Medical Sterilization Equipment | TIMEMECH0-60MIN | Timer, 60 min, mechanical | Timer, 60 min, mechanical | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Timer, 60 minutes, mechanical Product Description: Mechanical, spring driven countdown timer, professional use Time setting via rotating the housing/ hand Ring or dial, easy to read, graduated 0 to 60 minutes, in 5 and 1 minute intervals Loud long ring-alert at time elapse Robust construction, housing, spring and gears: shock resistant Can be positioned stable on a surface or clipped on clothes Smooth finishing allows for easy to clean Easy to transport, Light-weighted construction Material: Mechanics inside made of steel Housing: Housing made of easy-to-clean and resistant plastic Weight - Volume - estimated Estimated weight: 0.3kg Estimated volume: 0.0032m3 Supplied with: Manufacturer's instruction for use. Instructions for use: Mechanical timer for the laboratory Used at facilities with limited access to replacement batteries (laboratory network levels 1-II) Packaging & Labelling: Primary packaging: Unit presentation: 1(one) Timer, 60 minutes, mechanical Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 9001:2008 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General Description: Timer, 60 minutes, mechanical Product Description: Mechanical, spring driven countdown timer, professional use Time setting via rotating the housing/ hand Ring or dial, easy to read, graduated 0 to 60 minutes, in 5 and 1 minute intervals Loud long ring-alert at time elapse Robust construction, housing, spring and gears: shock resistant Can be positioned stable on a surface or clipped on clothes Smooth finishing allows for easy to clean Easy to transport, Light-weighted construction Material: Mechanics inside made of steel Housing: Housing made of easy-to-clean and resistant plastic Weight - Volume - estimated Estimated weight: 0.3kg Estimated volume: 0.0032m3 Supplied with: Manufacturer's instruction for use. Instructions for use: Mechanical timer for the laboratory Used at facilities with limited access to replacement batteries (laboratory network levels 1-II) Packaging & Labelling: Primary packaging: Unit presentation: 1(one) Timer, 60 minutes, mechanical Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 9001:2008
134 | PPEs | APRONDISPOSABLE_1 | Apron, surgical, disposable, water resistant, non-sterile | Apron, surgical, disposable, water resistant, non-sterile | Product description: Apron, disposable, non sterile This is a disposable, single use garment that must be disposed after one time use. Intended purpose of use: A garment designed to be worn by medical staff for general protection from soiling or spills on the ventral (front) aspect of the body, and typically to prevent penetration of infectious material, during the course of work or a procedure. Material: Plastic, Polyethylene etc. PVC is not recommended. Good resistance to tearing and perforation. Seamless liquid proof and stain resistant. Straight apron with bib. Antistatic. Thickness: not less than: 50 µm. Can resist water and disinfectant (ethanol 70% and chlorine solution 0.05% or 500 ppm). Back fastening and neckband should be strong and not detachable. Colour: Any color. Size: One-size-fit-all (120-150 x 70-90 cm) Single use, disposable. Non-sterile. Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): • EN ISO 13688 • EN 14126-B and partial body protection (EN 13034 or EN 14605) • EN 343 for water and breathability or equivalent Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Packaging and labelling: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: One (1) unit in a protective packaging. Manufacturer name and address. mfg date and expiry ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, manufacturing date and expiry date. Word ‘non-sterile.’ Word ‘single use' and/or 'disposable.’ | Product description: Apron, disposable, non sterile This is a disposable, single use garment that must be disposed after one time use. Intended purpose of use: A garment designed to be worn by medical staff for general protection from soiling or spills on the ventral (front) aspect of the body, and typically to prevent penetration of infectious material, during the course of work or a procedure. Material: Plastic, Polyethylene etc. PVC is not recommended. Good resistance to tearing and perforation. Seamless liquid proof and stain resistant. Straight apron with bib. Antistatic. Thickness: not less than: 50 µm. Can resist water and disinfectant (ethanol 70% and chlorine solution 0.05% or 500 ppm). Back fastening and neckband should be strong and not detachable. Colour: Any color. Size: One-size-fit-all (120-150 x 70-90 cm) Single use, disposable. Non-sterile. Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): • EN ISO 13688 • EN 14126-B and partial body protection (EN 13034 or EN 14605) • EN 343 for water and breathability or equivalent Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Packaging and labelling: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: One (1) unit in a protective packaging. Manufacturer name and address. mfg date and expiry ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, manufacturing date and expiry date. Word ‘non-sterile.’ Word ‘single use' and/or 'disposable.’
135 | PPEs | GLOVESNITR_L_100 | Gloves, exam, nitrile, powder free, non-sterile, large, P100 | Gloves, exam, nitrile, powder free, non-sterile, large, P100 | Intended Use: To prevent contamination of the hands when providing care for patients. General Description: Examination Gloves, nitrile, Powderfree, Large, disposable, box/100 Technical Specifications: • Material: nitrile • Powder free • Non-sterile • Single-use disposable • Size: large (8.5 to 9.5) • Fits either hand • Thickness at fingertips of >= 120um • Freedom from hole AQL 1.5 • Box of 100 gloves • CE marked - Conform to EN-455 and EN 374 • Protection against dangerous chemicals and microorganism Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): EN 455-1, EN 455-2, EN 455-3, EN 455-4 EN ISO 374-1:2016+A1:2018 Protective gloves against dangerous chemicals and micro-organisms. Terminology and performance requirements for chemical risks EN ISO 374-2:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to penetration EN ISO 374-4:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to degradation by chemicals OR equivalent set of standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Labelling on the primary packaging: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. | Intended Use: To prevent contamination of the hands when providing care for patients. General Description: Examination Gloves, nitrile, Powderfree, Large, disposable, box/100 Technical Specifications: • Material: nitrile • Powder free • Non-sterile • Single-use disposable • Size: large (8.5 to 9.5) • Fits either hand • Thickness at fingertips of >= 120um • Freedom from hole AQL 1.5 • Box of 100 gloves • CE marked - Conform to EN-455 and EN 374 • Protection against dangerous chemicals and microorganism Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): EN 455-1, EN 455-2, EN 455-3, EN 455-4 EN ISO 374-1:2016+A1:2018 Protective gloves against dangerous chemicals and micro-organisms. Terminology and performance requirements for chemical risks EN ISO 374-2:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to penetration EN ISO 374-4:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to degradation by chemicals OR equivalent set of standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Labelling on the primary packaging: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert.
136 | PPEs | GLOVESNITR_M_100 | Gloves, exam, nitrile, powder free, non-sterile, medium,P100 | Gloves, exam, nitrile, powder free, non-sterile, medium,P100 | Intended Use: To prevent contamination of the hands when providing care for patients. General Description: Examination Gloves, nitrile, Powderfree, Medium, disposable, box/100 Technical Specifications: • Material: nitrile • Powder free • Non-sterile • Single-use disposable • Size: medium (7 to 8) • Fits either hand • Thickness at fingertips of >= 120um • Freedom from hole AQL 1.5 • Box of 100 gloves • CE marked - Conform to EN-455 and EN 374 • Protection against dangerous chemicals and microorganism Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): EN 455-1, EN 455-2, EN 455-3, EN 455-4 EN ISO 374-1:2016+A1:2018 Protective gloves against dangerous chemicals and micro-organisms. Terminology and performance requirements for chemical risks EN ISO 374-2:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to penetration EN ISO 374-4:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to degradation by chemicals OR equivalent set of standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Labelling on the primary packaging: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert. | Intended Use: To prevent contamination of the hands when providing care for patients. General Description: Examination Gloves, nitrile, Powderfree, Medium, disposable, box/100 Technical Specifications: • Material: nitrile • Powder free • Non-sterile • Single-use disposable • Size: medium (7 to 8) • Fits either hand • Thickness at fingertips of >= 120um • Freedom from hole AQL 1.5 • Box of 100 gloves • CE marked - Conform to EN-455 and EN 374 • Protection against dangerous chemicals and microorganism Shelf life 3-5 years Applicable standards (submissions to be made as per latest revisions): EN 455-1, EN 455-2, EN 455-3, EN 455-4 EN ISO 374-1:2016+A1:2018 Protective gloves against dangerous chemicals and micro-organisms. Terminology and performance requirements for chemical risks EN ISO 374-2:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to penetration EN ISO 374-4:2019 Protective gloves against dangerous chemicals and micro-organisms. Determination of resistance to degradation by chemicals OR equivalent set of standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Labelling on the primary packaging: - Instruction for use in langauages English/ French/arabic/spanish or pictograms Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) (if applicable). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows). Number of units per primary packaging (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Manufacturer's instruction for use. Alternatively, the instruction for use can be indicated on a separate insert.
137 | PPEs | GOWNSURGICALSTERIL | Gown, surgical, non-woven, sterile, disposable | Gown, surgical, non-woven, sterile, disposable | Intended purpose of use: A single use garment, also known as an operating room (OR) gown, intended to be worn over a scrub suit to cover the arms, trunk, and upper legs, during a surgical procedure in order to prevent transmission of infectious agents. Technical specifications: Non-woven, multilayer (SMS or SMMS). Impervious supplementary protection layer (polyethylene film, nonwoven, etc.) in the critical zones (include the front of the body from top of shoulders to knees and the arms from the wrist cuff to above the elbow) = reinforced. Breathable For medium to high fluid level interventions. Wraparound. Long sleeves. Cuffs made of elastic jersey. Latex free. Sterile, for single use. Sizes available: Large, XLarge and XXLarge. Length (shoulder seam to mid calf): size Large: 140 cm (dimensions within 2% tolerance limits) size XLarge: 150 cm (dimensions within 2% tolerance limits) size XXLarge: 160 -170 cm Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions to be made as per latest revisions): • - EN 13795-1:2002: Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment. General requirements for manufacturers, processors and products - EN 13795-2:2004: Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. Test methods - EN 13795-3:2006: Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. Performance requirements and performance levels • AAMI PB70 for moderate risk and high risk situations Intended use: Strictly single use. Non-woven gowns are intended for situations where sterilization is problematic. Used in procedures where high level sterility is required. Packaging and labelling: One (1) unit with reverse folding in order to ensure an aseptic gowning technique. Wrapped in a drape. Double sterile packaging in peel-open pack. Labelling on primary packaging (one unit) must include: - Instruction for use in langauages English/ French/arabic/spanish or pictograms - Name and/or trademark of the manufacturer - Manufacturer address - Manufacturer's product reference (product code) - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) - Words ‘sterile with sterilization method, disposable, single use’ - Size - Batch/lot and expiry date. - CE mark with notified body number (+ EC REP), FDA, and equivalent Secondary packaging: Packaging of multiple units Labelling to be the same as primary packaging. Extra information required: Number of units per box | Intended purpose of use: A single use garment, also known as an operating room (OR) gown, intended to be worn over a scrub suit to cover the arms, trunk, and upper legs, during a surgical procedure in order to prevent transmission of infectious agents. Technical specifications: Non-woven, multilayer (SMS or SMMS). Impervious supplementary protection layer (polyethylene film, nonwoven, etc.) in the critical zones (include the front of the body from top of shoulders to knees and the arms from the wrist cuff to above the elbow) = reinforced. Breathable For medium to high fluid level interventions. Wraparound. Long sleeves. Cuffs made of elastic jersey. Latex free. Sterile, for single use. Sizes available: Large, XLarge and XXLarge. Length (shoulder seam to mid calf): size Large: 140 cm (dimensions within 2% tolerance limits) size XLarge: 150 cm (dimensions within 2% tolerance limits) size XXLarge: 160 -170 cm Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions to be made as per latest revisions): • - EN 13795-1:2002: Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment. General requirements for manufacturers, processors and products - EN 13795-2:2004: Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. Test methods - EN 13795-3:2006: Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. Performance requirements and performance levels • AAMI PB70 for moderate risk and high risk situations Intended use: Strictly single use. Non-woven gowns are intended for situations where sterilization is problematic. Used in procedures where high level sterility is required. Packaging and labelling: One (1) unit with reverse folding in order to ensure an aseptic gowning technique. Wrapped in a drape. Double sterile packaging in peel-open pack. Labelling on primary packaging (one unit) must include: - Instruction for use in langauages English/ French/arabic/spanish or pictograms - Name and/or trademark of the manufacturer - Manufacturer address - Manufacturer's product reference (product code) - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) - Words ‘sterile with sterilization method, disposable, single use’ - Size - Batch/lot and expiry date. - CE mark with notified body number (+ EC REP), FDA, and equivalent Secondary packaging: Packaging of multiple units Labelling to be the same as primary packaging. Extra information required: Number of units per box
138 | PPEs | MASKMEDICAL_I_P50 | Medical Mask, type I, for patients, disposable | Medical Mask, type I, for patients, disposable | Intended purpose of use: Mask is covering the nose, mouth and chin, designed to limit transmission of infectious agents exhaled by the nose and mouth of the patient. This mask (type I) is not intended for protection of the wearer from the viral infection (Note: this mask is NOT to be used by the healthcare workers). Surgical or medical mask, type I, for patients, disposable General description: Disposable surgical or medical mask, type I, for patients suspected or confirmed viral infection. Technical specifications: . Type I (EN 14683) surgical or medical mask. . Bacterial filtering efficiency (BFE): equal to or greater than 95% (AQL 4% minimum 5 samples as per ISO 2859-1) . Differential pressure (breathability)/Breathing resistance: lower than 40 Pa/cm2 as per EN 14683 or equivalent . Splash resistance: None. . Fabric, non-woven with outer layer impervious liquid splash resistant material, e.g. polyethylene. . Comprised of 3 non-woven folded layers, shape completely covering nose, mouth and chin. . Clearly identifiable inner and outer surfaces. . Malleable nose strip, made of aluminum, allowing a snug fit. . With attached 2 x 2 tie-straps, allowing correct fixation and securing at the back of the head, or ear loops. . Size (indicative): 15-19 cm x 9-11 cm (l x w). Unfolded 175 x 175 mm. . Latex-free, glass fibre-free . Non-sterile, microbial cleanliness ≤30 cfu/g . Single use, disposable Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions shall be as per current standards): • EN ISO 11737-1, The bioburden of the medical mask or as per Annex D of EN 14683:2019. • EN 14683 Type I • ASTM F2100 minimum level 1 or equivalent. • Biocompatibility as the item is a skin contact device ISO 10993-1, 5&10 (cytotoxicity, irritation and sensitization) Intended use: Worn by the contagious patient, suspected or verified COVID-19 patients. The surgical mask prevents the contamination spread to the people surrounding and the environment (air, surface, products…) around the wearer because the mask captures liquid droplets from the nose and mouth of the wearer. Safety instructions: Caution! This mask does not protect the wearer against liquid splashes, and it is not designed to prevent getting infected. Note! A surgical or medical mask does not protect the wearer against airborne infectious agents (coronavirus, TB, viral haemorrhagic fever, measles, varicella, SARS, avian influenza, etc.). Not for reuse after removing from the face. The mask is to be replaced at least every 3 hours. Appropriate hand hygiene is to be applied before fitting and after removing the mask. Packaging and labelling: Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: Multiple units (25, 50) per box. Manufacturer name and/or trademark, and address. Manufacturer's product reference. ISO 15223 CE mark, FDA and equivalent. Lot/batch, MFD and expiry date. Word ‘non-sterile, single use, disposable.’ Comes with instructions for use. Type I (EN 14683) is indicated. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate. Information for handling, if applicable (or equivalent harmonised symbol). | Intended purpose of use: Mask is covering the nose, mouth and chin, designed to limit transmission of infectious agents exhaled by the nose and mouth of the patient. This mask (type I) is not intended for protection of the wearer from the viral infection (Note: this mask is NOT to be used by the healthcare workers). Surgical or medical mask, type I, for patients, disposable General description: Disposable surgical or medical mask, type I, for patients suspected or confirmed viral infection. Technical specifications: . Type I (EN 14683) surgical or medical mask. . Bacterial filtering efficiency (BFE): equal to or greater than 95% (AQL 4% minimum 5 samples as per ISO 2859-1) . Differential pressure (breathability)/Breathing resistance: lower than 40 Pa/cm2 as per EN 14683 or equivalent . Splash resistance: None. . Fabric, non-woven with outer layer impervious liquid splash resistant material, e.g. polyethylene. . Comprised of 3 non-woven folded layers, shape completely covering nose, mouth and chin. . Clearly identifiable inner and outer surfaces. . Malleable nose strip, made of aluminum, allowing a snug fit. . With attached 2 x 2 tie-straps, allowing correct fixation and securing at the back of the head, or ear loops. . Size (indicative): 15-19 cm x 9-11 cm (l x w). Unfolded 175 x 175 mm. . Latex-free, glass fibre-free . Non-sterile, microbial cleanliness ≤30 cfu/g . Single use, disposable Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions shall be as per current standards): • EN ISO 11737-1, The bioburden of the medical mask or as per Annex D of EN 14683:2019. • EN 14683 Type I • ASTM F2100 minimum level 1 or equivalent. • Biocompatibility as the item is a skin contact device ISO 10993-1, 5&10 (cytotoxicity, irritation and sensitization) Intended use: Worn by the contagious patient, suspected or verified COVID-19 patients. The surgical mask prevents the contamination spread to the people surrounding and the environment (air, surface, products…) around the wearer because the mask captures liquid droplets from the nose and mouth of the wearer. Safety instructions: Caution! This mask does not protect the wearer against liquid splashes, and it is not designed to prevent getting infected. Note! A surgical or medical mask does not protect the wearer against airborne infectious agents (coronavirus, TB, viral haemorrhagic fever, measles, varicella, SARS, avian influenza, etc.). Not for reuse after removing from the face. The mask is to be replaced at least every 3 hours. Appropriate hand hygiene is to be applied before fitting and after removing the mask. Packaging and labelling: Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: Multiple units (25, 50) per box. Manufacturer name and/or trademark, and address. Manufacturer's product reference. ISO 15223 CE mark, FDA and equivalent. Lot/batch, MFD and expiry date. Word ‘non-sterile, single use, disposable.’ Comes with instructions for use. Type I (EN 14683) is indicated. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate. Information for handling, if applicable (or equivalent harmonised symbol).
139 | PPEs | MASKRESPIRATOR_NV | Respirator, FFP2/N95/KN95, mask, no valve, disposable | Respirator, FFP2/N95/KN95, mask, no valve, disposable | Intended purpose of use: For medical/surgical use It is recommended for use to reduce the Healthcare worker’s risk of inhaling hazardous airborne particles (including dust particles and infectious agents), gases, or vapors. Surgical respirator, high-filtration, FFP2/N95/KN95 or higher, mask, no valve, non-sterile, disposable General Description: Respirator mask protecting against airborne pathogens. Anti-penetration high filtration mask. Technical specifications • Material: non-woven filter layer •N95 (United States NIOSH-42CFR84) • FFP2 (Europe EN 149) • KN95 (China GB2626) • Filtration level: ≥94% FFP2 as per EN 149; ≥95% for N95 as per USA NIOSH ; ≥95% KN95 as per GB2626 China • Total Inward Leakage (TIL): <10% (N95) or <8% (FFP2/KN95) • Fluid resistance of 120 mmHg for surgical respirator: Minimum NIOSH approved (42 CFR Part 84)and FDA cleared “surgical N95” ; EN 149, minimum “FFP2“ and EN 14683 Type IIR ; GB 19083, minimum “Grade/Level 1” • Air permeability: > 2 mm H2O • Shape of the mask: duckbill, folded (horizontal) width-wise , shall not collapse against the face. • Without valve • Respirator mask fits all face shapes, without inspiration/expiration air-leakage • Upper edge has integrated easy malleable nose bridge strip reducing fogging of protective eye-wear • Size nose bridge strip: 4 x 90 mm (w x l) (+/-10%) • Two pre-attached elasticated straps, fitting (i) around top of the head,(ii) around base of the head • Colour: white • Non-sterile • Single use, disposable • Each mask bares clear identification of - protection provided FFP2/N95/KN95, - which side to wear up (nose), - manufacturer's name, and - model reference Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Respirator mask conforms to current standards: NIOSH N95 recommended by the CDC and WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, the part number (P/N), the protection provided by the filter (e.g. N-95),and "NIOSH." Or EN 149: 2001/FFP2 recommended by WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, EN 149: year/FFP2 and CE certification number. Must comply with PPE Regulation (EU) 2016/425 - category III. Or GB 2626, minimum KN95 recommended by WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, GB 2626/KN95 and product reference. Packaging and Labelling: IFU as per manufacturer, precautions and warnings shelf life, expiry date, storage confiitons of temperature and humidity, product reference number, manufacturer name and address , Labelling as per applicable standards for FFP2/N95/KN95. | Intended purpose of use: For medical/surgical use It is recommended for use to reduce the Healthcare worker’s risk of inhaling hazardous airborne particles (including dust particles and infectious agents), gases, or vapors. Surgical respirator, high-filtration, FFP2/N95/KN95 or higher, mask, no valve, non-sterile, disposable General Description: Respirator mask protecting against airborne pathogens. Anti-penetration high filtration mask. Technical specifications • Material: non-woven filter layer •N95 (United States NIOSH-42CFR84) • FFP2 (Europe EN 149) • KN95 (China GB2626) • Filtration level: ≥94% FFP2 as per EN 149; ≥95% for N95 as per USA NIOSH ; ≥95% KN95 as per GB2626 China • Total Inward Leakage (TIL): <10% (N95) or <8% (FFP2/KN95) • Fluid resistance of 120 mmHg for surgical respirator: Minimum NIOSH approved (42 CFR Part 84)and FDA cleared “surgical N95” ; EN 149, minimum “FFP2“ and EN 14683 Type IIR ; GB 19083, minimum “Grade/Level 1” • Air permeability: > 2 mm H2O • Shape of the mask: duckbill, folded (horizontal) width-wise , shall not collapse against the face. • Without valve • Respirator mask fits all face shapes, without inspiration/expiration air-leakage • Upper edge has integrated easy malleable nose bridge strip reducing fogging of protective eye-wear • Size nose bridge strip: 4 x 90 mm (w x l) (+/-10%) • Two pre-attached elasticated straps, fitting (i) around top of the head,(ii) around base of the head • Colour: white • Non-sterile • Single use, disposable • Each mask bares clear identification of - protection provided FFP2/N95/KN95, - which side to wear up (nose), - manufacturer's name, and - model reference Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Respirator mask conforms to current standards: NIOSH N95 recommended by the CDC and WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, the part number (P/N), the protection provided by the filter (e.g. N-95),and "NIOSH." Or EN 149: 2001/FFP2 recommended by WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, EN 149: year/FFP2 and CE certification number. Must comply with PPE Regulation (EU) 2016/425 - category III. Or GB 2626, minimum KN95 recommended by WHO for healthcare workers. The disposable respirator should be marked with the manufacturer's name, GB 2626/KN95 and product reference. Packaging and Labelling: IFU as per manufacturer, precautions and warnings shelf life, expiry date, storage confiitons of temperature and humidity, product reference number, manufacturer name and address , Labelling as per applicable standards for FFP2/N95/KN95.
140 | PPEs | MASKSURG_IIR_P50 MASKSURG_IIR_P25 | Surgical Mask, type IIR, healthcare worker, disposable | Surgical Mask, type IIR, healthcare worker, disposable | Intended purpose of use: Medical mask covering the nose, mouth and chin, designed to limit transmission of infectious agents exhaled by the nose and mouth of the wearer, and additionally to protect the wearer against liquid splashes. Surgical Mask, type IIR, Splash resistant, non-sterile, disposable General description: Mask, surgical, type IIR, tie strap or ear loops, disposable. MASKSURG_IIR_P50- pack of 50 MASKSURG_IIR_P25 - pack of 25 Technical specifications: . Splash resistant, type IIR surgical mask. . Bacterial filtering efficiency (BFE): equal to or greater than 98% (AQL 4% for BFE as per ISO 2859-1). . Differential pressure (breathability)/Breathing resistance: equal to or less than 49 Pa/cm2 as per ASTM F2100 (equal to or less than 60 Pa/cm2 EN 14638) . Splash resistance pressure: greater than 120 mmHg (tested in accordance with ASTM F1862 standard). . Fabric, non-woven with outer layer impervious liquid splash resistant material, f.e. polyethylene. . Comprised of 3 or 4 non-woven folded layers, shape completely covering nose, mouth and chin. . Clearly identifiable inner and outer surfaces. . Malleable nose strip, made of aluminum, allowing a snug fit. . With attached 2 x 2 tie-straps, allowing correct fixation and securing at the back of the head, or ear loops. . Size (indicative): 15-19 cm x 9-11 cm (l x w). Unfolded 175 x 175 mm. . Latex-free, glass fibre-free . Non-sterile, microbial cleanliness ≤30 cfu/g . Single use, disposable Shelf life 3-5 years . Conform requirements of EU Medical Devices Directive 93/42 (or equivalent internationally recognized marketing clearance). . In specific, compliant with the EN 14683 standard for type IIR (or equivalent international standard). Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Applicable standards (submissions shall be as per current standards): • EN 14683 Type IIR • ASTM F2100 minimum level 2 • Biocompatibility as the item is a skin contact device ISO 10993-1, 5&10 (cytotoxicity, irritation and sensitization) Intended use: Worn by the medical staff or most typically by the contagious patient. The surgical mask prevents the contamination spread to the people surrounding and the environment (air, surface, products…) around the wearer and protects the wearer against liquid splashes. Not to be reused after removing from the face. The mask must be replaced at least every 3 hours. Always wash hands before fitting and after removing the mask. Note! A surgical mask does not protect the wearer against airborne infectious agents (coronavirus, TB, viral haemorrhagic fever, measles, varicella, SARS, avian influenza, etc.). In such cases, it is advisable to wear a respiratory protective mask rated FFP2 (complies with european standards) or N95 (complies with american standards). Packaging and labelling: Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: One (1) unit in a protective packaging. EN 14683 type IIR or ASTM level 2 is mentioned on the box Manufacturer name and address. ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, MFD and expiry date. Word ‘non-sterile, single use, disposable.’ Comes with instructions for use. | Intended purpose of use: Medical mask covering the nose, mouth and chin, designed to limit transmission of infectious agents exhaled by the nose and mouth of the wearer, and additionally to protect the wearer against liquid splashes. Surgical Mask, type IIR, Splash resistant, non-sterile, disposable General description: Mask, surgical, type IIR, tie strap or ear loops, disposable. MASKSURG_IIR_P50- pack of 50 MASKSURG_IIR_P25 - pack of 25 Technical specifications: . Splash resistant, type IIR surgical mask. . Bacterial filtering efficiency (BFE): equal to or greater than 98% (AQL 4% for BFE as per ISO 2859-1). . Differential pressure (breathability)/Breathing resistance: equal to or less than 49 Pa/cm2 as per ASTM F2100 (equal to or less than 60 Pa/cm2 EN 14638) . Splash resistance pressure: greater than 120 mmHg (tested in accordance with ASTM F1862 standard). . Fabric, non-woven with outer layer impervious liquid splash resistant material, f.e. polyethylene. . Comprised of 3 or 4 non-woven folded layers, shape completely covering nose, mouth and chin. . Clearly identifiable inner and outer surfaces. . Malleable nose strip, made of aluminum, allowing a snug fit. . With attached 2 x 2 tie-straps, allowing correct fixation and securing at the back of the head, or ear loops. . Size (indicative): 15-19 cm x 9-11 cm (l x w). Unfolded 175 x 175 mm. . Latex-free, glass fibre-free . Non-sterile, microbial cleanliness ≤30 cfu/g . Single use, disposable Shelf life 3-5 years . Conform requirements of EU Medical Devices Directive 93/42 (or equivalent internationally recognized marketing clearance). . In specific, compliant with the EN 14683 standard for type IIR (or equivalent international standard). Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Applicable standards (submissions shall be as per current standards): • EN 14683 Type IIR • ASTM F2100 minimum level 2 • Biocompatibility as the item is a skin contact device ISO 10993-1, 5&10 (cytotoxicity, irritation and sensitization) Intended use: Worn by the medical staff or most typically by the contagious patient. The surgical mask prevents the contamination spread to the people surrounding and the environment (air, surface, products…) around the wearer and protects the wearer against liquid splashes. Not to be reused after removing from the face. The mask must be replaced at least every 3 hours. Always wash hands before fitting and after removing the mask. Note! A surgical mask does not protect the wearer against airborne infectious agents (coronavirus, TB, viral haemorrhagic fever, measles, varicella, SARS, avian influenza, etc.). In such cases, it is advisable to wear a respiratory protective mask rated FFP2 (complies with european standards) or N95 (complies with american standards). Packaging and labelling: Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging: One (1) unit in a protective packaging. EN 14683 type IIR or ASTM level 2 is mentioned on the box Manufacturer name and address. ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, MFD and expiry date. Word ‘non-sterile, single use, disposable.’ Comes with instructions for use.
141 | PPEs | SAFETY_GOGGLES | Goggles, panoramic, regular nose, indirect ventilation | Goggles, panoramic, regular nose, indirect ventilation | Intended purpose of use: Protective eyewear to prevent particulates, water or chemicals from striking the eyes. In haemorrhagic fever contexts it is recommended to use safety wrap around goggles. . General Description: Goggle, panoramic, regular nose, indirect ventilation. Good seal with the skin of the face, flexible frame to easily fit with all face contours with even pressure, enclose eyes and the surrounding areas, accommodate wearers with prescription glasses, clear plastic lens with fog- and scratch-resistant treatments. Adjustable band to secure firmly so as not to become loose during clinical activity. Indirect venting to avoid fogging. Reusable (provided appropriate arrangements for decontamination are in place) Technical specifications: Good seal with the skin of the face. Reusable. Markings written on the goggles (frame and lenses) according to the EN 166 specifications: Mechanical class = F = Minimum mechanical resistance (resistance to shocks of low-energy particles): withstands a bead of 6 mm and 0.85 g at 45m/s impact. Optic class 2 = applicable for intermittent use, for activities with medium visual requirements. K = Anti-Scratch N = Anti-Fog Frame Protection class = 3 Protection against liquid doplets. Lens: Material: mechanically strong but less flexible polycarbonate (PC) or a better chemical resistant acetate type (PA) are equivalent options. Clear lens for protection against UV light. Panoramic view 180° The optical zone for every eye must be at least 32 mm horizontally and 25 mm vertically. Replaceable. Treated against fogging and against scratching. Must be easy to disinfect and resistant to chlorine a 0.5% solution. Frame: Shape of nose bridge: regular (or alternatively wide) Covers big part of the cheek (important in VHF context: between hood and respirator). Parts that touch the skin are not allergenic or irritating. Easy to combine with personal spectacles/ corrective glasses. Made of thermoplastic elastomer. Flexible PVC frame to easily fit with all face contours with even pressure. Broad and fully adjustable headband. Adjustable band to secure firmly so as not to become loose during clinical activity. Must be easy to disinfect and resistant to chlorine à.5% solution. Indirect ventilation: a venting system that does not allow for direct contact of particles to the interior of the goggles. this is achieved by adding angled vents, (prefereably at the bottom), which face away from the front lens that the wearer looks through. Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions shall be as per current standards): PPE class 3 EN 166 : 2002 Personal eye-protection – Specification EN 167 : 2002 Personal eye-protection - optical test methods EN 168 : 2002 Personal eye-protection - non-optical test methods ANSI/ISEA Z87.1, or equivalent Instructions for use: Worn during procedures that may expose personnel to eye contamination by infectious or hazardous chemicals. The reusable goggles must be disinfected with chlorine solution and be easy to remove. The ocular should be cleaned regularly. It can be cleaned with water and soap, or cleaning fluid for glasses. In case the goggles should be cleaned with a chlorine solution, it is necessary to clean it afterwards with water and soap, and rinse it thoroughly with clear water. No traces of chlorine should be left over because these traces will quickly result in fogging. Scratched ocular should be replaced. If the elastic strap is damaged, the entire goggle needs to be replaced.Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging and labelling: Packaging: One (1) unit in a protective packaging. Manufacturer name and address. ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, MFD and expiry date | Intended purpose of use: Protective eyewear to prevent particulates, water or chemicals from striking the eyes. In haemorrhagic fever contexts it is recommended to use safety wrap around goggles. . General Description: Goggle, panoramic, regular nose, indirect ventilation. Good seal with the skin of the face, flexible frame to easily fit with all face contours with even pressure, enclose eyes and the surrounding areas, accommodate wearers with prescription glasses, clear plastic lens with fog- and scratch-resistant treatments. Adjustable band to secure firmly so as not to become loose during clinical activity. Indirect venting to avoid fogging. Reusable (provided appropriate arrangements for decontamination are in place) Technical specifications: Good seal with the skin of the face. Reusable. Markings written on the goggles (frame and lenses) according to the EN 166 specifications: Mechanical class = F = Minimum mechanical resistance (resistance to shocks of low-energy particles): withstands a bead of 6 mm and 0.85 g at 45m/s impact. Optic class 2 = applicable for intermittent use, for activities with medium visual requirements. K = Anti-Scratch N = Anti-Fog Frame Protection class = 3 Protection against liquid doplets. Lens: Material: mechanically strong but less flexible polycarbonate (PC) or a better chemical resistant acetate type (PA) are equivalent options. Clear lens for protection against UV light. Panoramic view 180° The optical zone for every eye must be at least 32 mm horizontally and 25 mm vertically. Replaceable. Treated against fogging and against scratching. Must be easy to disinfect and resistant to chlorine a 0.5% solution. Frame: Shape of nose bridge: regular (or alternatively wide) Covers big part of the cheek (important in VHF context: between hood and respirator). Parts that touch the skin are not allergenic or irritating. Easy to combine with personal spectacles/ corrective glasses. Made of thermoplastic elastomer. Flexible PVC frame to easily fit with all face contours with even pressure. Broad and fully adjustable headband. Adjustable band to secure firmly so as not to become loose during clinical activity. Must be easy to disinfect and resistant to chlorine à.5% solution. Indirect ventilation: a venting system that does not allow for direct contact of particles to the interior of the goggles. this is achieved by adding angled vents, (prefereably at the bottom), which face away from the front lens that the wearer looks through. Shelf life 3-5 years Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (submissions shall be as per current standards): PPE class 3 EN 166 : 2002 Personal eye-protection – Specification EN 167 : 2002 Personal eye-protection - optical test methods EN 168 : 2002 Personal eye-protection - non-optical test methods ANSI/ISEA Z87.1, or equivalent Instructions for use: Worn during procedures that may expose personnel to eye contamination by infectious or hazardous chemicals. The reusable goggles must be disinfected with chlorine solution and be easy to remove. The ocular should be cleaned regularly. It can be cleaned with water and soap, or cleaning fluid for glasses. In case the goggles should be cleaned with a chlorine solution, it is necessary to clean it afterwards with water and soap, and rinse it thoroughly with clear water. No traces of chlorine should be left over because these traces will quickly result in fogging. Scratched ocular should be replaced. If the elastic strap is damaged, the entire goggle needs to be replaced.Instruction for use in langauages English/ French/arabic/spanish or pictograms Packaging and labelling: Packaging: One (1) unit in a protective packaging. Manufacturer name and address. ISO 15223 CE mark (+EC REP), FDA and equivalent. Lot/batch, MFD and expiry date
142 | PPEs | COVERALLS_L_50 COVERALLS_M_50 COVERALLS_XL_50 | Coverall, L, 3b/4b/5b/6b, disposable Coverall, M, 3b/4b/5b/6b, disposable Coverall, XL, 3b/4b/5b/6b, disposable | Coverall, L, 3b/4b/5b/6b, disposable Coverall, M, 3b/4b/5b/6b, disposable Coverall, XL, 3b/4b/5b/6b, disposable | Coverall, PPE Cat. III, 3b/4b/5b/6b, disposable Coverall, protection, Category III, type 3b (liquid tight) Coverall, protection, Category III, type 4b (spray tight) Coverall, protection, Category III, type 5b (particle tight, limited splash-proof) Coverall, protection, Category III, type 6b (particle tight, limited splash-proof) General description: Spray/aerosol-penetration resistant, biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. Protective clothing (PPE) category III Technical specifications: Elasticated hood around face. Elasticated cuffs and ankles. Protective seams providing barrier equal to fabric. Zipper at the front with re-sealable flap protecting leakage through seams. Each coverall has a stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White/ yellow/orange Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa Non-sterile Single Use, disposable sizes, indicative only, cm: M: 166-174 ; L: 174-182; XL: 182-190 Shelf life 3-5 years Intended use: Disposable liquid-tight biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. After use, the coverall must be disposed of in a biohazard waste container, collected and destroyed. This is also applicable if the coverall is damaged (perforation, etc.). Please refer to WHO publication Safe management of waste for Health Care. Packaging and labelling: Instructions for use in English/Arabic/spanish/French, storage and disposal instructions, warnings, pictograms as applicable Packaging: One (1) unit in a plastic bag. Labelling on primary packaging (one unit) must include: - Name and/or trademark of the manufacturer - Manufacturer address - Manufacturer's product reference (product code) - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) - Words ‘Non-sterile, disposable, single use’ - CE mark (+ EC REP), FDA, and equivalent -Harmonized pictograms to indicate compliance to the type (3B, 4B, 5B, 6B or combinations) and applicable standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (all submissions as per current revision of standards) EN 14605:2005+A1:2009 Performance requirements for clothing with liquid-tight (Type 3) or spray-tight (Type 4) EN ISO 13982-1:2004+A1:2010 Performance requirements for protective clothing providing protection against airborne solid particulates (Type 5) EN 13034:2005+A1:2009 Performance requirements for protective clothing offering limited protective performance against liquid chemicals (Type 6) EN 14126:2003 Protective clothing against infective agent (as per type 3b/4b/5b/6b) Type 4 applicable standards: Category III • EN 14605 • EN ISO 13892-1 • EN 13034 • EN 14126 • EN 1073-2 • EN 1149-5 Type 5/6 applicable standards: Category III • EN 13034 • EN ISO 13892-1 • EN 1149-5 • EN 14126 • EN 1073-2 | Coverall, PPE Cat. III, 3b/4b/5b/6b, disposable Coverall, protection, Category III, type 3b (liquid tight) Coverall, protection, Category III, type 4b (spray tight) Coverall, protection, Category III, type 5b (particle tight, limited splash-proof) Coverall, protection, Category III, type 6b (particle tight, limited splash-proof) General description: Spray/aerosol-penetration resistant, biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. Protective clothing (PPE) category III Technical specifications: Elasticated hood around face. Elasticated cuffs and ankles. Protective seams providing barrier equal to fabric. Zipper at the front with re-sealable flap protecting leakage through seams. Each coverall has a stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White/ yellow/orange Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa Non-sterile Single Use, disposable sizes, indicative only, cm: M: 166-174 ; L: 174-182; XL: 182-190 Shelf life 3-5 years Intended use: Disposable liquid-tight biohazard-protective coverall, for use in EVD patient-isolation units for infection prevention and control against viral penetration. After use, the coverall must be disposed of in a biohazard waste container, collected and destroyed. This is also applicable if the coverall is damaged (perforation, etc.). Please refer to WHO publication Safe management of waste for Health Care. Packaging and labelling: Instructions for use in English/Arabic/spanish/French, storage and disposal instructions, warnings, pictograms as applicable Packaging: One (1) unit in a plastic bag. Labelling on primary packaging (one unit) must include: - Name and/or trademark of the manufacturer - Manufacturer address - Manufacturer's product reference (product code) - Type of product and main characteristics - If the packaging is not transparent, it must bear a diagram showing the essential parts of the product - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) - Information for handling, if applicable (or equivalent harmonised symbol) - Words ‘Non-sterile, disposable, single use’ - CE mark (+ EC REP), FDA, and equivalent -Harmonized pictograms to indicate compliance to the type (3B, 4B, 5B, 6B or combinations) and applicable standards Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards (all submissions as per current revision of standards) EN 14605:2005+A1:2009 Performance requirements for clothing with liquid-tight (Type 3) or spray-tight (Type 4) EN ISO 13982-1:2004+A1:2010 Performance requirements for protective clothing providing protection against airborne solid particulates (Type 5) EN 13034:2005+A1:2009 Performance requirements for protective clothing offering limited protective performance against liquid chemicals (Type 6) EN 14126:2003 Protective clothing against infective agent (as per type 3b/4b/5b/6b) Type 4 applicable standards: Category III • EN 14605 • EN ISO 13892-1 • EN 13034 • EN 14126 • EN 1073-2 • EN 1149-5 Type 5/6 applicable standards: Category III • EN 13034 • EN ISO 13892-1 • EN 1149-5 • EN 14126 • EN 1073-2
143 | PPEs | Face shield | Faceshield, fog-resistant, disposable | Faceshield, fog-resistant, disposable | Intended purpose of use: To prevent the user from contamination of mucous membranes of the mouth and nose by infective droplets Description: Made of clear plastic and providing good visibility to both the wearer and the patient. Adjustable band to attach firmly around the head and fit snuggly against the forehead, fog resistant. Completely cover the sides and length of the face. Technical specifications: Material, shield part: clear polycarbonate, thickness approx. 0.3 mm Size shield, down from headband, approx.: 26 x 34 cm (w x h) Adjustable length headband, integrated with the shield Width headband, approx.: 3 cm Front part of the headband is foam padded (length approx. 25 cm) Shield is anti-fog Outside is coated to prevent glare from reflection Shelf life 3-5 years Instructions for use: When worn together with a properly fitted surgical mask, face shields offer protection from contamination to others, as well as to the wearer. Where face masks are not available, face shields may be used as an alternative but with limited effectiveness. Instruction for use in langauages English/ French/arabic/spanish or pictograms Cleaning and disinfection procedure (Mass disinfection of multiple face shields or individual shield cleaning as applicable) Packaging and labelling: One (1) unit in a protective packaging. Manufacturer name and address. Lot/batch, MFD and expiry date. Product reference(s) of item(s) offered Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards ( submissions as per latest revision of standards:) EN 166 (if reusable) ANSI/ISEA Z87.1 (if reusable) Or alternative equivalent set of standards | Intended purpose of use: To prevent the user from contamination of mucous membranes of the mouth and nose by infective droplets Description: Made of clear plastic and providing good visibility to both the wearer and the patient. Adjustable band to attach firmly around the head and fit snuggly against the forehead, fog resistant. Completely cover the sides and length of the face. Technical specifications: Material, shield part: clear polycarbonate, thickness approx. 0.3 mm Size shield, down from headband, approx.: 26 x 34 cm (w x h) Adjustable length headband, integrated with the shield Width headband, approx.: 3 cm Front part of the headband is foam padded (length approx. 25 cm) Shield is anti-fog Outside is coated to prevent glare from reflection Shelf life 3-5 years Instructions for use: When worn together with a properly fitted surgical mask, face shields offer protection from contamination to others, as well as to the wearer. Where face masks are not available, face shields may be used as an alternative but with limited effectiveness. Instruction for use in langauages English/ French/arabic/spanish or pictograms Cleaning and disinfection procedure (Mass disinfection of multiple face shields or individual shield cleaning as applicable) Packaging and labelling: One (1) unit in a protective packaging. Manufacturer name and address. Lot/batch, MFD and expiry date. Product reference(s) of item(s) offered Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Applicable standards ( submissions as per latest revision of standards:) EN 166 (if reusable) ANSI/ISEA Z87.1 (if reusable) Or alternative equivalent set of standards
144 | PPEs | Surgical cap | Cap,surgical,bouffant,non-woven | Cap,surgical,bouffant,non-woven | Intended use: protection for head from contamination by infectious agents during surgical proceudre Description: A non-sterile cap to completely cover the hair and is intended to be worn by surgical staff during an operation to protect both the patient and themselves from the transfer of microorganisms, body fluids, and particulate material. It is an elasticated cap that is universal size for adults. Technical specifications Material : non-woven (polypropylene) One size (Ø 53 - 60 cm) Size and design shall accommodate and completely cover all hair types Elastic: latex-free elastic Lightweight 10 to 30g Impervious to liquid and non absorbent Breathable Non-sterile, for single use Shelf life 3-5 years shall be clean with bioburden less than 30 cfu/g Free from loose particles and fibrers. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Packaging and labelling: manufacturer name address, mfg date and expiry Instructions for use in English/Arabic, Spanish and French or pictograms | Intended use: protection for head from contamination by infectious agents during surgical proceudre Description: A non-sterile cap to completely cover the hair and is intended to be worn by surgical staff during an operation to protect both the patient and themselves from the transfer of microorganisms, body fluids, and particulate material. It is an elasticated cap that is universal size for adults. Technical specifications Material : non-woven (polypropylene) One size (Ø 53 - 60 cm) Size and design shall accommodate and completely cover all hair types Elastic: latex-free elastic Lightweight 10 to 30g Impervious to liquid and non absorbent Breathable Non-sterile, for single use Shelf life 3-5 years shall be clean with bioburden less than 30 cfu/g Free from loose particles and fibrers. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or For medical devices used as PPE, both PPE Regulation EU 2016/425 and EU 2017/745 (MDR) apply. For PPE, only the PPE Regulation EU 2016/425 (MDR) applies with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Packaging and labelling: manufacturer name address, mfg date and expiry Instructions for use in English/Arabic, Spanish and French or pictograms
145 | Medical supplies | AIRWAYGUEDEL0 | Airway, Guedel, paediatric, size 0 (50mm) | Airway, Guedel, paediatric, size 0 (50mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for maintenance of a clear airway through the buccal cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 50mm±2.5mm, ISO size 5, Inner diameter 3.5mm minimum, Color code:Blue, paediatric (Baby) Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Shelf life 5 years Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for maintenance of a clear airway through the buccal cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 50mm±2.5mm, ISO size 5, Inner diameter 3.5mm minimum, Color code:Blue, paediatric (Baby) Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Shelf life 5 years Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
146 | Medical supplies | AIRWAYGUEDEL00C | Airway, Guedel, neonate, size 00 (35-40mm) | Airway, Guedel, neonate, size 00 (35-40mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 40mm±2.5mm, ISO size 4, Inner diameter 3mm minimum, Color code:Pink, neonatal Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 40mm±2.5mm, ISO size 4, Inner diameter 3mm minimum, Color code:Pink, neonatal Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
147 | Medical supplies | AIRWAYGUEDEL1 | Airway, Guedel, paediatric, size 1 (60mm) | Airway, Guedel, paediatric, size 1 (60mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 60mm±2.5mm, ISO size 6, Inner diameter 4mm minimum, Color code:black, small child Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 1522CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 60mm±2.5mm, ISO size 6, Inner diameter 4mm minimum, Color code:black, small child Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 1522CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
148 | Medical supplies | AIRWAYGUEDEL2 | Airway, Guedel, child, size 2 (70mm) | Airway, Guedel, child, size 2 (70mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 70mm with tolerance +5.0 / -2.5mm, ISO size 7, Inner diameter 4mm minimum, Color code:white, Child Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 70mm with tolerance +5.0 / -2.5mm, ISO size 7, Inner diameter 4mm minimum, Color code:white, Child Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
149 | Medical supplies | AIRWAYGUEDEL3 | Airway, Guedel, adolescent, size 3 (80mm) | Airway, Guedel, adolescent, size 3 (80mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 80mm±5.0mm, ISO size 8, Inner diameter 4.5mm minimum, Color code:green, Adolescent Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 80mm±5.0mm, ISO size 8, Inner diameter 4.5mm minimum, Color code:green, Adolescent Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
150 | Medical supplies | AIRWAYGUEDEL4 | Airway, Guedel, adult, size 4 (90mm) | Airway, Guedel, adult, size 4 (90mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 90mm±5.0mm, ISO size 9, Inner diameter 4.5mm minimum, Color code:yellow, Adult small Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 90mm±5.0mm, ISO size 9, Inner diameter 4.5mm minimum, Color code:yellow, Adult small Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Supplied with: Manufacturer's instruction for use. Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
151 | Medical supplies | AIRWAYGUEDEL5 | Airway, Guedel, adult, size 5 (100mm) | Airway, Guedel, adult, size 5 (100mm) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 100mm±5.0mm, ISO size 10, Inner diameter 5mm minimum, Color code:red, Adult Large Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to maintain gas pathway through the oral cavity and the pharynx by preventing blockage by the tongue. Product description: Oropharyngeal airway/cannula, Guedel type, Color coded as per size Size: Length 100mm±5.0mm, ISO size 10, Inner diameter 5mm minimum, Color code:red, Adult Large Material: Medical grade plastic such as Polyethylene vinyl acetate (EVA) / Polyvinyl chloride (PVC) phathalate free Semi-rigid, clear and transparent. Surface shall be marked with the size and maufacturer name. Parts: curved and flattened distal (or pharyngeal) end with oval aperture Straight and reinforced proximal (or buccal) end​ to function as bite block proximal flange (collar) intended to be external to the teeth or gums. Atraumatic tip. Smoothly rounded edges. sterile, Single use only. The product shall be resistant to steam sterilization at 121 degC or 134 degC. Instructions for use: For maintenance of a clear oral airway by preventing blockage by the tongue. The choice of an oral-pharyngeal cannula must take into account the size and anatomical configuration of the oropharynx, as these can vary greatly from one person to another. Packaging and labelling: One (1) airway Guedel in a plastic bag. Symbols used according ISO 15223 CE mark Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : EN ISO 5364: 2016: Anaesthetic and respiratory equipment - Oropharyngeal airways (5th edition) ISO 18190: 2016: Anaesthetic and respiratory equipment - General requirements for airways and related equipment ISO10993-1: Biological safety testing
152 | Medical supplies | AMNIOHOOK_1 | Amniohook, plastic, sterile, single use, P1 | Amniohook, plastic, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Instrument ended by a hook used to rupture the amniotic membrane to induce labour Product description: Device provides a safe, comfortable method to rupture the amniotic membrane the first time, virtually every time. It safely “snags” the membrane, eliminating the need to poke or puncture. Device with traditional, straight amniotomes made of non-glare plastic, with flat handle for maneuverability and molded sides to ensure a nonslip grip and precise control. Material: medical grade plastic such as polypropylene Length: 25cm (approx.) Rounded tip and protected point. Latex-free. Sterile, individual packages. For single use Instructions for use: Used for rupturing the amniotic sac to induce labour. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Symbols used according ISO 15223 CE mark with notified body number Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Instrument ended by a hook used to rupture the amniotic membrane to induce labour Product description: Device provides a safe, comfortable method to rupture the amniotic membrane the first time, virtually every time. It safely “snags” the membrane, eliminating the need to poke or puncture. Device with traditional, straight amniotomes made of non-glare plastic, with flat handle for maneuverability and molded sides to ensure a nonslip grip and precise control. Material: medical grade plastic such as polypropylene Length: 25cm (approx.) Rounded tip and protected point. Latex-free. Sterile, individual packages. For single use Instructions for use: Used for rupturing the amniotic sac to induce labour. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Symbols used according ISO 15223 CE mark with notified body number Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : ISO10993-1: Biological safety testing
153 | Medical supplies | AMNIOHOOK_100 | Amniohook, plastic, sterile, single use, P100 | Amniohook, plastic, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Instrument ended by a hook used to rupture the amniotic membrane to induce labour Product description: Device provides a safe, comfortable method to rupture the amniotic membrane the first time, virtually every time. It safely “snags” the membrane, eliminating the need to poke or puncture. Device with traditional, straight amniotomes made of non-glare plastic, with flat handle for maneuverability and molded sides to ensure a nonslip grip and precise control. Material: medical grade plastic such as polypropylene Length: 25cm (approx.) Rounded tip and protected point. Latex-free. Sterile, individual packages. For single use Instructions for use: Used for rupturing the amniotic sac to induce labour. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packed and then into pack of 100 in the secondary package Symbols used according ISO 15223 CE mark with notified body number Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : ISO10993-1: Biological safety testing | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Instrument ended by a hook used to rupture the amniotic membrane to induce labour Product description: Device provides a safe, comfortable method to rupture the amniotic membrane the first time, virtually every time. It safely “snags” the membrane, eliminating the need to poke or puncture. Device with traditional, straight amniotomes made of non-glare plastic, with flat handle for maneuverability and molded sides to ensure a nonslip grip and precise control. Material: medical grade plastic such as polypropylene Length: 25cm (approx.) Rounded tip and protected point. Latex-free. Sterile, individual packages. For single use Instructions for use: Used for rupturing the amniotic sac to induce labour. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packed and then into pack of 100 in the secondary package Symbols used according ISO 15223 CE mark with notified body number Primary pack labelling shall be as per the device classification and regulatory compliance applied to the device. Supplied with: Manufacturer's instruction for use. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (as per latest revision): Must comply with the following standards : ISO10993-1: Biological safety testing
154 | Medical supplies | ANTISHOCKGARMENT | Non-pneumatic anti-shock garment | Non-pneumatic anti-shock garment | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: First aid to treat the hypovolemic shock that may arise during postpartum haemorrhage (PPH) Product description: The non-pneumatic anti-shock garment (NASG) is used to manage the uncontrollable postpartum haemorrhage (PPH) and to keep women alive until they can get the treatment they need. Lightweight, flexible and comfortable for the wearer. Reusable No metal parts: safe for X-rays and MRIs. Material: Made of a lightweight neoprene garment and Velcro. Shall be free from hazardous chemicals such as azocolorants, shall not cause irritation and allergy to users. The NASG must allow perineal access so that examinations and vaginal procedures can be performed without it being removed. Shape: - Made in the shape of trousers divided into five or six segments for ease of application to different parts of lower body below the diaphragm with Velcro fasteners. - The segment 1 must consist of 2 pieces to be used in each ankles, - The segment 2 must consist of 2 pieces to be used in each calf, - The segment 3 must consist of 2 pieces to be used in each thigh, - The segment 4 must consist of 1 piece to be used around the pelvis, - The segment 5 and/or 6 must consist of 1 piece to be used aground the abdomen/umbilicus. This segment must apply extra compression with a small foam ball. Sizes: Required in 3 different sizes (small, medium, and large). NOTE! Please verify the sizes supplied with the vendor before placing the PO. - The garment must be able to apply 30 to 50 mm Hg of pressure to the lower body pressure. - Life time: At least 40 uses or more based on wash Testing, and Stretch and Recovery Testing washable in washer machine. Instructions for use: Intended for women suffering from uncontrollable postpartum haemorrhage (PPH), to control the bleeding, reverse the shock, and stabilize the patient for safe transport to a comprehensive obstetric care facility. The NASG applies pressure to the lower body and abdomen, thereby stabilizing vital signs and resolving hypovolemic shock. When fitted correctly, the NASG reverses the shock by returning blood to the essential organs - heart, lungs, and brain. 1. Place the NASG under the woman; 2. Close the segments 1 tightly around the ankles; 3. Close the segments 2 tightly around each calf; 4. Close the segments 3 tightly around each thigh, leave knees free; 5. Close the segment 4 around pelvis with lower edge at level of pubic bone; 6. Close the segment 5 with pressure ball over the umbilicus and finally, close the NASG with segment 6; The segments 1, 2, 3 can be applied by two persons simultaneously, but segments 4, 5 and 6 should only be applied by one. Make sure the woman can breathe normally after close the NASG. It is extremely important not to remove the NASG before a woman receives IV fluids, blood and before all vital signs are restored. Early removal can be dangerous and even fatal. For cleaning: Mix a 0,05% bleach solution (as recommended by manufacturer), immerse the NASG in that solution and soak for 10min. Scrub to remove tissue and particles. Wash and rinse with clean and cool water. Squeeze and hang in the sun to dry it. Supplied with: Manufacture's instructions for use in English, French and Spanish. The instructions for use, removal, cleaning and folding shall be provided by a QR code printed on the product which allows users to follow guidelines at different locations. Packaging and labelling: Three NASG, one of each size, packaged individually with full instructions in English, French and Spanish. Labelling on the primary packaging: - Name and/or trade mark, and address of the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). -QR code for instructions for use printed on the product and on the primary package Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards ( as per the latest revisions available): Must comply with following standards ISO 13485 Medical devices -Quality management systems ISO 14971Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: Not use of PVC polymer Phthalates free | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: First aid to treat the hypovolemic shock that may arise during postpartum haemorrhage (PPH) Product description: The non-pneumatic anti-shock garment (NASG) is used to manage the uncontrollable postpartum haemorrhage (PPH) and to keep women alive until they can get the treatment they need. Lightweight, flexible and comfortable for the wearer. Reusable No metal parts: safe for X-rays and MRIs. Material: Made of a lightweight neoprene garment and Velcro. Shall be free from hazardous chemicals such as azocolorants, shall not cause irritation and allergy to users. The NASG must allow perineal access so that examinations and vaginal procedures can be performed without it being removed. Shape: - Made in the shape of trousers divided into five or six segments for ease of application to different parts of lower body below the diaphragm with Velcro fasteners. - The segment 1 must consist of 2 pieces to be used in each ankles, - The segment 2 must consist of 2 pieces to be used in each calf, - The segment 3 must consist of 2 pieces to be used in each thigh, - The segment 4 must consist of 1 piece to be used around the pelvis, - The segment 5 and/or 6 must consist of 1 piece to be used aground the abdomen/umbilicus. This segment must apply extra compression with a small foam ball. Sizes: Required in 3 different sizes (small, medium, and large). NOTE! Please verify the sizes supplied with the vendor before placing the PO. - The garment must be able to apply 30 to 50 mm Hg of pressure to the lower body pressure. - Life time: At least 40 uses or more based on wash Testing, and Stretch and Recovery Testing washable in washer machine. Instructions for use: Intended for women suffering from uncontrollable postpartum haemorrhage (PPH), to control the bleeding, reverse the shock, and stabilize the patient for safe transport to a comprehensive obstetric care facility. The NASG applies pressure to the lower body and abdomen, thereby stabilizing vital signs and resolving hypovolemic shock. When fitted correctly, the NASG reverses the shock by returning blood to the essential organs - heart, lungs, and brain. 1. Place the NASG under the woman; 2. Close the segments 1 tightly around the ankles; 3. Close the segments 2 tightly around each calf; 4. Close the segments 3 tightly around each thigh, leave knees free; 5. Close the segment 4 around pelvis with lower edge at level of pubic bone; 6. Close the segment 5 with pressure ball over the umbilicus and finally, close the NASG with segment 6; The segments 1, 2, 3 can be applied by two persons simultaneously, but segments 4, 5 and 6 should only be applied by one. Make sure the woman can breathe normally after close the NASG. It is extremely important not to remove the NASG before a woman receives IV fluids, blood and before all vital signs are restored. Early removal can be dangerous and even fatal. For cleaning: Mix a 0,05% bleach solution (as recommended by manufacturer), immerse the NASG in that solution and soak for 10min. Scrub to remove tissue and particles. Wash and rinse with clean and cool water. Squeeze and hang in the sun to dry it. Supplied with: Manufacture's instructions for use in English, French and Spanish. The instructions for use, removal, cleaning and folding shall be provided by a QR code printed on the product which allows users to follow guidelines at different locations. Packaging and labelling: Three NASG, one of each size, packaged individually with full instructions in English, French and Spanish. Labelling on the primary packaging: - Name and/or trade mark, and address of the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). -QR code for instructions for use printed on the product and on the primary package Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards ( as per the latest revisions available): Must comply with following standards ISO 13485 Medical devices -Quality management systems ISO 14971Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: Not use of PVC polymer Phthalates free
155 | Medical supplies | BIOHAZARDBAG_50 | Bag, Biohazard, yellow, 50 liters, P50 | Bag, Biohazard, yellow, 50 liters, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A non-sterile bag intended to collect and contain disposable materials/devices that need to be autoclaved before disposal. Product description: Plastic bag that is autoclavable & impervious to microorganisms Material: High molecular weight, high density (HMHD) Polythene or Polypropylene (PP). Colour: Yellow, Used for pathological waste such as tissues or body parts or other solid infectious waste such as gauze, dressings etc. Pattern: Imprinted in black with English language text 'Biohazard' and the 'black tri-sickle logo' according U+2623 both sides of the bag. Empty field for an end-user to add name and address. Size: approx. 600×800 mm Capacity: approx. 50 L Thickness: minimum 0.050 mm With metal closure strip, one per bag Bags resist punctures, tears, and leaks: Dart Impact Strength >530g and Elmendorf Tear Resistance 200g in MD and 980g in TD. Leak proof bottom seal for added safety. Bag can be autoclaved prior to the disposal in temperatures +121°C up to +135 °C. Steam autoclave indicator on the bag (During autoclaving the color changes providing proof that bag has been exposed to high temperatures for sterilisation) Shelf life 5 yrs Supplied with: N/A Accessories / Spare Parts / Consumables (available but not supplied): N/A Instructions for use: Instructions and warnings for use are printed onto the bag. Not to be used if product or packaging has been visually compromised. The bag must not be completely closed when autoclaving. Printed with a sterilisation patch that changes color when subject to steam sterilization. Storage conditions to be specified Packaging and Labelling: One (1) roll or box of fifty (50) individual bags. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not a medical device Safety and Product Standards: ASTM D1922 - Elmendorf Tear Resistance . ASTM 1709- Dart Impact Strength Environmental Requirements: ISO 14001: Environmental management systems Instructions to Country offices: Handling, storage, and transportation of health-care waste at https://www.who.int/water_sanitation_health/medicalwaste/061to076.pdf | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A non-sterile bag intended to collect and contain disposable materials/devices that need to be autoclaved before disposal. Product description: Plastic bag that is autoclavable & impervious to microorganisms Material: High molecular weight, high density (HMHD) Polythene or Polypropylene (PP). Colour: Yellow, Used for pathological waste such as tissues or body parts or other solid infectious waste such as gauze, dressings etc. Pattern: Imprinted in black with English language text 'Biohazard' and the 'black tri-sickle logo' according U+2623 both sides of the bag. Empty field for an end-user to add name and address. Size: approx. 600×800 mm Capacity: approx. 50 L Thickness: minimum 0.050 mm With metal closure strip, one per bag Bags resist punctures, tears, and leaks: Dart Impact Strength >530g and Elmendorf Tear Resistance 200g in MD and 980g in TD. Leak proof bottom seal for added safety. Bag can be autoclaved prior to the disposal in temperatures +121°C up to +135 °C. Steam autoclave indicator on the bag (During autoclaving the color changes providing proof that bag has been exposed to high temperatures for sterilisation) Shelf life 5 yrs Supplied with: N/A Accessories / Spare Parts / Consumables (available but not supplied): N/A Instructions for use: Instructions and warnings for use are printed onto the bag. Not to be used if product or packaging has been visually compromised. The bag must not be completely closed when autoclaving. Printed with a sterilisation patch that changes color when subject to steam sterilization. Storage conditions to be specified Packaging and Labelling: One (1) roll or box of fifty (50) individual bags. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not a medical device Safety and Product Standards: ASTM D1922 - Elmendorf Tear Resistance . ASTM 1709- Dart Impact Strength Environmental Requirements: ISO 14001: Environmental management systems Instructions to Country offices: Handling, storage, and transportation of health-care waste at https://www.who.int/water_sanitation_health/medicalwaste/061to076.pdf
156 | Medical supplies | BLOODGIVINGSET | Blood-giving set | Blood-giving set | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for intravascular, gravity-controlled transfusion of blood or blood components. Product description: Blood-giving set, also known as blood transfusion set, is used for the transfusion of blood and blood products. Components: Flexible tubing, perforator (spike) with air vent, drip-counter chamber, blood filter, flow regulator, injector port. All materials must be haemocompatible based on tests selected for the nature and duration contact. Shall be free from residual migrating chemicals. Material: - Tube: PVC-based material - Perforator: stainless steel, encased in polyamide or moulded from ABS or POM (polyacetal). - Air inlet: PVC - Drip-counting chamber: PVC - Blood filter: polyamide (fibres of PA) - Flow regulator: polystyrene, ABS - Injection port: rubber latex fee Tube: - Transparent (permits air bubble detection), supple, resistant to kinking. - Length : approx.: 50cm ( set has an overall length of 170cm). - Ends in a male Luer lock or slip connector, with or without a bung; permits its connection to another intravenous system. Perforator: - A hollow device situated at the proximal end of the set; permits the set to be connected to a bottle or flexible blood bag. - Comprised of a needle or a tapering tube mounted on a base; ensure a good seal between the set and the blood container. Air inlet: With integrated bacteriological filter. Drop counting chamber: Transparent drip chamber, calibrated at 20 drops/ml, with 15-20μm fluid filter. flow regulator: Smooth roller clamp to facilitate easy, safe control and adjustment of fluid rates Sterile and single use non-pyrogenic sterilized by ETO or Gamma radiation. Instructions for use: Used for the transfusion of blood and blood products Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. CE Mark with Notified Body number Supplied with: Manufacture's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Blood-giving set in an individual sterilised peel pack with information for use Secondary packaging: Protected unit. One (1) box of 50 giving set. Labelling on the primary packaging: - shall be compliant as per the regulation applied to the device -Name and/or trademark and address of the manufacturer. -Manufacturer's product reference. - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) - The words "for single use" (or equivalent harmonised symbol). - sterile symbol and sterilization method (or harmonized symbol) -"Pyrogenic free fluid pathway" , or equivalent harmonised symbol - "Do not use if there is any visible sign of deterioration " , or equivalent harmonised symbol - " Contains phthalate (DEHP) "or equivalent harmonised symbol - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 and ISO 3826-2 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (submissions as per the latest revision of standards below): Must comply with the following standards: ISO 1135-4: 2015 Transfusion equipment for medical use - Part 4: Transfusion sets for single use, gravity feed (edition 6) ISO 13485 Medical devices -- Quality management systems ISO 14971 Medical Devices- Application of risk management to medical devices EN ISO 3826-1:2003: Plastics collapsible containers for human blood and blood components Part 1: Conventional containers ISO 3826-3:2006 Plastics collapsible containers for human blood and blood components — Part 3: Blood bag systems with integrated features ISO 3826-2:2008 Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets ISO 10993-1 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-4:2002+ Amd 1:2006: Biological evaluation of medical devices — Part 4: Selection of tests for interactions with blood ISO 10993-11: 2006: Biological evaluation of medical devices — Part 11: Tests for systemic toxicity EN 15986:2011 Symbol for use in the labelling of medical devices — Requirements for labelling of medical devices containing phthalates | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for intravascular, gravity-controlled transfusion of blood or blood components. Product description: Blood-giving set, also known as blood transfusion set, is used for the transfusion of blood and blood products. Components: Flexible tubing, perforator (spike) with air vent, drip-counter chamber, blood filter, flow regulator, injector port. All materials must be haemocompatible based on tests selected for the nature and duration contact. Shall be free from residual migrating chemicals. Material: - Tube: PVC-based material - Perforator: stainless steel, encased in polyamide or moulded from ABS or POM (polyacetal). - Air inlet: PVC - Drip-counting chamber: PVC - Blood filter: polyamide (fibres of PA) - Flow regulator: polystyrene, ABS - Injection port: rubber latex fee Tube: - Transparent (permits air bubble detection), supple, resistant to kinking. - Length : approx.: 50cm ( set has an overall length of 170cm). - Ends in a male Luer lock or slip connector, with or without a bung; permits its connection to another intravenous system. Perforator: - A hollow device situated at the proximal end of the set; permits the set to be connected to a bottle or flexible blood bag. - Comprised of a needle or a tapering tube mounted on a base; ensure a good seal between the set and the blood container. Air inlet: With integrated bacteriological filter. Drop counting chamber: Transparent drip chamber, calibrated at 20 drops/ml, with 15-20μm fluid filter. flow regulator: Smooth roller clamp to facilitate easy, safe control and adjustment of fluid rates Sterile and single use non-pyrogenic sterilized by ETO or Gamma radiation. Instructions for use: Used for the transfusion of blood and blood products Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. CE Mark with Notified Body number Supplied with: Manufacture's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Blood-giving set in an individual sterilised peel pack with information for use Secondary packaging: Protected unit. One (1) box of 50 giving set. Labelling on the primary packaging: - shall be compliant as per the regulation applied to the device -Name and/or trademark and address of the manufacturer. -Manufacturer's product reference. - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) - The words "for single use" (or equivalent harmonised symbol). - sterile symbol and sterilization method (or harmonized symbol) -"Pyrogenic free fluid pathway" , or equivalent harmonised symbol - "Do not use if there is any visible sign of deterioration " , or equivalent harmonised symbol - " Contains phthalate (DEHP) "or equivalent harmonised symbol - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 and ISO 3826-2 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (submissions as per the latest revision of standards below): Must comply with the following standards: ISO 1135-4: 2015 Transfusion equipment for medical use - Part 4: Transfusion sets for single use, gravity feed (edition 6) ISO 13485 Medical devices -- Quality management systems ISO 14971 Medical Devices- Application of risk management to medical devices EN ISO 3826-1:2003: Plastics collapsible containers for human blood and blood components Part 1: Conventional containers ISO 3826-3:2006 Plastics collapsible containers for human blood and blood components — Part 3: Blood bag systems with integrated features ISO 3826-2:2008 Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets ISO 10993-1 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-4:2002+ Amd 1:2006: Biological evaluation of medical devices — Part 4: Selection of tests for interactions with blood ISO 10993-11: 2006: Biological evaluation of medical devices — Part 11: Tests for systemic toxicity EN 15986:2011 Symbol for use in the labelling of medical devices — Requirements for labelling of medical devices containing phthalates
157 | Medical supplies | BOX.SAFETY_1 | Safety box, disposal of used syringes & needles, 5L, P1 | Safety box, disposal of used syringes & needles, 5L, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe disposal container for safe handling and disposing of 100 (Auto Disable) syringes. Material: Recycled cardboard with plastic film, strong (waterproof and puncture-proof.) Thickness of walls: 1.1 - 4.4mm Container with 5 L capacity accommodating approximately 100 syringes Dimensions: approx. 284 x 162 x 129mm when set up. Notching system for hermetic sealing (inferior and superior parts) Diameter of syringe insert hole: 30-38mm Features: Boxes should be equipped with a carrying handle and have directions for use and destruction printed on the box. Distributed and stored in collapsible form Single use: to be destroyed with its contents by incineration. Instructions for use: Puncture resistant containers for collecting and disposing of used disposable and auto-disable syringes, needles and other injection materials must be provided and used in all immunization activities. These containers reduce the risk posed by contaminated needles and syringes to health staff and to the general public. Supplied with: Individual box, flat-packed for assembly by the end-user+ Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit presentation: 1 (one) Safety box f. used syrgs/ndls, 5 L Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: 93/42/EEC or EU MDR Class I – Self declaration Safety & product Standards: Compliance with WHO Performance Specification E10/IC.1. ASTM F2132 - 01(2008)e1 Standard Specification for Puncture Resistance of Materials Used in Containers for Discarded Medical Needles and Other Sharps Environmental requirements: Is recommended use other polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe disposal container for safe handling and disposing of 100 (Auto Disable) syringes. Material: Recycled cardboard with plastic film, strong (waterproof and puncture-proof.) Thickness of walls: 1.1 - 4.4mm Container with 5 L capacity accommodating approximately 100 syringes Dimensions: approx. 284 x 162 x 129mm when set up. Notching system for hermetic sealing (inferior and superior parts) Diameter of syringe insert hole: 30-38mm Features: Boxes should be equipped with a carrying handle and have directions for use and destruction printed on the box. Distributed and stored in collapsible form Single use: to be destroyed with its contents by incineration. Instructions for use: Puncture resistant containers for collecting and disposing of used disposable and auto-disable syringes, needles and other injection materials must be provided and used in all immunization activities. These containers reduce the risk posed by contaminated needles and syringes to health staff and to the general public. Supplied with: Individual box, flat-packed for assembly by the end-user+ Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit presentation: 1 (one) Safety box f. used syrgs/ndls, 5 L Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: 93/42/EEC or EU MDR Class I – Self declaration Safety & product Standards: Compliance with WHO Performance Specification E10/IC.1. ASTM F2132 - 01(2008)e1 Standard Specification for Puncture Resistance of Materials Used in Containers for Discarded Medical Needles and Other Sharps Environmental requirements: Is recommended use other polymer not PVC
158 | Medical supplies | BOX.SAFETY_25 | Safety box, disposal of used syringes & needles, 5L, P25 | Safety box, disposal of used syringes & needles, 5L, P25 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For disposal of used syringes, needles and scalpel blades. Product description: Syringe disposal container for safe handling and disposing of 100 (Auto Disable) syringes. Material: Recycled cardboard with plastic film, strong (waterproof and puncture-proof.) Thickness of walls: 1.1 - 4.4mm Container with 5 L capacity accommodating approximately 100 syringes Dimensions: approx. 284 x 162 x 129mm when set up. Notching system for hermetic sealing (inferior and superior parts) Diameter of syringe insert hole: 30-38mm Features: Boxes should be equipped with a carrying handle and have directions for use and destruction printed on the box. Distributed and stored in collapsible form Single use: to be destroyed with its contents by incineration. Upon incineration, the material of construction shall not release any ozone depleting chemicals. Marking instructions on each safety box: Bio-hazard marking: marked with the international bio-hazard warning not less than 50mm diameter, printed in black or red on each of the front and back faces of the box. Pictorial instructions: without writing printed on two sides of the container showing: • How to assemble the box. • How to use the box as a container for contaminated sharps; • Syringe disposal direction (needle down). • How to close the sharps aperture when the box is full. Instructions for use: Puncture resistant containers for collecting and disposing of used disposable and auto-disable syringes, needles and other injection materials must be provided and used in all immunization activities. These containers reduce the risk posed by contaminated needles and syringes to health staff and to the general public. To be destroyed with its contents by incineration. Packaging and labelling Primary packaging: None. One box not packed for construction at the field level. Secondary packaging 25 boxes in one carton are flat-packed for ease of shipment and storage + Manufacturer's instruction for use. Recyclable cardboard is to be used.Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not a medical device Safety & product Standards (all submissions must be as per latest revision of standards): Product verification done in accordance with PQS Verification Protocol E10/SB01.1 ASTM F2132 - 01(2008)e1 Standard Specification for Puncture Resistance of Materials Used in Containers for Discarded Medical Needles and Other Sharps Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For disposal of used syringes, needles and scalpel blades. Product description: Syringe disposal container for safe handling and disposing of 100 (Auto Disable) syringes. Material: Recycled cardboard with plastic film, strong (waterproof and puncture-proof.) Thickness of walls: 1.1 - 4.4mm Container with 5 L capacity accommodating approximately 100 syringes Dimensions: approx. 284 x 162 x 129mm when set up. Notching system for hermetic sealing (inferior and superior parts) Diameter of syringe insert hole: 30-38mm Features: Boxes should be equipped with a carrying handle and have directions for use and destruction printed on the box. Distributed and stored in collapsible form Single use: to be destroyed with its contents by incineration. Upon incineration, the material of construction shall not release any ozone depleting chemicals. Marking instructions on each safety box: Bio-hazard marking: marked with the international bio-hazard warning not less than 50mm diameter, printed in black or red on each of the front and back faces of the box. Pictorial instructions: without writing printed on two sides of the container showing: • How to assemble the box. • How to use the box as a container for contaminated sharps; • Syringe disposal direction (needle down). • How to close the sharps aperture when the box is full. Instructions for use: Puncture resistant containers for collecting and disposing of used disposable and auto-disable syringes, needles and other injection materials must be provided and used in all immunization activities. These containers reduce the risk posed by contaminated needles and syringes to health staff and to the general public. To be destroyed with its contents by incineration. Packaging and labelling Primary packaging: None. One box not packed for construction at the field level. Secondary packaging 25 boxes in one carton are flat-packed for ease of shipment and storage + Manufacturer's instruction for use. Recyclable cardboard is to be used.Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance Classification:93/42/EEC or EU MDR Class I – Self declaration Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Not a medical device Safety & product Standards (all submissions must be as per latest revision of standards): Product verification done in accordance with PQS Verification Protocol E10/SB01.1 ASTM F2132 - 01(2008)e1 Standard Specification for Puncture Resistance of Materials Used in Containers for Discarded Medical Needles and Other Sharps Environmental requirements: ISO 14001
159 | Medical supplies | CANNULAIV16G | Cannula, IV short, 16G, sterile, single use | Cannula, IV short, 16G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 16G (1.7 x 45mm) grey. Flow rate approx. 180ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 16G (1.7 x 45mm) grey. Flow rate approx. 180ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
160 | Medical supplies | CANNULAIV18G | Cannula, IV short, 18G, sterile, single use | Cannula, IV short, 18G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 18G (1.3 x 45mm) green. Flow rate approx. 80ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 18G (1.3 x 45mm) green. Flow rate approx. 80ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
161 | Medical supplies | CANNULAIV20G_1 | Cannula, IV short, 20G, sterile, single use, P1 | Cannula, IV short, 20G, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult and infants use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 20G (1 x 32mm) pink. Flow rate approx. 55ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult and infants use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 20G (1 x 32mm) pink. Flow rate approx. 55ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
162 | Medical supplies | CANNULAIV20G_50 | Cannula, IV short, 20G, sterile, single use, P50 | Cannula, IV short, 20G, sterile, single use, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult and infants use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 20G (1 x 32mm) pink. Flow rate approx. 55ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for adult and infants use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 20G (1 x 32mm) pink. Flow rate approx. 55ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
163 | Medical supplies | CANNULAIV22G | Cannula, IV short, 22G, sterile, single use | Cannula, IV short, 22G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for infant and children use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 22G (0.90 x 25mm) blue. Flow rate approx. 33ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for infant and children use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 22G (0.90 x 25mm) blue. Flow rate approx. 33ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
164 | Medical supplies | CANNULAIV24G | Cannula, IV short, 24G, sterile, single use | Cannula, IV short, 24G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for children and neonates use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 24G (0.70 x 19mm) yellow. Flow rate approx. 20ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For prolonged peripheral intravenous infusions. Product description: Sterile Intravenous (IV) cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products (for children and neonates use). Components: Protecting cap, trocar, cannula, stopper, injection port, Luer lock (proximal part) (all parts fit together to form a unit). Material for trocar: Stainless steel Material for cannula: PTFE (Poly Tetra Fluoro Ethylene), FEP (Fluorinated Ethylene Propylene), PUR (Polyurethane). DEHP free, should be non thrombogenic, Non Toxic and Non Pyrogenic. Size: 24G (0.70 x 19mm) yellow. Flow rate approx. 20ml/min Colour code/external diameter: Visible at the base of cannula. Cannula with fine and tapered walls, perfectly adjusted to the needle. Distal end: straight and tapering. Proximal end base fitted with luer lock connector, 2 grapping wings (butterfly), standardized colour code depending on the diameter. Fitted with a lateral injection port. Injection port fitted with silicone anti-reflux valve, cap with colour code and chimney of luer type. Protecting cap. Stopper fitted with a hydrophobic membrane with micro-perforations to let the air flow while stopping blood, male luer connection. Triple-bevelled trocar (ensures a good skin penetration and reduces vein trauma.) Transparent female luer base. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile IV cannula to be introduced into a peripheral vein for infusing fluids, administering IV drugs, transfusing blood and blood products. Safety process: IV cannula is for single use only. Rules of asepsis must be followed when inserting IV cannula. IV cannula should not be left in situ for more than 72 hours. It should be removed immediately in case of infection sign. Supplied with: Manufacturer's instruction for use. Shall contain instructions for storage, use, safe disposal, and warning to users. Warning shall include statements such as " protection of users:To prevent stick injuries, wear gloves. NEVER recap the trocar after use and dispose of the needle into a puncture proof safety container". Must be multilingual: English, French and Spanish Packaging and labelling: Primary packaging: One (1) IV cannula in an individual sterilised peel pack, preferably transparent to see the color, else color shall be printed on the primary package. Secondary packaging: Protected unit. One (1) box of 50 IV cannulas. Labelling on primary packaging: Shall be as per regulatory requirements applied to the device. Manufacture’s name and/or trademark, product description and reference number, "sterile" with method indicated, lot number, expiry date by year and month, "for single use", and CE mark and reference number for notifying body. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions shall be as per latest revisions of the satndards): Must comply with the following standards: ISO 10555-5: Intravascular catheters -- Sterile and single-use catheters -- Part 5: Over-needle peripheral catheters ISO 80369-7 small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications (ISO 594-1 & BS EN 20594-1 are withdrawn). ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
165 | Medical supplies | CATHETERFOLE12_1 | Catheter, urethral, Foley, CH12, sterile, P1 | Catheter, urethral, Foley, CH12, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
166 | Medical supplies | CATHETERFOLE12_10 | Catheter, urethral, Foley, CH12, sterile, P10 | Catheter, urethral, Foley, CH12, sterile, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
167 | Medical supplies | CATHETERFOLE12_50 | Catheter, urethral, Foley, CH12, sterile, P50 | Catheter, urethral, Foley, CH12, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH12 Length: 30 - 40cm, external diameter 4 mm, white Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
168 | Medical supplies | CATHETERFOLE14_10 | Catheter, urethral, Foley, CH14, sterile, P10 | Catheter, urethral, Foley, CH14, sterile, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
169 | Medical supplies | CATHETERFOLE14_50 | Catheter, urethral, Foley, CH14, sterile, P50 | Catheter, urethral, Foley, CH14, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
170 | Medical supplies | CATHETERFOLE18_10 | Catheter, urethral, Foley, adult, CH18, sterile, P10 | Catheter, urethral, Foley, adult, CH18, sterile, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
171 | Medical supplies | CATHETERFOLE18_50 | Catheter, urethral, Foley, adult, CH18, sterile, P50 | Catheter, urethral, Foley, adult, CH18, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
172 | Medical supplies | CATHETERFOLEY14_1 | Catheter, urethral, Foley, CH14, sterile, P1 | Catheter, urethral, Foley, CH14, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH14 Length: 30 - 40cm, external diameter 4.7 mm, green Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
173 | Medical supplies | CATHETERFOLEY18_1 | Catheter, urethral, Foley, adult, CH18, sterile, P1 | Catheter, urethral, Foley, adult, CH18, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH18 Length: 30 - 40cm, external diameter 6 mm, red Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to thecatheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. One (1) box of 50 Foley catheters. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
174 | Medical supplies | CATHETERURE12_1 | Catheter, urethral, CH12, sterile, P1 | Catheter, urethral, CH12, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the latest revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
175 | Medical supplies | CATHETERURE12_10 | Catheter, urethral, CH12, sterile, P10 | Catheter, urethral, CH12, sterile, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
176 | Medical supplies | CATHETERURE12_100 | Catheter, ureteral, CH12, sterile, P100 | Catheter, ureteral, CH12, sterile, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
177 | Medical supplies | CATHETERURIN12_50 | Catheter, ureteral, Nelaton, CH12, sterile, P50 | Catheter, ureteral, Nelaton, CH12, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Flexible tubular device which is inserted into the bladder through the urethra to drain away urine intermittently. Not to be used as indwelling catheter. General Description Catheter, urethral, CH12, sterile, disposable Technical Specifications Urethral catheter without balloon, nelaton tip. Atraumatic, soft, rounded closed distal end with 2 lateral eyes, proximal end with cup connector allowing catheter to be connected with urine bag. Central channel for urinary drainage Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter CH12, color coded Length: 30 - 40cm, Single-use Sterile. Shelf life 3 years Initial sterilization method: Ethylene oxide gas Packaging and labelling: Primary packaging: Unit of use One (1) catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. One (1) box of 50 urethral catheters. Instructions for use Sterile urethral catheter without balloon, nelaton tip, designed to be introduced into the bladder cavity, to drain off urine, for short term catheterization. Use strict aseptic techniques during insertion of catheter. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. Precautions such as "Not suitable for indwelling urinary catheterization which requires a Foley catheter with balloon". Storage conditions, disposal instructions etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
178 | Medical supplies | CATHETERWIRE5 | Catheter, urethral, CH5, with metal guide wire, sterile | Catheter, urethral, CH5, with metal guide wire, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Catheter intended for surgical programmes in case of ureter trauma. The device is mostly used for ureter dilation, bypass of partial obstructions, upper urinary tract drainage/irrigation during surgery, stone removal, or for retrograde urography etc Product description: A flexible tube designed for introduction into the ureters. It is a radiopaque tube with a straight tip configuration. Complete radio opaque for X ray visualization. Single lumen Marking at every cm and double marking at every 5cm. Two eyes for efficient drainage. Catheter with adapter and stylet. A luer lock adaptor is provided at proximal end of the catheter for fitting to syringe. With guide wire. Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC), Polyurethane or silicone. May be coated to reduce friction. Size selected: Diameter CH5 Length: 70cm Sterile, Single use. Shelf life 3 years minimum Instructions for use: As per manufacturer's instruction for use Safety process: When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal etc. Packaging and labelling: Packaging and labelling shall be as per regulations applied to the device. Primary packaging: Unit of use One Ureteric catheter in an individual sterilised peel pack. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 20697 : Sterile drainage catheters and accessory devices for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Catheter intended for surgical programmes in case of ureter trauma. The device is mostly used for ureter dilation, bypass of partial obstructions, upper urinary tract drainage/irrigation during surgery, stone removal, or for retrograde urography etc Product description: A flexible tube designed for introduction into the ureters. It is a radiopaque tube with a straight tip configuration. Complete radio opaque for X ray visualization. Single lumen Marking at every cm and double marking at every 5cm. Two eyes for efficient drainage. Catheter with adapter and stylet. A luer lock adaptor is provided at proximal end of the catheter for fitting to syringe. With guide wire. Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC), Polyurethane or silicone. May be coated to reduce friction. Size selected: Diameter CH5 Length: 70cm Sterile, Single use. Shelf life 3 years minimum Instructions for use: As per manufacturer's instruction for use Safety process: When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal etc. Packaging and labelling: Packaging and labelling shall be as per regulations applied to the device. Primary packaging: Unit of use One Ureteric catheter in an individual sterilised peel pack. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 20697 : Sterile drainage catheters and accessory devices for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
179 | Medical supplies | CATHETERWIRE6 | Catheter, urethral, size CH6, with metal guide wire, sterile | Catheter, urethral, size CH6, with metal guide wire, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Catheter intended for surgical programmes in case of ureter trauma. The device is mostly used for ureter dilation, bypass of partial obstructions, upper urinary tract drainage/irrigation during surgery, stone removal, or for retrograde urography etc Product description: A flexible tube designed for introduction into the ureters. It is a radiopaque tube with a straight tip configuration. Complete radio opaque for X ray visualization. Single lumen Marking at every cm and double marking at every 5cm. Two eyes for efficient drainage. Catheter with adapter and stylet. A luer lock adaptor is provided at proximal end of the catheter for fitting to syringe. With guide wire. Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC), Polyurethane or silicone. May be coated to reduce friction. Size selected: Diameter CH6 Length: 70cm Sterile, Single use. Shelf life 3 years minimum Instructions for use: As per manufacturer's instruction for use Safety process: When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal etc. Packaging and labelling: Packaging and labelling shall be as per regulations applied to the device. Primary packaging: Unit of use One Ureteric catheter in an individual sterilised peel pack. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 20697 : Sterile drainage catheters and accessory devices for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Catheter intended for surgical programmes in case of ureter trauma. The device is mostly used for ureter dilation, bypass of partial obstructions, upper urinary tract drainage/irrigation during surgery, stone removal, or for retrograde urography etc Product description: A flexible tube designed for introduction into the ureters. It is a radiopaque tube with a straight tip configuration. Complete radio opaque for X ray visualization. Single lumen Marking at every cm and double marking at every 5cm. Two eyes for efficient drainage. Catheter with adapter and stylet. A luer lock adaptor is provided at proximal end of the catheter for fitting to syringe. With guide wire. Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code cup connector. Material: Polyvinyl chloride (PVC), Polyurethane or silicone. May be coated to reduce friction. Size selected: Diameter CH6 Length: 70cm Sterile, Single use. Shelf life 3 years minimum Instructions for use: As per manufacturer's instruction for use Safety process: When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal etc. Packaging and labelling: Packaging and labelling shall be as per regulations applied to the device. Primary packaging: Unit of use One Ureteric catheter in an individual sterilised peel pack. Symbols used according ISO 15223 CE Mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 20697 : Sterile drainage catheters and accessory devices for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
180 | Medical supplies | CATHETFOLEYCH16_1 | Catheter, urethral, Foley, CH16, sterile, P1 | Catheter, urethral, Foley, CH16, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter CH16, external diameter 5.3mm, Orange, Length: 30 - 40cm Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Must be multilingual. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter CH16, external diameter 5.3mm, Orange, Length: 30 - 40cm Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Must be multilingual. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
181 | Medical supplies | CATHETFOLEYCH20_1 | Catheter, urethral, Foley, CH20, sterile, P1 | Catheter, urethral, Foley, CH20, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH20 Length: 30 - 40cm, external diameter 6.7 mm, yellow Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Must be multilingual. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Indwelling catheter to drain urine from the urinary bladder. For use with adults for routine procedure and/or post operative care. Product description: Materials: Silicone-coated natural rubber latex. Shall be free from biological hazard when tested for biocompatibility. Urethral catheter with balloon: Foley catheter 2-way with nelaton tip. Central channel for urinary drainage, straight rounded distal end with lateral eyes (bladder side), proximal end with cup connector allowing catheter to be connected to urine bag. Side channel for inflating balloon, ending by a non-return valve with Luer tip connector. Diameter expressed in Charriere, French gauge. Length expressed in cm. Balloon capacity expressed in ml. Colour code / external diameter visible on connector. Size selected: Diameter: CH20 Length: 30 - 40cm, external diameter 6.7 mm, yellow Balloon capacity, 5-15ml: the amount of sterile water to be inserted into the balloon is usually printed on the side of the nozzle end. 5 ml for paediatric balloon 10 ml balloon for routine drainage 30 ml or larger balloons specifically for post-operative use Single-use. Sterile. Shelf life 3 years Initial sterilisation method: ethylene oxide gas or equivalent Instructions for use: Sterile urethral catheter device designed to be introduced into the bladder cavity, to drain off urine, instill a liquid or irrigate the bladder. If lubrication is necessary to help to introduce the catheter, the substance used must not cause physical or chemical alterations to the catheter. If necessary, use sterile water. The balloon should be inflated with water or physiological saline. The size has been chosen as being the most commonly used. Safety process: The urethral catheter is for SINGLE USE only. When inserting a urethral catheter, use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Must be multilingual. Packaging and labelling: Primary packaging: Unit of use One (1) Foley catheter in an individual sterilised peel pack. Labelling on the primary package shall be as per regulations applied to the device. Symbols as per ISO 15223. Secondary packaging: Protected unit. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 20696 Sterile urethral catheters for single use ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
182 | Medical supplies | COLOSTOMYBAG_10 | Colostomy bag, transparent, drainable, w/filter, P10 | Colostomy bag, transparent, drainable, w/filter, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to collect waste from the stoma. Product description: The colostomy pouch is one piece open system, with deodorising filter and pre-cut opening. It is a medical device that consists of a bag attached to the patient's skin by an adhesive material. The pouch is drainable from the bottom and kept closed by a removable safety clamp. Bag: Made of flexible transparent plastic with a non woven protection layer on skin side. Shape: Anatomical shaped with smooth edges. Adhesive attachment system: 12 x 12cm, with hydrocolloid skin protector, to ensure, safe continuity between bag and colostomy. Cut-to-fit ostomy bags allow the pouch to fit the shape and size of most stomas Hygenically packed for medical use. Non sterile, Single use only. Measurements: Size: Bag: Length 19 x width 14 cm with: Window/aperture: starts at Ø +/- 14mm then pre-cut marking up to 64mm. Pre-cut marking: 25/38/51 and 64mm. Filter: for absorbing gases and unpleasant smell produced by bowels. Clip: with secure closing system. Material: Bag: Polyvinyl chloride (PVC), Low density polyethylene (LDPE), Ethylene vinyl acetate (EVA). Filter: Charcoal granulates. Clip: ll Plastic Shelf life 3 years. Packing and labelling: Shall be as per regulatory compliance applied to the device. ll closing clip by 20 pieces. Each carton to be clearly identified with product name & characteristics, manufacturer addressee, lot No, number of units per carton and multilingual user manual. Instructions for use: Supplied with IFU in multilingual. Instructions for use such as cutting the adhesive disc, preparation of skin, application on the skin and removal after use. Shall include storage and disposal instructions. Shall include any warnings, precuations for use etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 8670-2: Ostomy collection bags -- Part 2: Requirements and test methods ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to collect waste from the stoma. Product description: The colostomy pouch is one piece open system, with deodorising filter and pre-cut opening. It is a medical device that consists of a bag attached to the patient's skin by an adhesive material. The pouch is drainable from the bottom and kept closed by a removable safety clamp. Bag: Made of flexible transparent plastic with a non woven protection layer on skin side. Shape: Anatomical shaped with smooth edges. Adhesive attachment system: 12 x 12cm, with hydrocolloid skin protector, to ensure, safe continuity between bag and colostomy. Cut-to-fit ostomy bags allow the pouch to fit the shape and size of most stomas Hygenically packed for medical use. Non sterile, Single use only. Measurements: Size: Bag: Length 19 x width 14 cm with: Window/aperture: starts at Ø +/- 14mm then pre-cut marking up to 64mm. Pre-cut marking: 25/38/51 and 64mm. Filter: for absorbing gases and unpleasant smell produced by bowels. Clip: with secure closing system. Material: Bag: Polyvinyl chloride (PVC), Low density polyethylene (LDPE), Ethylene vinyl acetate (EVA). Filter: Charcoal granulates, carbon filter, or any other type of filter with similar or superior characteristics. Clip: Plastic Shelf life 3 years. Packing and labelling: Shall be as per regulatory compliance applied to the device. Closing clip by 10 pieces. Each carton to be clearly identified with product name & characteristics, manufacturer addressee, lot No, number of units per carton and multilingual user manual. Instructions for use: Supplied with IFU in multilingual. Instructions for use such as cutting the adhesive disc, preparation of skin, application on the skin and removal after use. Shall include storage and disposal instructions. Shall include any warnings, precuations for use etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 8670-2: Ostomy collection bags -- Part 2: Requirements and test methods ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
183 | Medical supplies | COLOSTOMYBAG_20 | Colostomy bag, transparent, drainable, w/filter, P20 | Colostomy bag, transparent, drainable, w/filter, P20 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to collect waste from the stoma. Product description: The colostomy pouch is one piece open system, with deodorising filter and pre-cut opening. It is a medical device that consists of a bag attached to the patient's skin by an adhesive material. The pouch is drainable from the bottom and kept closed by a removable safety clamp. Bag: Made of flexible transparent plastic with a non woven protection layer on skin side. Shape: Anatomical shaped with smooth edges. Adhesive attachment system: 12 x 12cm, with hydrocolloid skin protector, to ensure, safe continuity between bag and colostomy. Cut-to-fit ostomy bags allow the pouch to fit the shape and size of most stomas Hygenically packed for medical use. Non sterile, Single use only. Measurements: Size: Bag: Length 19 x width 14 cm with: Window/aperture: starts at Ø +/- 14mm then pre-cut marking up to 64mm. Pre-cut marking: 25/38/51 and 64mm. Filter: for absorbing gases and unpleasant smell produced by bowels. Clip: with secure closing system. Material: Bag: Polyvinyl chloride (PVC), Low density polyethylene (LDPE), Ethylene vinyl acetate (EVA). Filter: Charcoal granulates. Clip: ll Plastic Shelf life 3 years. Packing and labelling: Shall be as per regulatory compliance applied to the device. ll closing clip by 20 pieces. Each carton to be clearly identified with product name & characteristics, manufacturer addressee, lot No, number of units per carton and multilingual user manual. Instructions for use: Supplied with IFU in multilingual. Instructions for use such as cutting the adhesive disc, preparation of skin, application on the skin and removal after use. Shall include storage and disposal instructions. Shall include any warnings, precuations for use etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 8670-2: Ostomy collection bags -- Part 2: Requirements and test methods ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to collect waste from the stoma. Product description: The colostomy pouch is one piece open system, with deodorising filter and pre-cut opening. It is a medical device that consists of a bag attached to the patient's skin by an adhesive material. The pouch is drainable from the bottom and kept closed by a removable safety clamp. Bag: Made of flexible transparent plastic with a non woven protection layer on skin side. Shape: Anatomical shaped with smooth edges. Adhesive attachment system: 12 x 12cm, with hydrocolloid skin protector, to ensure, safe continuity between bag and colostomy. Cut-to-fit ostomy bags allow the pouch to fit the shape and size of most stomas Hygenically packed for medical use. Non sterile, Single use only. Measurements: Size: Bag: Length 19 x width 14 cm with: Window/aperture: starts at Ø +/- 14mm then pre-cut marking up to 64mm. Pre-cut marking: 25/38/51 and 64mm. Filter: for absorbing gases and unpleasant smell produced by bowels. Clip: with secure closing system. Material: Bag: Polyvinyl chloride (PVC), Low density polyethylene (LDPE), Ethylene vinyl acetate (EVA). Filter: Charcoal granulates, carbon filter, or any other type of filter with similar or superior characteristics. Clip: Plastic Shelf life 3 years. Packing and labelling: Shall be as per regulatory compliance applied to the device. Closing clip by 20 pieces. Each carton to be clearly identified with product name & characteristics, manufacturer addressee, lot No, number of units per carton and multilingual user manual. Instructions for use: Supplied with IFU in multilingual. Instructions for use such as cutting the adhesive disc, preparation of skin, application on the skin and removal after use. Shall include storage and disposal instructions. Shall include any warnings, precuations for use etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per the laetst revision of standards): Must comply with the following standards: ISO 8670-2: Ostomy collection bags -- Part 2: Requirements and test methods ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. Environmental requirements: ISO 14001
184 | Medical supplies | COMPRESSGAUZE_100 | Gauze, compress, 10 x 10cm, non-sterile, P100 | Gauze, compress, 10 x 10cm, non-sterile, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use : For uses such as disinfection of skin, cleaning wounds, Used as wound dressing when sterilized Product description: A non-medicated, dressing item in the form of a patch (also referred to as a sponge) made from woven material and primarily designed to absorb fluids for medical purposes. Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton, medical grade Components: Bleached, purified textile, plain weave. Not detectable by X-ray. Steam sterilization recommended. Shel life Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Non-sterile and single use. Hygenically packed for medical use. Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc; non-sterile gauze compress of 17 threads and 12 folds has been selected as it can also be used for surgery in sterile condition (after steam sterilisation). The size has been chosen as being the most commonly used Shall include storage conditions, disposal instructions etc. Supplied with: Manufacturer's instruction for use. Multilingual Packaging and labelling: Primary packaging: Unit of use. One (1) plastic or paper pack of 100 compresses gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use : For uses such as disinfection of skin, cleaning wounds, Used as wound dressing when sterilized Product description: A non-medicated, dressing item in the form of a patch (also referred to as a sponge) made from woven material and primarily designed to absorb fluids for medical purposes. Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton, medical grade Components: Bleached, purified textile, plain weave. Not detectable by X-ray. Shel life Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Non-sterile and single use. Hygenically packed for medical use. Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc; non-sterile gauze compress of 17 threads and 12 folds has been selected as it can also be used for surgery in sterile condition (after steam sterilisation). The size has been chosen as being the most commonly used Shall include storage conditions, disposal instructions etc. Supplied with: Manufacturer's instruction for use. Multilingual Packaging and labelling: Primary packaging: Unit of use. One (1) plastic or paper pack of 100 compresses gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
185 | Medical supplies |  | Gauze, compress, 10 x 10cm, sterile, P50 | Gauze, compress, 10 x 10cm, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch Secondary packaging: Protected unit. One (1) box of 50 packs of 5 gauze compresses. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch Secondary packaging: Protected unit. One (1) box of 50 packs of 5 gauze compresses. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
186 | Medical supplies | COMPRESSGAUZE_20 | Gauze, compress, 10 x 10cm, sterile, P20 | Gauze, compress, 10 x 10cm, sterile, P20 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch Secondary packaging: Protected unit. One (1) box of 4 packs of 5 gauze compresses. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch Secondary packaging: Protected unit. One (1) box of 4 packs of 5 gauze compresses. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
187 | Medical supplies | COMPRESSGAUZE_5 | Gauze, compress, 10 x 10cm, sterile, P5 | Gauze, compress, 10 x 10cm, sterile, P5 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch No secondary packaging. Only One (1) primary pack of Five (5) gauze compresses in an individual sterilised peel pouch. . Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: non-medicated, sterile gauze made from woven material and primarily designed to absorb fluids for medical purposes and wound care. Product description: Gauze compress, surgical folding, i.e. no free threads apparent after folding or when the first outside fold is opened. Material: Absorbent gauze, 100% cotton. Components: Bleached, purified textile, plain weave. Compress gauze folded. Non-detectable by X-ray. Thread count warp: 95 to 105 threads/dm, weft 66 to 74 threads/dm. Weight: 23g/m2. Type of gauze: 17 threads/cm2 (grammage 23g/m2). Number of folds (thickness): 12. Width: approx. 10cm. Length: approx. 10cm. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or ionising radiation. Shelf life Instructions for use: Dressing material used for cleaning wounds, disinfecting healthy skin, protecting wounds, etc. Shall indicate storage conditions, disposal instructions etc. Sterile gauze compress: 17 threads, 12 folds (same quality and same standard as the non-sterile gauze compress), reserved for emergencies and surgical departments, until a reliable sterilisation system is set up; the size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Shall be as per regulations applied ot the device Primary packaging: Unit of use. Five (5) gauze compresses in an individual sterilised peel pouch No secondary packaging. Only One (1) primary pack of Five (5) gauze compresses in an individual sterilised peel pouch. . Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): BS EN 14079 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze or as per pharmacopeia specifications for bleached cotton gauze EN 13726-1 Test methods for primary wound dressings - Part 1: Aspects of absorbency ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. You should select the tests appropriate for the duration and level of contact.
188 | Medical supplies | COTTONWOOL | Cotton wool, 500g, roll, non-sterile | Cotton wool, 500g, roll, non-sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Absorbent, non-sterile mass of cotton in roll forml and used as a general-purpose absorbent or padding for wound care. Product description: Absorbent cotton used for skin disinfection before injections, wound dressing etc. Surgical quality 100% cotton. Material: 100 % surgical hydrophilic cotton, which has been carefully purified, bleached, and carded. Roll of 500 gram net weight. Not pre-cut. Non-sterile cotton wool: can also be used in sterile condition (after steam sterilisation). Single use. Rolled with or without carft paper to facilitate unwinding. Hygenically packed for medical use. Shelf life required. Capable of undergoing steam sterilization as and when needed by users. Instructions for use: Dressing material with high absorption used for cleaning wounds. Non-sterile cotton wool: can also be used in sterile condition (after steam sterilisation). The size has been chosen as being the most commonly used. Safety process: The cotton wool is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use Packaging and labelling: One (1) roll of cotton wool is individually packed to maintain the integrity, cleanliness of the product during trasport and storage. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): Complies to European, US pharmacopea or equivalent | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Absorbent, non-sterile mass of cotton in roll forml and used as a general-purpose absorbent or padding for wound care. Product description: Absorbent cotton used for skin disinfection before injections, wound dressing etc. Surgical quality 100% cotton. Material: 100 % surgical hydrophilic cotton, which has been carefully purified, bleached, and carded. Roll of 500 gram net weight. Not pre-cut. Non-sterile cotton wool: can also be used in sterile condition (after steam sterilisation). Single use. Rolled with or without carft paper to facilitate unwinding. Hygenically packed for medical use. Shelf life required. Capable of undergoing steam sterilization as and when needed by users. Instructions for use: Dressing material with high absorption used for cleaning wounds. Non-sterile cotton wool: can also be used in sterile condition (after steam sterilisation). The size has been chosen as being the most commonly used. Safety process: The cotton wool is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use Packaging and labelling: One (1) roll of cotton wool is individually packed to maintain the integrity, cleanliness of the product during trasport and storage. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standards): Complies to European, US pharmacopea or equivalent
189 | Medical supplies | CPDABLOODBAG250_1 | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: One (1) blood bag is provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
190 | Medical supplies | CPDABLOODBAG250_10 | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 250ml | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
191 | Medical supplies | CPDABLOODBAG250_5 | Bloodbag+CPDA-1(citrate-phosphate-dextrose-adenine),250ml,P5 | Bloodbag+CPDA-1(citrate-phosphate-dextrose-adenine),250ml,P5 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 250ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 35 ml/250 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
192 | Medical supplies | CPDABLOODBAG350_1 | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),350ml,P1 | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),350ml,P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 350ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 49 ml/350 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 350ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 49 ml/350 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: One (1) blood bag is provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
193 | Medical supplies | CPDABLOODBAG350_10 | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 350ml | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 350ml | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 350ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 49 ml/350 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 350ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 49 ml/350 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
194 | Medical supplies | CPDABLOODBAG450_1 | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),450ml,P1 | Blood bag+CPDA-1 (citrate-phosph.-dextrose-adenine),450ml,P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 450ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 63 ml/450 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 450ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 63 ml/450 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: One (1) blood bag is provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
195 | Medical supplies | CPDABLOODBAG450_10 | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 450ml | Blood bag+CPDA-1 (citrate-phosphate-dextrose-adenine), 450ml | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 450ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 63 ml/450 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to collect whole blood from a donor and to be used later for transfusions after testing. Product description: Single blood bag with CPDA-1 anticoagulant, tubing and a needle to be used for blood collection, testing and storage (transfusions). It includes an integrated diversion pouch. Sterile. Non-toxic and non-pyrogenic. Shelf life of unused bag: minimum 2 years. Shelf life of bag with blood: approx. 35 days. Bag: Capacity 450ml. Material: medical grade PVC (preferably DEHP free). An integrated diversion pouch to collect first 30ml of blood, to prevent bacteria infections. Bag is transparent, almost colorless to allow inspection of the product. With anti-coagulant CPDA-1. Anticoaglulant Citrate Phosphate Dextrose Adenine solution (CPDA-1) complies with the current PhEur or USP monograph. Volume of CPDA-1 per bag: 63 ml/450 ml. Bag has tubing (giving line) connected to a needle for blood collection, and two outlets to connect the blood transfusion set (with protective cap). Blank label for essential data (minimum: bag number, collection date, expiry date, type of product, ABO group + Rh) Needle: Material: Stainless steel. Size: 16G to 20G (approx. 1.0 x 50mm) - for adults. Connected to the tubing at the donor's end. With a protective cap. Needle guard, mounted on the tubing to prevent accidental needlestick injury. The quard slides over the soiled needle after use and self-locks. The design of the needle and needle guard must be capable of being withdrawn from the venepuncture site smoothly, in a single step. Sterile, single use. Tubing: Material: medical grade polymer (preferably DEHP free). Flexible but kink resistant. Length: minimum 80 cm. Fitted with two clamps. Supplied with: One unit includes a single blood bag with anticoagulant, tubing and a needle with protective cap, and an integrated diversion pouch. Manufacturer's instructions for use. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: To be used for blood storage and transfusions, and to collect blood from an adult donor. Single use device supplied in sterile packaging, check the integrity of each unit before use, do not use if the packaging is damaged. Packaging and Labelling: Primary packaging: One (1) blood bag packed in an individual sterilised protective packaging. Secondary packaging: Ten (10) blood bags are provided in aluminium foil package with airtight wrapping, or two packs of Five (5) blood bags are provided in aluminium foil package with airtight wrapping. Foil pack is supplied inside a strong carton box. Protected against light. Storage instructions to be provided. Labelling requirement: Device should be labelled in accordance with British Transfusion Services Guidelines https://www.transfusionguidelines.org/red-book/chapter-26-specification-for-blood-pack-base-labels. Must meet the requirements of ISO 3826-1 or equivalent. Label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223 and ISO 3826-2. CE mark. The word “Sterile” (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words “For single use” (or equivalent harmonised symbol). The words “Do not reuse"" or ""Single use only” (or equivalent harmonised symbol). The words “Pyrogenic free fluid pathway” (or equivalent harmonised symbol). The words ""check the package integrity before use"", or similar warning. The words “contains DEHP” or ""DEHP free"" (or equivalent harmonised symbol). The words ""Destroy after use"", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 This is a combination product requiring a completed 'Questionnaire for Blood Products' for the CPDA-1 component. Certificate of analysis of the latest three batches of CPDA-1 must be submitted for evaluation. Classification: Classified under EU MDR 2017/745 as Class IIb device Applicable standards ( must comply to the latets revisions of the standards): ISO 3826-1: Plastics collapsible containers for human blood and blood components - Part 1: Conventional containers. edition: 3 ISO 3826-2: Plastics collapsible containers for human blood and blood components — Part 2: Graphical symbols for use on labels and instruction leaflets. edition 1 ISO 3826-3: Plastics collapsible containers for human blood and blood components - Part 3: Blood bag systems with integrated feature. edition 1 ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (tests for biocompatibility, extractables, toxicity etc) ISO 10993-3 Biological evaluation of medical devices. Part 3: Tests for genotoxicity, carcinogenicity and reproductive toxicity ISO 10993-4: Biological evaluation of medical devices. Part 4: Selection of tests for interactions with blood ISO 10993-5: Biological evaluation of medical devices. Part 5: Tests for in vitro cytotoxicity ISO 10993-17: Biological evaluation of medical devices. Part 17: Establishment of allowable limits for leachable substances ISO 10993-18: Biological evaluation of medical devices. Part 18: Chemical characterisation of materials. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to steriliz. method. ISO 11607-1: Packaging for terminally sterilized MDs- Part 1 ISO 11607-2: Packaging for terminally sterilized MDs - Part 2
196 | Medical supplies | DRUGS.BAG | Bag (envelope), plastic, for drugs, 10 x 15cm | Bag (envelope), plastic, for drugs, 10 x 15cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Plastic tablet bag presented with three white write-on horizontal stripes, clearly and cleanly printed on one side of the sachet, of approx. 100 x 15mm for writing: Name of drug Quantity Dosage instructions Date Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as ziplock, gripzeal or re-closable bags). Moisture resistant closure. Active length: approx. 150mm (i.e. the length of the bag below the seal) Width: approx. 100mm. Thickness: 0.05mm (50 micorns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One plastic pack of 100 sachets. Regulation & conformity requirements QMS ISO 9001 certificate Regulation & conformity requirements: NA Classification: NA Safety & product Standards: ISO 9001:2008 Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Use other Polymer materials not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Plastic tablet bag presented with three white write-on horizontal stripes, clearly and cleanly printed on one side of the sachet, of approx. 100 x 15mm for writing: Name of drug Quantity Dosage instructions Date Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as ziplock, gripzeal or re-closable bags). Moisture resistant closure. Active length: approx. 150mm (i.e. the length of the bag below the seal) Width: approx. 100mm. Thickness: 0.05mm (50 micorns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One plastic pack of 100 sachets. Regulation & conformity requirements QMS ISO 9001 certificate Regulation & conformity requirements: NA Classification: NA Safety & product Standards: ISO 9001:2008 Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Use other Polymer materials not PVC
197 | Medical supplies | DRUGS.BAG2_1 | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P1 | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Packaging used for patient-prescribed drug distribution. Plastic tablet envelope presented with blue stripes on one side for writing: - Name of drug - Quantity - Dosage instructions - Date -Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as zip lock, grip zeal or re-closable bags). Moisture resistant closure Active length: approx. 280mm (i.e. the length of the bag below the seal) Width: approx. 180mm. Thickness: 0.05mm (50 microns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for user. Packaging and labelling: One plastic pack of 100 bags. Labelling: Name and/or trademark and address of the manufacturer. Regulation & conformity requirements: NA Classification: NA Safety & product Standards: Must comply with the following standards: ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Packaging used for patient-prescribed drug distribution. Plastic tablet envelope presented with blue stripes on one side for writing: - Name of drug - Quantity - Dosage instructions - Date -Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as zip lock, grip zeal or re-closable bags). Moisture resistant closure Active length: approx. 280mm (i.e. the length of the bag below the seal) Width: approx. 180mm. Thickness: 0.05mm (50 microns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for user. Packaging and labelling: One plastic pack of 100 bags. Labelling: Name and/or trademark and address of the manufacturer. Regulation & conformity requirements: NA Classification: NA Safety & product Standards: Must comply with the following standards: ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Not use PVC polymer
198 | Medical supplies | DRUGS.BAG2_100 | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P100 | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Packaging used for patient-prescribed drug distribution. Plastic tablet envelope presented with blue stripes on one side for writing: - Name of drug - Quantity - Dosage instructions - Date -Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as zip lock, grip zeal or re-closable bags). Moisture resistant closure Active length: approx. 280mm (i.e. the length of the bag below the seal) Width: approx. 180mm. Thickness: 0.05mm (50 microns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for user. Packaging and labelling: One plastic pack of 100 bags. Labelling: Name and/or trademark and address of the manufacturer. Regulation & conformity requirements: NA Classification: NA Safety & product Standards: Must comply with the following standards: ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to store drug and to distribute to patients. Product description: Packaging used for patient-prescribed drug distribution. Plastic tablet envelope presented with blue stripes on one side for writing: - Name of drug - Quantity - Dosage instructions - Date -Patient’s Name Material: Virgin polyethylene. Self-sealing with minigrip type zip lock (also known as zip lock, grip zeal or re-closable bags). Moisture resistant closure Active length: approx. 280mm (i.e. the length of the bag below the seal) Width: approx. 180mm. Thickness: 0.05mm (50 microns) Instructions for use: The bag sealable is used to the distribution of drugs to the patients. Supplied with: Manufacturer's instruction for user. Packaging and labelling: One plastic pack of 100 bags. Labelling: Name and/or trademark and address of the manufacturer. Regulation & conformity requirements: NA Classification: NA Safety & product Standards: Must comply with the following standards: ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes Technical Standard: Product must comply to BS EN 71-3:2005 Safety of Toys - Part 3: Specifications for migration of certain elements. Environmental requirements: Not use PVC polymer
199 | Medical supplies | EXTRACTORMUCUS_1 | Extractor, mucus, 20ml, sterile, single use, P1 | Extractor, mucus, 20ml, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001
200 | Medical supplies | EXTRACTORMUCUS_10 | Extractor, mucus, 20ml, sterile, single use, P10 | Extractor, mucus, 20ml, sterile, single use, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001
201 | Medical supplies | EXTRACTORMUCUS_100 | Extractor, mucus, 20ml, sterile, single use, P100 | Extractor, mucus, 20ml, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001
202 | Medical supplies | EXTRACTORMUCUS_50 | Extractor, mucus, 20ml, sterile, single use, P50 | Extractor, mucus, 20ml, sterile, single use, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways pharynx in newborn babies to ensure free respiration. Components: Transparent container to permit immediate visual examination of the mucus. Catheter, size selected: CH12, with open end smooth round tip for trauma free insertion, distal end with conical tip. Catheter tip ensures atrumatic intubation. 2 catheters, one from the reservoir to operator’s mouth and the second one for suction. With filter to prevent entry of mucus to users mouth during suction. Single chamber container, plastic Self-adhesive identification label for newborn identification Material: Container: polypropylene, medical grade Catheters: polyvinyl chloride, medical grade Capacity: approx. 20 ml. Sterile and single use. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile mucus extractor device used for aspirating secretions or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Safety process: The mucus extractor is for single use only. When inserting a suction tube, use clean techniques. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, contraindications if any. Storage conditions and disposal instructions to be included. Packaging and labelling: Shall be as per regulations applied to the device. Primary packaging: Unit of use One (1) mucus extractor in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 mucus extractors. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards (must be as per latest revisions of the standards): Must comply with the following standards: ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 10993-1 Biological Evaluation of medical devices Environmental requirements: ISO 14001
203 | Medical supplies | GLOVESEXAML_100 | Gloves, exam, latex, large, non-sterile, single use, P100 | Gloves, exam, latex, large, non-sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile, Single-use disposable Powderfree Size: large Total length: minimum 240mm. Width: 110 mm, +/- 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, large, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile, Single-use disposable Powderfree Size: large Total length: minimum 240mm. Width: 110 mm, +/- 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, large, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process
204 | Medical supplies | GLOVESEXAMM_1 | Gloves, exam, latex, medium, non-sterile | Gloves, exam, latex, medium, non-sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: medium. Total length: minimum 240mm. Width: 95 mm, ± 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, medium, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: medium. Total length: minimum 240mm. Width: 95 mm, ± 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, medium, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces
205 | Medical supplies | GLOVESEXAMM_100 | Gloves, exam, latex, medium, non-sterile, P100 | Gloves, exam, latex, medium, non-sterile, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: medium. Total length: minimum 240mm. Width: 95 mm, ± 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, medium, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: medium. Total length: minimum 240mm. Width: 95 mm, ± 10mm. Thickness: fingers minimum 0.12mm; palm minimum 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: 1(one) Gloves, exam, latex, medium, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces
206 | Medical supplies | GLOVESEXAMS_100 | Gloves, exam, latex, small, non-sterile, single use, P100 | Gloves, exam, latex, small, non-sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: small. Total length: minimum 240mm. Width: 80 mm, +/- 10mm. Thickness: fingers: 0.12mm; palm: 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: Unit presentation: 1(one) Gloves, exam, latex, small, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gloves intended for use in the medical field to protect patient and user from cross-contamination Product description: Glove for clinical examinations and routine clinical laboratory work contains of 5 fingers, palm and a sleeve Ambidextrous (Fits either hand). Material: Natural rubber latex. Waterproof. Non-sterile Single-use disposable powderfree Size: small. Total length: minimum 240mm. Width: 80 mm, +/- 10mm. Thickness: fingers: 0.12mm; palm: 0.1mm. AQL for holes = 1.5 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Instructions for use: Used for multiple purposes Used in wide range of laboratory services from basic laboratory to reference centre. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: shall be as per regulations applicable to the device. Unit presentation: Unit presentation: 1(one) Gloves, exam, latex, small, disp, box/100 Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 21171:2006 Medical gloves - Determination of removable surface powder ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management proces
207 | Medical supplies | GLOVESGYNAECOLL1 | Gloves, gynaecological, large, powder-free, sterile, 1 pair | Gloves, gynaecological, large, powder-free, sterile, 1 pair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: large (8.5-9). Total length: 400mm. Width: 108±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: large (8.5-9). Total length: 400mm. Width: 108±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
208 | Medical supplies | GLOVESGYNAECOLL25; GLOVESGYNAECOLL50 | Gloves, gynaecological, large, powder-free, sterile | Gloves, gynaecological, large, powder-free, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: large (8.5-9). Total length: 400mm. Width: 108±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: large (8.5-9). Total length: 400mm. Width: 108±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
209 | Medical supplies | GLOVESGYNAECOLM1 | Gloves, gynaecological, medium, powder-free, sterile, 1 pair | Gloves, gynaecological, medium, powder-free, sterile, 1 pair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: medium (7.5-8). Total length: approximately 400mm. Palm width: approximately 95±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: medium (7.5-8). Total length: approximately 400mm. Palm width: approximately 95±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
210 | Medical supplies | GLOVESGYNAECOLM25; GLOVESGYNAECOLM50 | Gloves, gynaecological, medium, powder-free, sterile | Gloves, gynaecological, medium, powder-free, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: medium (7.5-8). Total length: approximately 400mm. Palm width: approximately 95±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: medium (7.5-8). Total length: approximately 400mm. Palm width: approximately 95±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
211 | Medical supplies | GLOVESGYNAECOLS1 | Gloves, gynaecological, small, powder-free, sterile, 1 pair | Gloves, gynaecological, small, powder-free, sterile, 1 pair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
212 | Medical supplies | GLOVESGYNAECOLS25 | Gloves, gynaecological, small,powder-free, sterile, P25 | Gloves, gynaecological, small,powder-free, sterile, P25 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
213 | Medical supplies | GLOVESGYNAECOLS50 | Gloves, gynaecological, small,powder-free, sterile, P50 | Gloves, gynaecological, small,powder-free, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be worn on the hands in surgical settings for Gynaecological use, provide a barrier against potentially infectious material and other contaminants Product description: Long cuff gloves Pair of gynaecological gloves with long cuff, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural latex. Powder-free. Size selected: Gynaecological gloves, size: small (6.5-7). Total length: approximately 400mm. Palm width: approximately 83±5mm. Thickness: cuff minimum 0.12mm; palm minimum 0.16mm; fingers minimum 0.18mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves with long cuff used for gynaecological and obstetric procedures to protect both patient and medical staff from cross-contamination. Gynaecological gloves, also named long cuff gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Use intructions, warnings, storage conditions, disposal instructions etc. Packaging and labelling: Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 25 or 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
214 | Medical supplies | GLOVESSURG7_1 | Gloves, surgical, size 7, powder-free, sterile, single use, | Gloves, surgical, size 7, powder-free, sterile, single use, | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7 Total length: ≥270mm Width: approximately 89 ± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7 Total length: ≥270mm Width: approximately 89 ± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
215 | Medical supplies | GLOVESSURG7_50 | Gloves, surgical, size 7, powder-free, sterile, single use | Gloves, surgical, size 7, powder-free, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7 Total length: ≥270mm Width: approximately 89 ± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7 Total length: ≥270mm Width: approximately 89 ± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
216 | Medical supplies | GLOVESSURG8_50 | Gloves, surgical, size 8, powder-free, sterile, single use, | Gloves, surgical, size 8, powder-free, sterile, single use, | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8 Total length: ≥270mm Width: approximately 102± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8 Total length: ≥270mm Width: approximately 102± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
217 | Medical supplies | GLOVESSURGICAL8_1 | Gloves, surgical, size 8, powder-free, sterile, single use, | Gloves, surgical, size 8, powder-free, sterile, single use, | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7.5 Total length: ≥270mm Width: approximately 95± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8 Total length: ≥270mm Width: approximately 95± 5mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
218 | Medical supplies | GLOVESSURGICAL7.5 | Gloves, surgical, size 7,5, powder-free, sterile, single use | Gloves, surgical, size 7,5, powder-free, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8 Total length: ≥270mm Width: approximately 102± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 7,5 Total length: ≥270mm Width: approximately 102± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of 50 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
219 | Medical supplies | GLOVESSURGICAL8.5 | Gloves, surgical, size 8.5, powder-free, sterile, single use | Gloves, surgical, size 8.5, powder-free, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8.5 Total length: ≥280mm Width: approximately 108± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to be worn on the hands in surgical settings to provide barrier against potentially infectious agents Product description: Pair of surgical gloves, anatomically shaped: 1 right-handed, 1 left-handed. Waterproof. Stretch proof. Material: Natural rubber latex. Powder-free Size selected: Surgical gloves, size: 8.5 Total length: ≥280mm Width: approximately 108± 6mm Minimum thickness: Smooth area: 0,10mm; Textured area: 0,13mm Sterile and single use. AQL for holes = 0.65 (water tightness test) Low residual extractable protein level Shelf life to be indicated. Initial sterilisation method: Ethylene oxide gas or Gamma radiation. Instructions for use: Sterile single use gloves used for surgical procedures to protect both patient and medical staff from cross-contamination. Surgical gloves are sterile and strictly SINGLE USE This item must be ordered in sufficient quantity to encourage its use The sizes have been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use.Use instructions, storage instructions, disposal instructions etc. Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) pair of gloves in an individual sterilised peel pack Supplied with double packaging: one interior layer plus sterilised peel-pack. Packed in the interior layer: right and left hand with the sleeves reversed up to the thumb. Labelled with: size, right hand, left hand and position of the sleeves. Secondary packaging: Protected unit. One (1) box of any pack size from 25 to 100 pairs gloves. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (all submissions must be as per latest revisions of standards): Must comply with the following applicable standards or equivalent ensuring conformance to the latest revisions BS EN 455-1:2020 Medical gloves for single use. Requirements and testing for freedom from holes BS EN 455-2:2015 Medical gloves for single use. Requirements and testing for physical properties BS EN 455-3:2015 Medical gloves for single use. Requirements and testing for biological evaluation BS EN 455-4:2009 Medical gloves for single use. Requirements and testing for shelf life determination ISO 12243:2003 Medical gloves made from natural rubber latex -- Determination of water-extractable protein using the modified Lowry method or ASTM D 5712 ISO 2859-1 – Sampling procedures ISO 10993-1: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process ISO 11607-1&2 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems
220 | Medical supplies | HANDBRUSH_1 | Brush, hand, scrubbing, plastic, P1 | Brush, hand, scrubbing, plastic, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used by healthcare workers to scrub their hands, fingers, and forearms prior to surgery or other intervention where a high degree of personal hygiene is required Product description: To be used for scrubbing hand prior to surgical intervention. Brush, nylon bristles, plastic block. Soft bristles: Minimum 5 rows. Material: - Head: Polypropylene. - Bristles: Nylon. Dimension: Length: head approx. 8 - 10 cm. Width: approx. 3 - 5 cm. Height: approx. 1 cm. Reusable, can be sterilized by autoclave Non-sterile Instructions for use: Plastic brush for scrubbing hands in general, and specifically prior to surgery. This brush must be a plastic type that can be autoclaved after cleaning and disinfection. Do not use a brush with wooden head as cracks in the wood may harbour contamination. The size has been chosen as being the most commonly used. Safety process: This item is used as a «clean» or «sterile» item. The item must be cleaned, disinfected and sterilized in a steam sterilizer as often as necessary. Protection of users: Washing hands before and after each medical act is extremely important to limit the risk of cross-contamination. Wash hands as often as necessary. Prior to any medical act: Use water and soap. Time approx. 3 minutes. Prior to surgical purpose: Use distilled water and soap or disinfectant. Time approx. 8 minutes. Supplied with: Manufacturer's instruction for use. to provide cleaning, disinfection before autoclave. Packaging and labelling: One ( 1) or two (2) scrubbing brush in a plastic bag with manufacturer's instruction for user in English. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit Ten (10) scrubbing brushes in a box with manufacturer's instruction for use (when applicable). Labelling on the primary packaging: Name and/or trademark of the manufacturer. Symbols used according ISO 15223 Secondary packaging: Ten (10) scrubbing brushes in a box with manufacturer's instruction for use (when applicable). Labelling on the secondary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Number of units per secondary packaging. Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: NA Safety & product Standards: Must comply with the following standards: ISO 9001 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used by healthcare workers to scrub their hands, fingers, and forearms prior to surgery or other intervention where a high degree of personal hygiene is required Product description: To be used for scrubbing hand prior to surgical intervention. Brush, nylon bristles, plastic block. Soft bristles: Minimum 5 rows. Material: - Head: Polypropylene. - Bristles: Nylon. Dimension: Length: head approx. 8 - 10 cm. Width: approx. 3 - 5 cm. Height: approx. 1 cm. Reusable, can be sterilized by autoclave Non-sterile Instructions for use: Plastic brush for scrubbing hands in general, and specifically prior to surgery. This brush must be a plastic type that can be autoclaved after cleaning and disinfection. Do not use a brush with wooden head as cracks in the wood may harbour contamination. The size has been chosen as being the most commonly used. Safety process: This item is used as a «clean» or «sterile» item. The item must be cleaned, disinfected and sterilized in a steam sterilizer as often as necessary. Protection of users: Washing hands before and after each medical act is extremely important to limit the risk of cross-contamination. Wash hands as often as necessary. Prior to any medical act: Use water and soap. Time approx. 3 minutes. Prior to surgical purpose: Use distilled water and soap or disinfectant. Time approx. 8 minutes. Supplied with: Manufacturer's instruction for use. to provide cleaning, disinfection before autoclave. Packaging and labelling: One ( 1) or two (2) scrubbing brush in a plastic bag with manufacturer's instruction for user in English. Alternatively, the instruction for use can be indicated on a separate insert. Secondary packaging: Protected unit Ten (10) scrubbing brushes in a box with manufacturer's instruction for use (when applicable). Labelling on the primary packaging: Name and/or trademark of the manufacturer. Symbols used according ISO 15223 Secondary packaging: Ten (10) scrubbing brushes in a box with manufacturer's instruction for use (when applicable). Labelling on the secondary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Number of units per secondary packaging. Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: NA Safety & product Standards: Must comply with the following standards: ISO 9001 Environmental requirements: ISO 14001
221 | Medical supplies | INFUSION.SET_1 | Infusion set, sterile, single use, 1 piece | Infusion set, sterile, single use, 1 piece | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001
222 | Medical supplies | INFUSION.SET_100 | Infusion set, sterile, single use, P100 | Infusion set, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001
223 | Medical supplies | INFUSION.SET_400 | Infusion set, sterile, single use, P400 | Infusion set, sterile, single use, P400 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: noninvasive sterile device to conduct fluids (except blood and blood products) from an intravenous fluid container to a patient's venous system. Product description: A collection of sterile devices designed to conduct fluids from an intravenous (IV) fluid container to a patient's venous system; used for gravitational intravenous administration. Components: Tubing Spike or perforator Air inlet with air filter and protective cap. Drip-counting chamber. Flow regulator. Y-injection site. Luer-lock connection to vascular system Materials: The material of construction shall be medical grade and must not modify the properties of the fluids or solutions passing through the infusion set. Tube: plastic (PVC: polyvinyl chloride), medical grade. DEHP free. Transparent (allowing the detection of air bubbles). Resistant to kinking. Length is approximately 150cm (overall IV giving set length approximately 170cm). Internal / external diameter is 3mm / 4.1mm. Plastic Perforator: Medical grade plastic ABS (Acrylonitrile Butadiene Styrene), POM (polyacetal), polystirole. Double Channel, Hollow device located at the proximal end of the infusion set allowing it to be connected to the infusion bottle or bag. Composed of a tapering tube mounted on a base ensuring a good seal between the set and the infusion bottle or bag. Fitted with a protecting cap and air inlet. Air inlet: plastic (ABS: acrylonitrile butadiene styrene) Incorporated into the perforator. Fitted with an air filter (bacteriological filter). Drop-counting chamber: Materials: medical-grade PVC and polypropylene. Transparent, supple, compressible bag with calibrated tube (upper end) providing even drop formation: 20 drops of ±1ml. - Bacterial filter 15-20µm (chamber, lower end) designed to retain particles. Flow regulator: plastic (ABS: acrylonitrile butadiene styrene). Ensures the flow regularity during the IV administration. (Clamp to control administration speed/volume) Injection portal: plastic (ABS: acrylonitrile butadiene styrene)+ synthetic rubber. Allows introducing precise volume of injectable medicines into the infusion flow. Natural rubber latex-free Y-shaped, self-sealing membrane made of synthetic elastomer Terminal connection (stopper): plastic (ABS: acrylonitrile butadiene styrene). With Luer lock connector. Includes expiration date (shelf life) Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Instructions for use: Sterile item, used for parenteral administration of injectable preparations. Sterile infusion giving set with graduated chamber enabling a slow intravenous administration of a precise volume of infusion or injectable drug over a given time. It is to be connected between the infusion and the intravenous catheter. An infusion giving set should not be left in situ for more than 72 hours. Not for use with blood or blood products or light sensitive products. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions, precautions for use, warnings etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) Infusion giving set in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 10 Infusion giving set units. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (all submissions must be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 8536-4: Infusion equipment for medical use - Part 4: Infusion sets for single use, gravity feed ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Special Instructions to CO: Provision for needle to be made separately Environmental requirements: ISO 14001
224 | Medical supplies | NEEDLE.SPINAL25G | Needle, spinal, 25G, sterile, single use | Needle, spinal, 25G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General information For injection of a local anaesthetic and for spinal anaesthesia. For taking spinal fluid samples. Specifications Components & materials: Cannula: stainless steel, without silicone. Stylet: stainless steel, without silicone. Base: polypropylene or polyamide. Cap: rigid plastic. Product description: Needle, spinal, 25G (0.53 x 90mm) disposable, sterile. Quincke bevel spinal needle. Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). Needle core hub with lock to keep the surface of needle bevel joint well. External diameter expressed in Gauge and mm. Length expressed in mmColour code/external diameter: Visible at the base of the spinal needle. Components: Needle, stylet and sheath. Material: Stainless steel needle and stylet without silicone. Clear ridged polycarbonate hub. Protective plastic sheath. Size selected: 25G (0.53 x 90mm) . Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 25 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Includes multilingual user instructions: English, French and Spanish, others when available. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards: EN-ISO 20594-1: Conical fittings with a 6% (Luer) taper for syringes, needles and certain other medical equipment - Part 1: General requirements ISO 80369-7 : Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications EN-ISO 7864-1: Sterile hypodermic needles for single use ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: Use of other polymer instead of PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. General information For injection of a local anaesthetic and for spinal anaesthesia. For taking spinal fluid samples. Specifications Components & materials: Cannula: stainless steel, without silicone. Stylet: stainless steel, without silicone. Base: polypropylene or polyamide. Cap: rigid plastic. Product description: Needle, spinal, 25G (0.53 x 90mm) disposable, sterile. Quincke bevel spinal needle. Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). Needle core hub with lock to keep the surface of needle bevel joint well. External diameter expressed in Gauge and mm. Length expressed in mmColour code/external diameter: Visible at the base of the spinal needle. Components: Needle, stylet and sheath. Material: Stainless steel needle and stylet without silicone. Clear ridged polycarbonate hub. Protective plastic sheath. Size selected: 25G (0.53 x 90mm) . Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas. Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 25 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Includes multilingual user instructions: English, French and Spanish, others when available. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards: EN-ISO 20594-1: Conical fittings with a 6% (Luer) taper for syringes, needles and certain other medical equipment - Part 1: General requirements ISO 80369-7 : Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications EN-ISO 7864-1: Sterile hypodermic needles for single use ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: Use of other polymer instead of PVC
225 | Medical supplies | NEEDLELUER19G | Needle, luer, 19G, sterile, P100 | Needle, luer, 19G, sterile, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: plastic, medical grade. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 19G (1.1 x 40mm), Color coded (Cream). Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle between 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be indicated Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: plastic, medical grade. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 19G (1.1 x 40mm), Color coded (Cream). Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle between 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be indicated Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001
226 | Medical supplies | NEEDLELUER21G | Needle, luer, 21G, sterile, P100 | Needle, luer, 21G, sterile, P100 | Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 21G (0.8 x 40mm), color coded green Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle betweeGn 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be included Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001 | Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 21G (0.8 x 40mm), color coded green Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle betweeGn 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be included Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001
227 | Medical supplies | NEEDLELUER23G | Needle, luer, 23G, sterile, P100 | Needle, luer, 23G, sterile, P100 | Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 23G (0.6 x 30mm), color coded blue Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle between 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be included Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001 | Intended purpose of use: to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient. Product description: Hypodermic needle Needle with base (Luer type fitting), and protective cap. Material: Needle: Stainless steel. Base and protective cap: medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in Gauge and millimeters. Length expressed preferably in millimetres. Size: 23G (0.6 x 30mm), color coded blue Cannula and tube: Made of stainless steel conform to ISO 9626-1. Normal or thin-walled. Long bevelled (angle inferior or equal to 13°), or short bevelled (angle between 13° and 19°). Silicone-coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropylene (PP). Internationally recognized colour code system (ISO 6009) allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made of polypropylene (PP). Shelf life to be included Instructions for use: For intravenous (IV) injection: adults, children For intramuscular (IM) injection: adults. For intra-articular: adults, children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer's instruction for use. Indications for use, Storage instructions, disposal instructions, precautions for use such as "To prevent accidental needlestick injuries, never recap the needle after use". Packaging and labelling: Primary packaging: Unit of use. 1 needle packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: Protected unit. 1 box of 100 blister packed needles. Over packaging: Packaging unit. X (protected units) packed in (extra packaging) rigid packaging with or without plastic film. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (all submissions shall be as per latest revisions of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 7864: Sterile hypodermic needles for single use -Requirements and test methods ISO 9626: Stainless steel needle tubing for the manufacture of medical devices - Requirements and test methods. ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications - Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1- Biological evaluation of medical devices— the general principles governing the biological evaluation of medical devices within a risk management process Environmental requirements: ISO 14001
228 | Medical supplies | NEEDLESCLPVN21100 | Needle, scalp vein, 21G, sterile, P100 | Needle, scalp vein, 21G, sterile, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
229 | Medical supplies | NEEDLESCLPVN2150 | Needle, scalp vein, 21G, sterile, single use, P50 | Needle, scalp vein, 21G, sterile, single use, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
230 | Medical supplies | NEEDLESCLPVN21G3 | Needle, scalp vein, 21G, sterile, single use | Needle, scalp vein, 21G, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to puncture a vein to collect blood and/or for transient (<24 hours) administration of intravenous fluids. Product description: Scalp vein set, also known as scalp vein needle or butterfly needle. Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use, most commonly used for adults. Components: needle with protecting cap, Butterfly wings, tube including Luer connector with cap ( all parts fit together to form a unit) Needle: - Material: stainless steel - silicone -coated for atraumatic insertion. - mounted on a green base with two flexible fins that are used for anchoring and attaching the needle. - butterfly wings for better fixations and handling (Material: PVC), medical grade Tube: -Material: Medical grade Polyurethane (PUR) or Plastized PVC DEHP free - soft, flexible - Length: approx. 150-300 mm - including Luer connection with cap Connector: - Female Luer - usually interlocking - with a bung that permits the attachment of a infusion set. external diameter expressed in Gauge and mm, inscribed i the fins length expressed in mm size selected : 21 G ( 0,8x 20 mm) green disposable, single use. sterile (ETO) Shelf life to be included Instructions for use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 20mm) green, for adults. - 25G (0.50 x 20mm) orange, for children. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if the packaging is undamaged. Safety process: The scalp vein needle is for single use only. Rules of asepsis must be followed when inserting scalp vein needle. Protection of users: To prevent stick injuries, wear gloves. NEVER recap the needle after use and dispose of the needle into a puncture proof safety container. Collect and destroy puncture proof containers either by incinerating them in controlled surroundings or dispose of them in a safe burial pit in compliance with national laws and regulation on health care waste management. Supplied with: Manufacture's information for use. Indications for use, storage instructions, disposal instructions, precautions of use etc Packaging and labelling: shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) scalp vein needle in an individual sterilised peel pack. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT"(or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply as per the latest revision of standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10555-1: Sterile, single-use intravascular catheters — Part 1: General requirements ISO 6009: Hypodermic needles for single use -- Colour coding for identification ISO 11607-1Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
231 | Medical supplies | NEEDLESPINAL22100 | Needle, spinal, 22G, sterile, single use, P100 | Needle, spinal, 22G, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001
232 | Medical supplies | NEEDLESPINAL22G1 | Needle, spinal, 22G, sterile, single use, P1 | Needle, spinal, 22G, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001
233 | Medical supplies | NEEDLESPINAL22G25 | Needle, spinal, 22G, sterile, single use, P25 | Needle, spinal, 22G, sterile, single use, P25 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001
234 | Medical supplies | NEEDLESPINAL22G50 | Needle, spinal, 22G, sterile, single use, P50 | Needle, spinal, 22G, sterile, single use, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection of a local anaesthetic and for spinal anaesthesia and For taking spinal fluid samples. Product description: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Quincke bevel spinal needle. Components & materials: Cannula: stainless steel, without silicone coating. Stylet: stainless steel, without silicone coating. Base: polypropylene or polyamide. Cap: rigid plastic -The cannula is cylindrical; at one end it has a base that is either opaque or translucent, and at the other end a short bevelled tip, called a "Quinck". This short spine bevelled tip, used around the spine, is designed to separate the dural fibres rather than cutting them, thus reducing damage and the loss of cerebrospinal fluid. - The stylet is the same length as the tube; it fits properly into the needle lumen and indicates the bevel position of the needle. - The base allows the needle to be connected to a syringe and is formed by a conical female Luer connection. - hub (base) with standardized colour code corresponding to the external diameter of the needle -Transparent hub, to allow easy visualization of cerebrospinal fluid (CSF). External diameter expressed in Gauge and mm. Length expressed in mm. Size selected: 22G (0.70 x 90mm) black, for spinal anaesthesia, adults. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile spinal needle that can be introduced into the sub-arachnoid space to withdraw cerebrospinal fluid (CSF) for diagnosis purpose, or for injecting local anaesthetic into the CSF for spinal anaesthesia (surgical purpose). Must be used by qualified personnel only. Safety process: The spinal needle is for single use only. Strict rules of asepsis must be followed when inserting spinal needle. To prevent accidental needlestick injuries, never recap the needle after use. Supplied with: Manufacturer's instruction for use. Indications for use, warnings, storage conditions, disposal instructions etc Packaging and labelling: Shall be as per regulations applied to the device Primary packaging: Unit of use. One (1) spinal needle in an individual sterilised peel pack. Secondary packaging: Protected unit. One (1) box of 50 spinal needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to latest revisions of the standards): ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7864:1993: Sterile hypodermic needles for single use ISO 80369-7:Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmentally requirements: ISO 14001
235 | Medical supplies | NEEDLESURGICAL14 | Surgical needle, 1/2 circle, round, spring eye, size14, sterile | Surgical needle, 1/2 circle, round, spring eye, size14, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: 1/2 circle suture needle, round cross section with spring eye. Suture needle size 14 or equivalent. Sterile, Single use Material: stainless steel, martensitic. Shelf life to be included Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied with: Manufacturer's instruction for use. Indications for use, storage conditions, disposal instructions, precautions for use etc Packaging and labelling: Shall be as per regulations applied to the device. Unit packing: Individually packed in sterile peel packs One (1) box of 12 suture needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: 1/2 circle suture needle, round cross section with spring eye. Suture needle size 14 or equivalent. Sterile, Single use Material: stainless steel, martensitic. Shelf life to be included Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied with: Manufacturer's instruction for use. Indications for use, storage conditions, disposal instructions, precautions for use etc Packaging and labelling: Shall be as per regulations applied to the device. Unit packing: Individually packed in sterile peel packs One (1) box of 12 suture needles. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001
236 | Medical supplies | NEEDLESURGICAL27 | Needle, surgical, 1/2, 27mm, round, spring eye, P24 | Needle, surgical, 1/2, 27mm, round, spring eye, P24 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: Surgical Needle, stainless steel, sterile, 1/2 circle. Material: hardened high-quality stainless steel (AISI 420/420+) Size: 1.00 x 27mm, spring eye, needle and tip round bodied, for skin and muscle suturing. Sterile and single use. Shelf life to be included Initial sterilisation method: Ethylene oxide gas. Packaging and labelling: shall be as per regulations applied to the device One (1) suture needle in individual sterile packaging. Box of twelve (12) needles. CE mark with Notified Body number. Symbols used according ISO 15223. Markings Sterile and Single use, or equivalent symbols with Notified Body number. Method of sterilization. Lot, manufacturing and expiry date. Manufacturer name and address. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: Surgical Needle, stainless steel, sterile, 1/2 circle. Material: hardened high-quality stainless steel (AISI 420/420+) Size: 1.00 x 27mm, spring eye, needle and tip round bodied, for skin and muscle suturing. Sterile and single use. Shelf life to be included Initial sterilisation method: Ethylene oxide gas. Packaging and labelling: shall be as per regulations applied to the device One (1) suture needle in individual sterile packaging. Box of twelve (12) needles. CE mark with Notified Body number. Symbols used according ISO 15223. Markings Sterile and Single use, or equivalent symbols with Notified Body number. Method of sterilization. Lot, manufacturing and expiry date. Manufacturer name and address. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001
237 | Medical supplies | NEEDLESURGICAL63 | Needle, surgical, 1/2, 63mm, round, spring eye, P24 | Needle, surgical, 1/2, 63mm, round, spring eye, P24 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: Surgical Needle, stainless steel, sterile, 1/2 circle. Material: hardened high-quality stainless steel (AISI 420/420+) Size: 1.30x63mm, spring eye, needle and tip round bodied, for skin and muscle suturing. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Packaging and labelling: shall be as per regulations applied to the device One (1) suture needle in individual sterile packaging. Box of twelve (12) needles. CE mark with Notified Body number. Symbols used according ISO 15223. Markings Sterile and Single use, or equivalent symbols with Notified Body number. Method of sterilization. Lot, manufacturing and expiry date. Manufacturer name and address. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for the obstetric fistulas repairs Product description: Surgical Needle, stainless steel, sterile, 1/2 circle. Material: hardened high-quality stainless steel (AISI 420/420+) Size: 1.30x63mm, spring eye, needle and tip round bodied, for skin and muscle suturing. Sterile and single use. Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Packaging and labelling: shall be as per regulations applied to the device One (1) suture needle in individual sterile packaging. Box of twelve (12) needles. CE mark with Notified Body number. Symbols used according ISO 15223. Markings Sterile and Single use, or equivalent symbols with Notified Body number. Method of sterilization. Lot, manufacturing and expiry date. Manufacturer name and address. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1:2009 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 13402 Surgical and dental hand instruments — Determination of resistance against autoclaving, corrosion and thermal exposure ASTM F3014 Standard Test Method for Penetration Testing of Needles Used in Surgical Sutures ASTM F1874 Standard Test Method for Bend Testing of Needles Used in Surgical Sutures Environmental requirements: ISO 14001
239 | Medical supplies | PLASTERWOUND_25 | Adhesive bandage, wound plaster, waterproof | Adhesive bandage, wound plaster, waterproof | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile, peel pack, rectangular shaped, Each plaster individually wrapped, Pack of 25 or 50 Elastic fabric, strong adhesion, hypoallergenic adhesive Waterproof With or without ventilation holes Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (for the adhesive that is coming in contact with skin) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile, peel pack, rectangular shaped, Each plaster individually wrapped, Pack of 25 or 50 Elastic fabric, strong adhesion, hypoallergenic adhesive Waterproof With or without ventilation holes Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must be compliant ass per the latest revisions of strandards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 10993-1 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process (for the adhesive that is coming in contact with skin)
240 | Medical supplies | PRESSURECUFF | Pressure cuff, infusion/transfusion, for bags 500-1000 ml | Pressure cuff, infusion/transfusion, for bags 500-1000 ml | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to carry out pressurized infusion/transfusion in emergency situations. Product description: Pneumatic cuff designed for 500/1000ml infusion pouch/blood bags, to carry out pressurized infusion/transfusion in emergency situations. Components: - Cuff sleeve with net window and a pneumatic bag. - Loop: rear suspension and font suspension - Manometer pressure gauge - Air-release valve - inflation bulb Material: - cuff: tear proof cotton - net: polypropylene. - pneumatic bag, connecting tubes: rubber latex free - bulb: polyurethane o rubber latex free - pressure gauge and air-release valve: Stainless steel, chrome-plated - pressure gauge membrane: Copper beryllium. Specifications: -Pressure cuff can be 500ml or 500 to 1000ml - Cuff: washable up to 60°C. Size approx. 35.5x17cm (inflated size 26x13cm) - Reusable Pneumatic bag with 2 side connections. - tube to pressure gauge: approx. 7cm - tube to bulb: aprox.80cm Transparent net window: 25x15cm Suspension system: 2 loops, one to maintain the container and the other to suspend the device from an infusion stand. Dependent on manufacturer and cuff size the suspension system may vary slightly Manometer: - Diameter 49mm, - Aluminium scale from 0 to 300mmHg Instructions for use: used to carry out pressurized infusions or transfusions in emergency situations, for a 500-1000ml infusion pouch or blood bags make sure that pressure gauge is clear of any obstructions and allowed to move freely Warming: - rough handling or abuse may affect the gauge accuracy Supplied with: user and maintenance manual, warranty, in English. Accessories/ spare parts/consumables: - Manometer 300mmHg. - Inflator bulb with release valve Packaging and labelling: Individually wrapped in a plastic bag with a manufacturer manual Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to carry out pressurized infusion/transfusion in emergency situations. Product description: Pneumatic cuff designed for 500/1000ml infusion pouch/blood bags, to carry out pressurized infusion/transfusion in emergency situations. Components: - Cuff sleeve with net window and a pneumatic bag. - Loop: rear suspension and font suspension - Manometer pressure gauge - Air-release valve - inflation bulb Material: - cuff: tear proof cotton - net: polypropylene. - pneumatic bag, connecting tubes: rubber latex free - bulb: polyurethane o rubber latex free - pressure gauge and air-release valve: Stainless steel, chrome-plated - pressure gauge membrane: Copper beryllium. Specifications: -Pressure cuff can be 500ml or 500 to 1000ml - Cuff: washable up to 60°C. Size approx. 35.5x17cm (inflated size 26x13cm) - Reusable Pneumatic bag with 2 side connections. - tube to pressure gauge: approx. 7cm - tube to bulb: aprox.80cm Transparent net window: 25x15cm Suspension system: 2 loops, one to maintain the container and the other to suspend the device from an infusion stand. Dependent on manufacturer and cuff size the suspension system may vary slightly Manometer: - Diameter 49mm, - Aluminium scale from 0 to 300mmHg Instructions for use: used to carry out pressurized infusions or transfusions in emergency situations, for a 500-1000ml infusion pouch or blood bags make sure that pressure gauge is clear of any obstructions and allowed to move freely Warming: - rough handling or abuse may affect the gauge accuracy Supplied with: user and maintenance manual, warranty, in English. Accessories/ spare parts/consumables: - Manometer 300mmHg. - Inflator bulb with release valve Packaging and labelling: Individually wrapped in a plastic bag with a manufacturer manual Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001
241 | Medical supplies | ROLLGAUZE_10 | Gauze, 8cm x 4m, roll, non-sterile, P10 | Gauze, 8cm x 4m, roll, non-sterile, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001
242 | Medical supplies | ROLLGAUZE_12 | Gauze, 8cm x 4m, roll, non-sterile, P12 | Gauze, 8cm x 4m, roll, non-sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001
243 | Medical supplies | ROLLGAUZE_15 | Gauze, 8cm x 4m, roll, non-sterile, P15 | Gauze, 8cm x 4m, roll, non-sterile, P15 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001
244 | Medical supplies | ROLLGAUZE_5 | Gauze, 8cm x 4m, roll, non-sterile, P5 | Gauze, 8cm x 4m, roll, non-sterile, P5 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Dressing material used for holding a compress in place or to cover or isolate a wound. Material: Absorbent gauze, 100 % cotton. Bleached, purified textile, plain weave. Gauze bandage with selvage. Non-elastic, non-adhesive, non-detectable by X-ray. Thread count: Warp 12 threads / cm, weft 8 threads / cm. Weight: Approx. 27.5g / m2. Width: approx. 8 cm. Length: approx. 400 cm. Non-sterile. Single use. Hygenically packed for medical use. Shelf life to be included Instructions for use: Dressing material used for holding a compress in place or to cover or isolate a wound. Gauze bandage with selvedge must be selected to avoid remaining threads on wounds, burns etc. The size has been chosen as being the most commonly used. Safety process: The gauze bandage is for single use only. Collect and destroy by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Storage conditions, disposal instructions etc Packaging and labelling: Primary packaging: Unit of use. One (1) gauze bandage in a plastic bag. Secondary packaging: Protected unit. Ten (10) gauze bandages in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must be compliant as per the latest revisions of the standards): BS EN 14079:2003 Non-active medical devices. Performance requirements and test methods for absorbent cotton gauze and absorbent cotton and viscose gauze EN 13726-1:2002 Test methods for primary wound dressings - Part 1: Aspects of absorbency or equivalent standards such as USP Environmental requirements: ISO 14001
245 | Medical supplies | SUCTIONTUBEYANKAU | Tube, suction, Yankauer, 28cm | Tube, suction, Yankauer, 28cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to provide aspiration (suction) to a surgical site and this device is connected to an aspiration device using suction tubing. Product description: Suction tube with removable tip and tubing connector used for aspirating fluids/secretions from the operative area. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic stell composition: 16 to 18% chromium; 10 to 14% nickel; 2 to 3% molybdenum. Curved. Length: approx. 280mm. Diameter: approx. 6mm or 8mm. Autoclavable Distal end of the cannula must be atraumatic Instructions for use: Used for aspirating fluids / secretions from the operative area. Safety process: This item must be cleaned, disinfected after each use,and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Cleaning, disinfection and sterilization instructions Packaging and labelling: Primary packaging: One (1) Suction Tube in a plastic bag Secondary packaging: Protected unit. Ten (10) suction tubes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to provide aspiration (suction) to a surgical site and this device is connected to an aspiration device using suction tubing. Product description: Suction tube with removable tip and tubing connector used for aspirating fluids/secretions from the operative area. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic stell composition: 16 to 18% chromium; 10 to 14% nickel; 2 to 3% molybdenum. Curved. Length: approx. 280mm. Diameter: approx. 6mm or 8mm. Autoclavable Distal end of the cannula must be atraumatic Instructions for use: Used for aspirating fluids / secretions from the operative area. Safety process: This item must be cleaned, disinfected after each use,and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Cleaning, disinfection and sterilization instructions Packaging and labelling: Primary packaging: One (1) Suction Tube in a plastic bag Secondary packaging: Protected unit. Ten (10) suction tubes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Environmental requirements: ISO 14001
246 | Medical supplies | SUTURA2.3/8.26_12 | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 2 (3-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular Length of needle: approx. 26mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 2 (3-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular Length of needle: approx. 26mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
247 | Medical supplies | SUTURA2.3/8.26_36 | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile | Suture, absorbable, DEC 2 (3-0), 3/8, 26mm, triangular, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 2 (3-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular Length of needle: approx. 26mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 2 (3-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular Length of needle: approx. 26mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
248 | Medical supplies | SUTURA3.1/2.30_12 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P12 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 1/2 circle. Point: round. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 1/2 circle. Point: round. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
249 | Medical supplies | SUTURA3.1/2.30_36 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P36 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 1/2 circle. Point: round. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. Length of thread: approx. 75cm. Type of needle: curved needle, 1/2 circle. Point: round. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
250 | Medical supplies | SUTURA3.3/8.30_12 | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P12 | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
251 | Medical supplies | SUTURA3.3/8.30_36 | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P36 | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
252 | Medical supplies | SUTURA3.3/8.50_12 | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P12 | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol) Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) The words "for single use" (or equivalent harmonised symbol) The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol) CE mark with Notified Body number Secondary packaging: Protected unit One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol) Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol) Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) The words "for single use" (or equivalent harmonised symbol) The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol) CE mark with Notified Body number Secondary packaging: Protected unit One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol) Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
253 | Medical supplies | SUTURA3.3/8.50_36 | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P36 | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol) Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) The words "for single use" (or equivalent harmonised symbol) The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol) CE mark with Notified Body number Secondary packaging: Protected unit One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol) Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 3 (2-0). Must comply to USP or Eur Ph - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: round (non-cutting) has no edges that can cut through tissue - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. Round needle is less traumatic for the majority of tissues. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol) Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol) The words "for single use" (or equivalent harmonised symbol) The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol) CE mark with Notified Body number Secondary packaging: Protected unit One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging Extra information required: - Number of units per secondary packaging Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol) Information for handling, if applicable (or equivalent harmonised symbol) Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
254 | Medical supplies | SUTURA3.SPOOL_12 | Suture, absorbable, DEC 3 (2-0), spool, sterile, P12 | Suture, absorbable, DEC 3 (2-0), spool, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Ligatures (thread without needle) used to tie up a bleeding artery (for hemostasis). Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 150cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Must comply to USP or Eur Ph Thread without needle. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread has a length of 150 cm on average. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, storage conditions, disposal instructions, contraindications, precautions for use etc shall be provided. Packaging and labelling: shall be as per regulations applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Ligatures (thread without needle) used to tie up a bleeding artery (for hemostasis). Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 150cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Must comply to USP or Eur Ph Thread without needle. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread has a length of 150 cm on average. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, storage conditions, disposal instructions, contraindications, precautions for use etc shall be provided. Packaging and labelling: shall be as per regulations applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
255 | Medical supplies | SUTURA3.SPOOL_36 | Suture, absorbable, DEC 3 (2-0), spool, sterile, P36 | Suture, absorbable, DEC 3 (2-0), spool, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Ligatures (thread without needle) used to tie up a bleeding artery (for hemostasis). Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 150cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Must comply to USP or Eur Ph Thread without needle. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread has a length of 150 cm on average. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, storage conditions, disposal instructions, contraindications, precautions for use etc shall be provided. Packaging and labelling: shall be as per regulations applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Ligatures (thread without needle) used to tie up a bleeding artery (for hemostasis). Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 150cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Must comply to USP or Eur Ph Thread without needle. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread has a length of 150 cm on average. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, storage conditions, disposal instructions, contraindications, precautions for use etc shall be provided. Packaging and labelling: shall be as per regulations applied to the device. Labelling shall have information such as Product name and reference number. Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
256 | Medical supplies | SUTURA4.1/2.30_12 | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P12 | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 75cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC4 corresponding to 1. Type of needle: Shape: curved needle, 1/2 circle. Point: round. Length: approximately 30mm. Disposable. Sterile. Must comply to USP or Eur Phar. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 75cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC4 corresponding to 1. Type of needle: Shape: curved needle, 1/2 circle. Point: round. Length: approximately 30mm. Disposable. Sterile. Must comply to USP or Eur Phar. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
257 | Medical supplies | SUTURA4.1/2.30_36 | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P36 | Suture, absorbable, DEC 4 (1), 1/2, 30mm, round, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 75cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC4 corresponding to 1. Type of needle: Shape: curved needle, 1/2 circle. Point: round. Length: approximately 30mm. Disposable. Sterile. Must comply to USP or Eur Phar. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Technical Specifications: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Length: Approximately 75cm. Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC4 corresponding to 1. Type of needle: Shape: curved needle, 1/2 circle. Point: round. Length: approximately 30mm. Disposable. Sterile. Must comply to USP or Eur Phar. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Needle type and size. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
258 | Medical supplies | SUTURA4.3/8.36_12 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P12 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 corresponding to size 1. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: triangular in section (to pass through the skin more easily) - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 36mm. Minimum 5-year shelf life. Sterile and single use. Must comply to the USP or Eur Phar Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needle is triangular in section , to pass through the skin more easily. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 corresponding to size 1. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: triangular in section (to pass through the skin more easily) - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 36mm. Minimum 5-year shelf life. Sterile and single use. Must comply to the USP or Eur Phar Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needle is triangular in section , to pass through the skin more easily. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
259 | Medical supplies | SUTURA4.3/8.36_36 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P36 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 corresponding to size 1. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: triangular in section (to pass through the skin more easily) - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 36mm. Minimum 5-year shelf life. Sterile and single use. Must comply to the USP or Eur Phar Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needle is triangular in section , to pass through the skin more easily. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: A suture holds a wound together, in good apposition, until such time as the natural healing process is sufficiently well-advanced as to render support from suture materials unnecessary and redundant. Thread: - Type : Absorbable, synthetic braided thread - Material: PGA, polyglycolic acid, polyglactine. - Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 corresponding to size 1. - Length: approx. 75cm. Needle: - Type: eyeless with attached needle, thread mounted on needle - Point: triangular in section (to pass through the skin more easily) - Shape: curved needle, 3/8 circle. - Material: treated stainless steel - Length: approx. 36mm. Minimum 5-year shelf life. Sterile and single use. Must comply to the USP or Eur Phar Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Instructions for use: For internal suture. The thread with needle has a length of 75 cm on average. The needle is triangular in section , to pass through the skin more easily. This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. It has been chosen in preference to catgut, which is used less and less often in any case. Catgut is absorbed in a few days, sometimes less than 48 hours, and occasionally causes inflammation or its rejection. There is a risk of contamination because catgut is made of sheep gut (sheep trembles). Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Safety process: The suture is for single use only. Do not re-sterilize: once opened discard all unused sutures. When suturing: use strict aseptic techniques. Protection of users: To prevent stick injuries: Sutures with curved needles and with round-point needles are less dangerous than sutures with straight needles and with triangular-point needles. Use surgical gloves, a needle holder and a tissue standard forceps for suturing. After use, put the used sutures with needles into a sealed container. Collect and destroy the sealed containers by incineration in a controlled environment. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Labelling on primary packaging Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Product description and thread composition Thread gauge size needle type and size The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the integrity of the individual sterilised pack before use" ( or equivalent harmonized symbol). CE mark with Notified Body number Secondary packaging: Protected unit. One (1) box of 12 absorbable sutures Labelling on a secondary packaging: Labelling to be the same as primary packaging. Extra information required: - Number of units per secondary packaging. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Manufacturer's instruction for use. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
260 | Medical supplies | SUTURA4.3/8.50_36 | Suture, absorbable, DEC 4 (1), 3/8, 50mm, round, sterile, P36 | Suture, absorbable, DEC 4 (1), 3/8, 50mm, round, sterile, P36 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 (1). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: round Length of needle: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number and in the old system of thread numbering DEC 4 (1). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: round Length of needle: approx. 50mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
261 | Medical supplies | SUTUREA1.5_FISTULA | Suture, absorbable, DEC 1.5 (4/0), with needle, fistula repair | Suture, absorbable, DEC 1.5 (4/0), with needle, fistula repair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 1.5 in the old system of thread size corresponding to 4-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 1.5 in the old system of thread size corresponding to 4-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
262 | Medical supplies | SUTUREA2_FISTULA | Suture, absorbable, DEC 2 (3/0), with needle, fistula repair | Suture, absorbable, DEC 2 (3/0), with needle, fistula repair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 2 in the old system of thread size corresponding to 3-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days.Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 2 in the old system of thread size corresponding to 3-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days.Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
263 | Medical supplies | SUTUREA3_FISTULA | Suture, absorbable, DEC 3 (2/0), with needle, fistula repair | Suture, absorbable, DEC 3 (2/0), with needle, fistula repair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 3 in the old system of thread size corresponding to 2-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 3 in the old system of thread size corresponding to 2-0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
264 | Medical supplies | SUTUREA3.5_FISTULA | Suture, absorbable, DEC 3.5 (0), with needle, fistula repair | Suture, absorbable, DEC 3.5 (0), with needle, fistula repair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 3.5 in the old system of thread size corresponding to 0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended ourpose of use: intended to join the edges of a soft-tissue wound or incision by stitching or to ligate soft tissues. Product description: Type of thread: Absorbable, synthetic braided thread (PGA, polyglycolic acid, polyglactine). Gauge of thread: expressed in decimal gauge DEC number 3.5 in the old system of thread size corresponding to 0 Material: Thread: polyglactin 910 (copolymer composed of 90% glycoside and 10% L-lacticide)Thread coating: mixture of equal parts copolymer of glycoside and lactide (polyglactin 370) and calcium stearate. Needle: treated stainless steel Needle shape: 1/2 circle, round bodied Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or EU Phar Instructions for use: For peritoneal, digestive suture. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
265 | Medical supplies | SUTUREA4_FISTULA | Suture, absorbable, DEC 4 (1), with needle, fistula repair | Suture, absorbable, DEC 4 (1), with needle, fistula repair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: re-fixation of the pubo-cervical fascia. Product description: Type of thread: Absorbable, monofilament, synthetic braided Polydioxanone (PDO) suture for re-fixation of the pubo-cervical fascia. PDSII 1 or equivalent Material: Thread: Prepared from polyester and (p-dioxanone) Needle: treated stainless steel Needle shape: 1/2 circle, round bodied. Needle length: 35mm approx or equivalent suitable for intended purpose Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For re-fixation of the pubo-cervical fascia.The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: re-fixation of the pubo-cervical fascia. Product description: Type of thread: Absorbable, monofilament, synthetic braided Polydioxanone (PDO) suture for re-fixation of the pubo-cervical fascia. PDSII 1 or equivalent Material: Thread: Prepared from polyester and (p-dioxanone) Needle: treated stainless steel Needle shape: 1/2 circle, round bodied. Needle length: 35mm approx or equivalent suitable for intended purpose Minimum 5-year shelf life. Sterile and single use. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For re-fixation of the pubo-cervical fascia.The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc shall be provided. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia: Ph. Eur. Monograph 0667 USP : Absorbable surgical suture Environmental requirements: ISO 14001
266 | Medical supplies | SUTUREN2.3/8_12 | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Non-absorbable, monofilament suture with swaged needle. Type of thread: non-absorbable, synthetic monofilament (nylon, polyester, polyamide, polypropylene). Gauge of thread: expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC 2 (3-0). Length of thread: approx. 750 mm. Type of needle: atraumatic curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 18 mm Minimum 5-year shelf life. Sterile and single use. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Non-absorbable, monofilament suture with swaged needle. Type of thread: non-absorbable, synthetic monofilament (nylon, polyester, polyamide, polypropylene). Gauge of thread: expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC 2 (3-0). Length of thread: approx. 750 mm. Type of needle: atraumatic curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 18 mm Minimum 5-year shelf life. Sterile and single use. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
267 | Medical supplies | SUTUREN2.3/8_36 | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | Suture, non-absorbable, DEC 2 (3-0), 3/8, triangular, sterile, | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Non-absorbable, monofilament suture with swaged needle. Type of thread: non-absorbable, synthetic monofilament (nylon, polyester, polyamide, polypropylene). Gauge of thread: expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC 2 (3-0). Length of thread: approx. 750 mm. Type of needle: atraumatic curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 18 mm Minimum 5-year shelf life. Sterile and single use. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Non-absorbable, monofilament suture with swaged needle. Type of thread: non-absorbable, synthetic monofilament (nylon, polyester, polyamide, polypropylene). Gauge of thread: expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC 2 (3-0). Length of thread: approx. 750 mm. Type of needle: atraumatic curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 18 mm Minimum 5-year shelf life. Sterile and single use. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
268 | Medical supplies | SUTURENAPA3.1/2.R | Suture, non-absorbable PA, DEC 3 (2/0), 1/2 circle, round | Suture, non-absorbable PA, DEC 3 (2/0), 1/2 circle, round | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non absorbable, synthetic monofilament (polyamide) for skin closure size: USP 2-0 Type of needle: Shape: curved needle, 1/2 circle. Point: round. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Must comply to USP or Eu Phar. Instructions for use: For closure of the abdominal fascia. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non absorbable, synthetic monofilament (polyamide) for skin closure size: USP 2-0 Type of needle: Shape: curved needle, 1/2 circle. Point: round. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Must comply to USP or Eu Phar. Instructions for use: For closure of the abdominal fascia. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
269 | Medical supplies | SUTURENAPA4.1/2.R | Suture, non-absorbable PA, DEC 4 (1), 1/2 circle, round | Suture, non-absorbable PA, DEC 4 (1), 1/2 circle, round | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non absorbable, synthetic monofilament (polyamide) for closure of the abdominal fascia. Size:1 Type of needle: Shape: curved needle, 1/2 circle. Point: round. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Must comply to USP or Eu Phar. Instructions for use: For closure of the abdominal fascia. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non absorbable, synthetic monofilament (polyamide) for closure of the abdominal fascia. Size:1 Type of needle: Shape: curved needle, 1/2 circle. Point: round. Disposable. Sterile. Initial sterilisation method: Ethylene oxide gas / Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Must comply to USP or Eu Phar. Instructions for use: For closure of the abdominal fascia. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). The suture is synthetic, of slow absorption type, which is absorbed after about 21 days. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
270 | Medical supplies | SUTURN3.3/8.30_12 | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non-absorbable, synthetic monofilament (Nylon, polyester, polyamide, polypropylene) Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non-absorbable, synthetic monofilament (Nylon, polyester, polyamide, polypropylene) Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2-0). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
271 | Medical supplies | SUTURN3.3/8.30_36 | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | Suture, non-absorbable, DEC 3 (2-0), 3/8, triang., 30mm,sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non-absorbable, synthetic monofilament (Nylon, polyester, polyamide, polypropylene) Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2/0). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: intended to join the edges of a soft-tissue wound or incision by stitching, non absorbable type. Product description: Type of thread: Non-absorbable, synthetic monofilament (Nylon, polyester, polyamide, polypropylene) Gauge of thread: Expressed in Decimal gauge DEC number and in the old system of thread numbering: DEC3 (2/0). Length of thread: approx. 75cm. Type of needle: curved needle, 3/8 circle. Point: triangular. Length of needle: approx. 30mm. Minimum 5-year shelf life. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Suture elicits minimal acute inflammatory reactions and shall not adhere to the tissues while suture removal Initial sterilization method: Ethylene oxide gas/Gamma irradiation. Must comply to USP or Eu Phar. Instructions for use: For tendons, skin face. The thread with needle has a length of 75 cm on average. The needles are either triangular in section (to pass through the skin more easily) or round (less traumatic for the majority of tissues). This suture is inert thread, synthetic, monofilament. Silk and linen have been abandoned in favour of this thread type, which reduces the risk of inflammation or rejection. Safety process: The suture is for single use only. When suturing: use strict aseptic techniques. Supplied with: Manufacturer's instruction for use. Intended use, contraindications, warnings, precautions for use etc. Storage:Maximum temperature 25°C, protect from direct sunlight or as per manufacturer's instructions. warnings such as: Do not re-sterilize: once opened discard all unused sutures and ligatures. Re-sterilizing causes materials to lose their strength and pliability. Packaging and labelling: as per regualtions applied to the device. Labelling shall have information such as Product name and reference number. Needle size and shape, Product description and thread composition. Thread gauge size in USP and EP. Sterile and Sterilization method (ethylene oxide). Expiration date, lot and batch number. CE mark and reference number of notifying body. Primary packaging: Unit of use. One (1) Non-absorbable suture in an individual sterilized peel pack Secondary packaging: Protected unit. One (1) box of 36 non-absorbable sutures Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of the standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems European Pharmacopoeia or USP Environmental requirements: ISO 14001
272 | Medical supplies | SWAB.ALCOHOL | Swab, alcohol, sterile, single use | Swab, alcohol, sterile, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for reliable site preparation for injections, other laboratory and household uses. Product description: Swab with 70% v/v isopropyl alcohol solution. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation Minimum 4 log reduction for disinfection purpose. Instructions for use: Alcohol Swabs help fight bacteria and cleanse the skin. To be used when prepping the skin for an injection, as well as cleaning wounds before bandaging. Intended use, warnings, storage instructions and disposal instructions shall be provided Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Packaging and labelling: shall be as per regulations applied to the device. Primary packaging: Unit of use. One (1) alcohol swab in an individual sealed sterilized peel pack which provides integrity of the pack untill opened Labelling shall have information such as name and physical address of manufacturer; b) name of the product; c) instructions for use; d) swab size e) % alcohol content; f) precaution, “for external use only” “for single use”; and “don’t use when package is damaged”, g) lot/batch number; h) country of origin; i) date of manufacture; and j) expiry date i)CE mark with NB number Secondary packaging: Protected unit. One (1) box of 100 individual alcohol swab Symbols used according ISO 15223 Safety & product Standards: Must comply with the following standards (must comply to the latest revision of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems BS EN 16615:2015 - Chemical disinfectants and antiseptics. Quantitative test method for the evaluation of bactericidal and yeasticidal activity or EN 13727 Chemical disinfectants and antiseptics - Quantitative suspension test for the evaluation of bactericidal activity in the medical area or equivalent | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for reliable site preparation for injections, other laboratory and household uses. Product description: Swab with 70% v/v isopropyl alcohol solution. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma irradiation Minimum 4 log reduction for disinfection purpose. Instructions for use: Alcohol Swabs help fight bacteria and cleanse the skin. To be used when prepping the skin for an injection, as well as cleaning wounds before bandaging. Intended use, warnings, storage instructions and disposal instructions shall be provided Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Packaging and labelling: shall be as per regulations applied to the device. Primary packaging: Unit of use. One (1) alcohol swab in an individual sealed sterilized peel pack which provides integrity of the pack untill opened Labelling shall have information such as name and physical address of manufacturer; b) name of the product; c) instructions for use; d) swab size e) % alcohol content; f) precaution, “for external use only” “for single use”; and “don’t use when package is damaged”, g) lot/batch number; h) country of origin; i) date of manufacture; and j) expiry date i)CE mark with NB number Secondary packaging: Protected unit. One (1) box of 100 individual alcohol swab Symbols used according ISO 15223 Safety & product Standards: Must comply with the following standards (must comply to the latest revision of standards): ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems BS EN 16615:2015 - Chemical disinfectants and antiseptics. Quantitative test method for the evaluation of bactericidal and yeasticidal activity or EN 13727 Chemical disinfectants and antiseptics - Quantitative suspension test for the evaluation of bactericidal activity in the medical area or equivalent
273 | Medical supplies | SYR0.5MLW/NDL_100 | Syringe, 0.5ml, w/permanently att. needle,sterile,single use | Syringe, 0.5ml, w/permanently att. needle,sterile,single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For accurate administration of low volume injections such as VItamin K for newborns and infants Product Description: Syringe (0.5ml) preaffixed permanently to a needle with a protective cap Sterile, single use, disposable. Sterilization method: ethylene oxide or gamma radiation. Non-toxic. Pyrogen free. Colour code: visible at the base of the needle and complies with ISO 6009 Shelf life 5 yrs Syringe: Capacity: 0.5 ml. Graduated scale on the barrel with 0.01 ml (minor) and 0.1 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Needle: Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Normal and thin walled. External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 27G (approx. 0.4 x 13mm) Needle tip: regular bevel. Silicone coated to facilitate insertion. Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the needle, and preserves the sterility of the needle. Material: polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For intramuscular (IM) injection: children, newborns. To administer low volume injectables, such as Vitamin K 2mg/0.2ml solution for newborns (0.040 to 0.100 ml volumes). This syringe type is typically used for insulin or tuberculin injections. Safety process: The needle is for single use only. Use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) syringe packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) box of hundred (100) blister packed syringes. Primary package label includes: Device intity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquidand gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 7886-1:1993 Syringes for single-use ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management system | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For accurate administration of low volume injections such as VItamin K for newborns and infants Product Description: Syringe (0.5ml) preaffixed permanently to a needle with a protective cap Sterile, single use, disposable. Sterilization method: ethylene oxide or gamma radiation. Non-toxic. Pyrogen free. Colour code: visible at the base of the needle and complies with ISO 6009 Shelf life 5 yrs Syringe: Capacity: 0.5 ml. Graduated scale on the barrel with 0.01 ml (minor) and 0.1 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Needle: Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Normal and thin walled. External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 27G (approx. 0.4 x 13mm) Needle tip: regular bevel. Silicone coated to facilitate insertion. Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the needle, and preserves the sterility of the needle. Material: polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For intramuscular (IM) injection: children, newborns. To administer low volume injectables, such as Vitamin K 2mg/0.2ml solution for newborns (0.040 to 0.100 ml volumes). This syringe type is typically used for insulin or tuberculin injections. Safety process: The needle is for single use only. Use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) syringe packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) box of hundred (100) blister packed syringes. Primary package label includes: Device intity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquidand gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 7886-1:1993 Syringes for single-use ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management system
274 | Medical supplies | SYR0.5MLW/NDL_100 (ADS) | Auto Disabled Syringe, 0.5ml, w/permanently att. needle,sterile,single use | Auto Disabled Syringe, 0.5ml, w/permanently att. needle,sterile,single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Auto Disabled Syringe (0.5ml) preaffixed permanently to a needle with a protective cap. For accurate administration of low volume injectables. Sterile, single use, disposable. Sterilization method: ethylene oxide or gamma radiation. Non-toxic. Pyrogen free. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908.Fixed volume. Primary package consists of an individual blister pack. Pre-sterilized. Syringe: Capacity: 0.5 ml. Graduated scale on the barrel with 0.01 ml (minor) and 0.1 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Needle: Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Normal and thin walled. External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 27G (approx. 0.4 x 13mm) Needle tip: regular bevel. Silicone coated to facilitate insertion. Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the needle, and preserves the sterility of the needle. Material: polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For intramuscular (IM) injection: children, newborns. To administer low volume injectables, such as Vitamin K 2mg/0.2ml solution for newborns (0.040 to 0.100 ml volumes). This syringe type is typically used for insulin or tuberculin injections. Safety process: The needle is for single use only. Use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) syringe packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) box of hundred (100) blister packed syringes. Primary package label includes: Device intity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquidand gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management s | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Auto Disabled Syringe (0.5ml) preaffixed permanently to a needle with a protective cap. For accurate administration of low volume injectables. Sterile, single use, disposable. Sterilization method: ethylene oxide or gamma radiation. Non-toxic. Pyrogen free. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908.Fixed volume. Primary package consists of an individual blister pack. Pre-sterilized. Syringe: Capacity: 0.5 ml. Graduated scale on the barrel with 0.01 ml (minor) and 0.1 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Needle: Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Normal and thin walled. External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 27G (approx. 0.4 x 13mm) Needle tip: regular bevel. Silicone coated to facilitate insertion. Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the needle, and preserves the sterility of the needle. Material: polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For intramuscular (IM) injection: children, newborns. To administer low volume injectables, such as Vitamin K 2mg/0.2ml solution for newborns (0.040 to 0.100 ml volumes). This syringe type is typically used for insulin or tuberculin injections. Safety process: The needle is for single use only. Use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) syringe packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) box of hundred (100) blister packed syringes. Primary package label includes: Device intity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquidand gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management s
275 | Medical supplies | SYRINGE.10ML | Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended pupose of use: To deliver an accurate dose of a drug through injection or for various other purposes. Product description: Syringe: Two pieces: barrel with Luer nozzle and piston OR three pieces: barrel with luer nozzle, piston and stopper. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 10ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Easily read graduations: every 0.01 ml (1 ml syringe), every 0.1 ml (3 ml syringe), every 0.2 ml (5 ml syringe), every 0.5 ml (10 ml syringe) Concentric or eccentric luer nozzle. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years Safety mechanism, if offered: Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose, shall remain safe through disposal, WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container" Supplied with: Manufacturer's instruction for use, intended use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 7886-4: 2019: Sterile hypodermic syringes for single use - Part 4: Syringes with re-use prevention feature ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended pupose of use: To deliver an accurate dose of a drug through injection or for various other purposes. Product description: Syringe: Two pieces: barrel with Luer nozzle and piston OR three pieces: barrel with luer nozzle, piston and stopper. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 10ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Easily read graduations: every 0.01 ml (1 ml syringe), every 0.1 ml (3 ml syringe), every 0.2 ml (5 ml syringe), every 0.5 ml (10 ml syringe) Concentric or eccentric luer nozzle. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years Safety mechanism, if offered: Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose, shall remain safe through disposal, WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container" Supplied with: Manufacturer's instruction for use, intended use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 7886-4: 2019: Sterile hypodermic syringes for single use - Part 4: Syringes with re-use prevention feature ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
276 | Medical supplies | SYRINGE.10ML (ADS) | Auto Disabled Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | Auto Disabled Syringe, luer, 10 ml, w/21G needle, sterile, single use,P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 10ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 10ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
277 | Medical supplies | SYRINGE.2ML | Syringe, luer, 2 ml, w/21G needle, sterile, single use, P100 | Syringe, luer, 2 ml, w/21G needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 2ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 2ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 2ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 2ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
278 | Medical supplies | SYRINGE.2ML (ADS) | Auto Disabled Syringe, luer, 2ml, w/21G needle, sterile, single use,P100 | Auto Disabled Syringe, luer, 2ml, w/21G needle, sterile, single use,P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 2ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 2ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
279 | Medical supplies | SYRINGE.50CAT_1 | Syringe, feeding, catheter tip, 50 ml, sterile, P1 | Syringe, feeding, catheter tip, 50 ml, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Product description: Syringe, feeding, 50ml, catheter tip, sterile 3 pieces: barrel, plunger and gasket. BARREL: clear and transparent, made of polypropylene, medical grade capacity: 50 or 60 ml easily read graduations every ml male cup connector/conical tip ending in a concentric or eccentric tip PLUNGER: Plunger slides inside the barrel. medical grade polypropylene, polyethylene perfectly fit to slide easily and smoothly inside the barrel Safe plunger backstop. GASKET: synthetic elastomer (latex free) Good seal gasket located at the end of the plunger. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. shelf life 5 years Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Indications for use, precuations for use, contraindications, warnings, storage , disposal instructions etc shall be provided Packaging and labelling: shall be as per regulations applied to the device Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 20 feeding syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN ISO 7886-1:1997 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Product description: Syringe, feeding, 50ml, catheter tip, sterile 3 pieces: barrel, plunger and gasket. BARREL: clear and transparent, made of polypropylene, medical grade capacity: 50 or 60 ml easily read graduations every ml male cup connector/conical tip ending in a concentric or eccentric tip PLUNGER: Plunger slides inside the barrel. medical grade polypropylene, polyethylene perfectly fit to slide easily and smoothly inside the barrel Safe plunger backstop. GASKET: synthetic elastomer (latex free) Good seal gasket located at the end of the plunger. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. shelf life 5 years Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Indications for use, precuations for use, contraindications, warnings, storage , disposal instructions etc shall be provided Packaging and labelling: shall be as per regulations applied to the device Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 1 feeding syringe or one (1) box of 10 feeding syringes Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN ISO 7886-1:1997 Environmental requirements: ISO 14001
280 | Medical supplies | SYRINGE.50CAT_1 (ADS) | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P1 | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 2ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. 3 pieces: barrel, plunger and gasket. Material of plunger: polypropylene (PP). Material of gasket: elastomer. Clear and transparent syringe barrel. Precise graduation in ml for the volume dosage. Permanent marking. Easy readability. Centric or eccentric catheter tip. Safe plunger backstop. Cylindrical. Plunger slides inside the barrel. Good seal gasket located at the end of the plunger. Capacity: 50ml or 60ml. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 1 feeding syringe or one (1) box of 10 feeding syringes Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
281 | Medical supplies | SYRINGE.50CAT_50 | Syringe, feeding, catheter tip, 50 ml, sterile, P50 | Syringe, feeding, catheter tip, 50 ml, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Product description: Syringe, feeding, 50ml, catheter tip, sterile 3 pieces: barrel, plunger and gasket. BARREL: clear and transparent, made of polypropylene, medical grade capacity: 50 or 60 ml easily read graduations every ml male cup connector/conical tip ending in a concentric or eccentric tip PLUNGER: Plunger slides inside the barrel. medical grade polypropylene, polyethylene perfectly fit to slide easily and smoothly inside the barrel Safe plunger backstop. GASKET: synthetic elastomer (latex free) Good seal gasket located at the end of the plunger. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. shelf life 5 years Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Indications for use, precuations for use, contraindications, warnings, storage , disposal instructions etc shall be provided Packaging and labelling: shall be as per regulations applied to the device Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 20 feeding syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN ISO 7886-1:1997 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Product description: Syringe, feeding, 50ml, catheter tip, sterile 3 pieces: barrel, plunger and gasket. BARREL: clear and transparent, made of polypropylene, medical grade capacity: 50 or 60 ml easily read graduations every ml male cup connector/conical tip ending in a concentric or eccentric tip PLUNGER: Plunger slides inside the barrel. medical grade polypropylene, polyethylene perfectly fit to slide easily and smoothly inside the barrel Safe plunger backstop. GASKET: synthetic elastomer (latex free) Good seal gasket located at the end of the plunger. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. shelf life 5 years Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Indications for use, precuations for use, contraindications, warnings, storage , disposal instructions etc shall be provided Packaging and labelling: shall be as per regulations applied to the device Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 50 feeding syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN ISO 7886-1:1997 Environmental requirements: ISO 14001
282 | Medical supplies | SYRINGE.50CAT_50 (ADS) | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P50 | Auto Disabled Syringe, feeding, catheter tip, 50 ml, sterile, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 2ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Sterile syringe used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. 3 pieces: barrel, plunger and gasket. Material of plunger: polypropylene (PP). Material of gasket: elastomer. Clear and transparent syringe barrel. Precise graduation in ml for the volume dosage. Permanent marking. Easy readability. Centric or eccentric catheter tip. Safe plunger backstop. Cylindrical. Plunger slides inside the barrel. Good seal gasket located at the end of the plunger. Capacity: 50ml or 60ml. Sterile and single use. Initial sterilization method: Ethylene oxide gas/Gamma radiation. Instructions for use: The sterile syringe 50/60ml catheter tip can be used for gastroenteral feeding, manual aspiration of liquids or gas from stomach, manual pharyngeal or tracheal aspiration/suction. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: one (1) feeding syringe in an individual sterilized peel pack. Secondary packaging: Protected unit. One (1) box of 50 feeding syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
283 | Medical supplies | SYRINGE.5ML | Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 21G (0.80 x 38-40 mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , green for 21G (0.80 x 40 mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device
284 | Medical supplies | SYRINGE.5ML (ADS) | Auto Disabled Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 5 ml, w/21G needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap. Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 21G (0.80 x 38-40 mm). Needle with base (conical Luer type fitting) and protecting cap. Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
285 | Medical supplies | SYRINGE.5ML.23G | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P100 | Syringe, luer, 5ml, with needle, 23G, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
286 | Medical supplies | SYRINGE.5ML.23G (ADS) | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
287 | Medical supplies | SYRINGE.5ML.23G_1 | Syringe, luer, 5ml, w/23G needle, sterile, single use, P1 | Syringe, luer, 5ml, w/23G needle, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
288 | Medical supplies | SYRINGE.5ML.23G_1 (ADS) | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P1 | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
289 | Medical supplies | SYRINGE.5ML.23G_25 | Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Preferably with retractable needle to prevent reuse and needle stick injury Capacity: 5ml BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); transparent for easy reading and for detecting air bubbles. Graduated scale on the barrel, easy to read, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Concentric luer nozzle for 5ml capacity. piston : Medical grade polypropylene, polyethylene, fit to slide easily and smoothly inside the barrel needle: 23 G (0.6 x 30mm) , stainless steel, tapered bevel, lubricated, with protective cap HUB : plastic medical grade, colour coded as per ISO 6009 , blue for 23 G (0.6 x 30mm) SAFETY MECHANISM, if offered : Integrated and retractable, Needle retracts automatically into the barrel upon administration of full dose. WHO PQ'ed as per Category E013 Hypodermic syringe with reuse prevention feature (RUP). stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. Shelf life 5 years. Instructions for use as per the item features offered: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. Disposal instructions such as "NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container", as applicable to the item offered. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 25 bi-packed syringes with needles. Labelling same as per primary packaging Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use EN ISO 7864: 2016: Sterile hypodermic needles for single use ISO 6009: 2016: Hypodermic needles for single use - Colour coding for identification ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
290 | Medical supplies | SYRINGE.5ML.23G_25 (ADS) | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | Auto Disabled Syringe, luer, 5ml, w/23G needle, sterile, single use, P25 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 5ml Graduated scale on the barrel, with scale interval of 0.2 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Size selected: 23G Needle with base (conical Luer type fitting) and protecting cap Material: Needle: stainless steel; base and protecting cap: plastic. Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe and needle bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes with needles. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
291 | Medical supplies | SYRINGE20ML_1 | Syringe, luer lock, 20 ml, sterile, single use, P1 | Syringe, luer lock, 20 ml, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Materials for barrel and piston: medical grade plastic (Polyethylene (PE)/ polypropylene (PP)/ polystyrene (PS). The Luer Lock nozzle shall be situated centrally. plunger tip: the three part syringes end in a synthetic elastomer plunger tip or gasket, flat or dome shaped ensuring a perfect plunger barrel seal integrity and a flawless gliding behaviour. Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: shall be as per regulations applied to the device Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Materials for barrel and piston: medical grade plastic (Polyethylene (PE)/ polypropylene (PP)/ polystyrene (PS). The Luer Lock nozzle shall be situated centrally. plunger tip: the three part syringes end in a synthetic elastomer plunger tip or gasket, flat or dome shaped ensuring a perfect plunger barrel seal integrity and a flawless gliding behaviour. Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: shall be as per regulations applied to the device Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management systems
292 | Medical supplies | SYRINGE20ML_1 (ADS) | Auto Disabled Syringe, luer lock, 20 ml, sterile, single use, P1 | Auto Disabled Syringe, luer lock, 20 ml, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer Lock nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single usONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer Lock nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single usONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management systems
293 | Medical supplies | SYRINGE20ML_100 | Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Materials: medical grade plastic (Polyethylene (PE)/ polypropylene (PP)/ polystyrene (PS). The Luer Lock nozzle shall be situated centrally. plunger tip: the three part syringes end in a synthetic elastomer plunger tip or gasket, flat or dome shaped ensuring a perfect plunger barrel seal integrity and a flawless gliding behaviour. Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: shall be as per regulations applied to the device Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: inject fluids into, or withdraw fluids from, the body Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Materials: medical grade plastic (Polyethylene (PE)/ polypropylene (PP)/ polystyrene (PS). The Luer Lock nozzle shall be situated centrally. plunger tip: the three part syringes end in a synthetic elastomer plunger tip or gasket, flat or dome shaped ensuring a perfect plunger barrel seal integrity and a flawless gliding behaviour. Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: shall be as per regulations applied to the device Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: ISO 14001: Environmental management systems
294 | Medical supplies | SYRINGE20ML_100 (ADS) | Auto Disabled Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | Auto Disabled Syringe, luer lock, 20ml, w/o needle,sterile,single use,P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mecahnisms either standalone or in combination e.g. Auto Disposable (ADS), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standard ISO 7886 and ISO 23908. Primary package consistes individual blister pack. Pre-sterilized. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer Lock nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single usONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management systems | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product Description: Syringe, two pieces: barrel with Luer Lock nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Without needle. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mecahnisms either standalone or in combination e.g. Auto Disposable (ADS), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standard ISO 7886 and ISO 23908. Primary package consistes individual blister pack. Pre-sterilized. Capacity: 20 ml. Graduated scale on the barrel with 1.0 ml (minor) and 5.0 ml (major) increment lines. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer Lock nozzle shall be situated centrally. Material: Medical grade plastic; Polyethylene (PE), polypropylene (PP), polystyrene (PS). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Syringe is used with an injection needle with compatible luer fitting. Injection safety: The syringe is sterile, ready for immediate use and is for single usONLY. Check the integrity of the packaging before opening. Do not use the syringe if the package integrity is compromised. Packaging and Labelling: Primary packaging: One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Secondary packaging: One (1) carton box of hundred (100) syringes to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Product and packaging complies with WHO PQS E08. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management systems
295 | Medical supplies | SYRINGEBLADDER25 | Bladder syringe 60-100 mls (with long nozzle, not luer lock) | Bladder syringe 60-100 mls (with long nozzle, not luer lock) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for bladder irrigation Product description: Bladder syringe 100 ml Materials: Made of polypropylene (PP). 3-peice syringe: Barrel: Transparent with contrast graduation: permanent markings and good legibility. Plunger stopper with backstop. Double sealing ring on plunger. Graduation in ml for the volume dosage. With centric-catheter fitting cone for a tube with a funnel connection (e.g., gastric tube). With long nozzle Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use irrigation of the urinary bladder, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Packaging: 1 syringe, sterile, peel pack, single-use. Secondary packaging: Protected unit. One (1) box of 25 syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for bladder irrigation Product description: Bladder syringe 100 ml Materials: Made of polypropylene (PP). 3-peice syringe: Barrel: Transparent with contrast graduation: permanent markings and good legibility. Plunger stopper with backstop. Double sealing ring on plunger. Graduation in ml for the volume dosage. With centric-catheter fitting cone for a tube with a funnel connection (e.g., gastric tube). With long nozzle Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Shelf life 5 years Supplied with: Manufacturer's instruction for use irrigation of the urinary bladder, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Packaging: 1 syringe, sterile, peel pack, single-use. Secondary packaging: Protected unit. One (1) box of 25 syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety and Product Standards (shall comply to the current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for iscular or hypodermic applications. ISO 7886-1: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental requirements: ISO 14001
296 | Medical supplies | SYRINGEBLADDER25 (ADS) | Auto Disabled Bladder syringe 60-100 mls (with long nozzle, not luer lock) | Auto Disabled Bladder syringe 60-100 mls (with long nozzle, not luer lock) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Bladder syringe 60-100 mls for washing out the stomach or the administration of liquid foods and for bladder irrigation Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Materials: Made of polypropylene (PP). 3-peice syringe: Barrel: Transparent with contrast graduation: permanent markings and good legibility. Plunger stopper with backstop. Double sealing ring on plunger. Graduation in ml for the volume dosage. With centric-catheter fitting cone for a tube with a funnel connection (e.g., gastric tube). With long nozzle Instruction for use: Instrument for irrigation of the urinary bladder. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: 1 syringe, sterile, peel pack, single-use. Secondary packaging: Protected unit. One (1) box of 25 syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards: ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Bladder syringe 60-100 mls for washing out the stomach or the administration of liquid foods and for bladder irrigation Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Materials: Made of polypropylene (PP). 3-peice syringe: Barrel: Transparent with contrast graduation: permanent markings and good legibility. Plunger stopper with backstop. Double sealing ring on plunger. Graduation in ml for the volume dosage. With centric-catheter fitting cone for a tube with a funnel connection (e.g., gastric tube). With long nozzle Instruction for use: Instrument for irrigation of the urinary bladder. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: 1 syringe, sterile, peel pack, single-use. Secondary packaging: Protected unit. One (1) box of 25 syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards: ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
297 | Medical supplies | SYRINGEONLY10_100 | Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Capacity: 10ml, supplied without needle. BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Concentric or eccentric luer nozzle for 10ml capacity. piston/plunger : Medical grade polypropylene, fit to slide easily and smoothly inside the barrel stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. SHelf life 5 years Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Capacity: 10ml, supplied without needle. BARREL with finger tip: Materials: medical grade PE (polyethylene) or PP (polypropylene); Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. Concentric or eccentric luer nozzle for 10ml capacity. piston/plunger : Medical grade polypropylene, fit to slide easily and smoothly inside the barrel stopper (for 3-piece syringe): Material is elastomer, medical grade, Located at the end of the plunger, Ensures a secure and safe seal. Sterile and disposable. Sterilization method: ethylene oxide. SHelf life 5 years Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling as per regulations applied to the device: Primary packaging: each syringe bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems Environmental requirements: ISO 14001
298 | Medical supplies | SYRINGEONLY10_100 (ADS) | Auto Disabled Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 10ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Syringe, two pieces: barrel with Luer nozzle and piston or three pieces barrel with luer nozzle, piston and stopper. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Capacity: 10ml Graduated scale on the barrel, with scale interval of 0.5 and 1.0ml increment between graduation lines to be numbered. Barrel sufficiently transparent to allow easy measurement of the volume contained in the syringe and detection of air bubbles. The Luer nozzle shall be situated centrally. Material: Polyethylene (PEF) or polypropylene (PP). Sterile and disposable. Sterilization method: ethylene oxide. Instructions for use: Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: each syringe bi-packed in an individual sterilized peel-off pack made of paper and/or plastic. Secondary packaging: Protective packaging - 1 carton of 100 bi-packed syringes. Symbols used according ISO 15223 CE Mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems EN 20594-1:1994, ISO 594-1:1986 EN ISO 7886-1:1997 Product and packaging complies with WHO PQS E08. Environmental requirements: Polymer not PVC
299 | Medical supplies | SYRINGEONLY1ML | Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 1 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 1 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
300 | Medical supplies | SYRINGEONLY1ML (ADS) | Auto Disabled Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 1ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 1 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 1 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer
301 | Medical supplies | SYRINGEONLY2ML | Syringe, luer, 2ml, w/o needle, sterile, single use, P100 | Syringe, luer, 2ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 2 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 2 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
302 | Medical supplies | SYRINGEONLY2ML (ADS) | Auto Disabled Syringe, luer, 2ml, w/o needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 2ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - with finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval with intervals 0.5 and 1.0ml increment between graduation lines to be numbered. - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. - Concentric Luer nozzle Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 2 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilized peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonized symbol). - Sterilization method (or equivalent harmonized symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonized symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonized symbol). - The words "or single use" (or equivalent harmonized symbol). - The words "check the integrity of the individual sterilized pack before use" (or equivalent harmonized symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - with finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval with intervals 0.5 and 1.0ml increment between graduation lines to be numbered. - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. - Concentric Luer nozzle Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 2 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilized peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonized symbol). - Sterilization method (or equivalent harmonized symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonized symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonized symbol). - The words "or single use" (or equivalent harmonized symbol). - The words "check the integrity of the individual sterilized pack before use" (or equivalent harmonized symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer
303 | Medical supplies | SYRINGEONLY5ML | Syringe, luer, 5ml, w/o needle, sterile, single use, P100 | Syringe, luer, 5ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 5 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For injection and various other uses, including mixing. Product description: Syringe, two pieces: barrel with Luer nozzle and piston/plunger or three pieces: barrel with luer nozzle, stopper and piston/plunger. Barrel: - With finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval of 0.01 ml ( 3-pieces syringe) or 0.05 ml ( 2 -piece syringe) - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket (stopper): - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 5 ml, without needle Sterile, single use Sterilisation method: Ethylene oxide. Shelf life 5 yrs Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): EN ISO 7886-1: 2017: Sterile hypodermic syringes for single use - Part 1: Syringes for manual use ISO 80369-7:2021 Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001
304 | Medical supplies | SYRINGEONLY5ML (ADS) | Auto Disabled Syringe, luer, 5ml, w/o needle, sterile, single use, P100 | Auto Disabled Syringe, luer, 5ml, w/o needle, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - with finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval with intervals 0.5 and 1.0ml increment between graduation lines to be numbered. - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. - Concentric Luer nozzle Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 5 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: For injection and various other uses, including mixing. Safety Engineered Feature : The syringe has an auto-disable technology feature that passively activates upon starting administration of the prescribed injectable dose to prevent subsequent re-use of the syringe. Various mechanisms either standalone or in combination e.g. Auto Disposable (AD), Reusable Prevention (RUP), Sharp Injury Prevention (SIP), Retractable Needle (RTN) may be considered. Mandatory product compliance standards ISO 7886 and ISO 23908. Primary package consists of an individual blister pack. Pre-sterilized. Syringe, two pieces: barrel with Luer nozzle and plunger or three pieces barrel with luer nozzle plunger and piston. Barrel: - with finger grip. - Materials: medical grade PE (polyethylene) or PP (polypropylene); - transparent for easy reading and for detecting air bubbles. - Cylindrical shape. - Graduation scale: clear, water resistant, with scale interval with intervals 0.5 and 1.0ml increment between graduation lines to be numbered. - Ends in a male Luer tip (conical with 6% taper). - The Luer nozzle shall be situated centrally. - Concentric Luer nozzle Plunger: - Material: PP (polypropylene) - Slides inside the barrel. Gasket: - Material: elastomer, medical grade. - Located at the end of the plunger. - Ensures a secure and safe seal Capacity: 5 ml Sterile, single use Sterilisation method: Ethylene oxide. Instructions for use: For injection and various other uses, including mixing. Injection safety: The syringe is sterile, ready for immediate use and is for single use ONLY. Check the integrity of the packaging before opening. Do not use the syringe if the packaging is not sealed or is pierced. NEVER recap the needle after use and IMMEDIATELY dispose of the mounted syringe and needle into a puncture-proof safety container. Supplied with: Manufacturer's instruction for use. Accessories/ spare parts/consumables: - syringe needle Packaging and labelling: Primary packaging: -One (1) syringe packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Secondary packaging: Protected unit. 1 box of 100 blister packed syringes. Secondary packaging: - Protected unit, one (1) box 100-unit presentation; - labelling is the same as that of the primary packaging. Number of units per secondary packaging, information for product-specific storage conditions (e.g., temperature, pressure, light, humidity, etc.). Symbols used according ISO 15223 and EN 980 CE mark Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 7886-1: 1993/Coor 1 :1995 Sterile hypodermic syringe for single use: part 1: syringe for manual use ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: Not use PVC polymer
305 | Medical supplies | TAPE.ADHESIVE_1 | Tape, adhesive, zinc oxide, 2.5cm x 5m, P1 | Tape, adhesive, zinc oxide, 2.5cm x 5m, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 2 - 2.5cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, 2.5cm x 5m, for basic and small wound dressing. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 2 - 2.5cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, 2.5cm x 5m, for basic and small wound dressing. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001
306 | Medical supplies | TAPE.ADHESIVE_10 | Tape, adhesive, zinc oxide, 2.5cm x 5m, P10 | Tape, adhesive, zinc oxide, 2.5cm x 5m, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 2 - 2.5cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, 2.5cm x 5m, for basic and small wound dressing. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 2 - 2.5cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, 2.5cm x 5m, for basic and small wound dressing. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001
307 | Medical supplies | TAPE.ADHESIVEL_1 | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P1 | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 10cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, perforated, 10 cm x 5 m, for important wound dressing and surgery. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 10cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, perforated, 10 cm x 5 m, for important wound dressing and surgery. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001
308 | Medical supplies | TAPE.ADHESIVEL_10 | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P10 | Tape, adhesive, zinc oxide, perforated, 10cmx5m, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 10cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, perforated, 10 cm x 5 m, for important wound dressing and surgery. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For wound dressing and general securement of tubing and devices Product description: Components: High adhesion, breathable cloth tape coated with adhesive for a range of applications Hypoallergenic and skin friendly. Non irritant and non sensitizing. Non-stretch. Can be torn by hand. Water resistant. Must adhere strongly when applied to the skin, but can be removed without causing skin trauma, Easy to unroll till the tape core. Material: Textile strip: woven viscose, cotton or similar Adhesive: Rubber based, medical grade Traditionally incorporates zinc oxide Colour: White or flesh coloured. Length: approximately 5m. Width: approximately 10cm. Non-sterile and single use. SHelf life to be included Hygenically packed for medical use. Instructions for use: Adhesive tape used for fixing dressings and general securement of critical tubes and devices to the skin. Select the correct size and type of tape for the application intended. Clip hair, if necessary. Apply the tape without tension to clean, dry skin, gently but firmly stroking the surface to maximize adhesion. When fixing dressings, tape should extend at least one-half inch beyond the edge of the dressing to hold the dressing in place. Zinc oxide tape has been selected because it is inexpensive. The quality is difficult to judge and the adhesive can deteriorate seriously under exposure to heat and dampness. Make sure the storage conditions are observed, so as to obtain maximum use of the item and avoid waste. The size has been chosen as being the most commonly used. Tape adhesive zinc oxide, perforated, 10 cm x 5 m, for important wound dressing and surgery. Supplied with: Manufacturer's instruction for use. Indications for use, contraindications, warnings, storage and disposal instructions shall be provided Packaging and labelling: Primary packaging: Unit of use. One (1) adhesive tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) adhesive tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (must comply with latest revision of standards): ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 29862: Self adhesive tapes — Determination of peel adhesion properties or equivalent Complies to USP Environmental requirements: ISO 14001
309 | Medical supplies | TAPE.UMBILICAL_1 | Tape, umbilical, 3mmx50m, non-sterile, single use, P1 | Tape, umbilical, 3mmx50m, non-sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: thin, non-adhesive cotton tape for the tying of umbilical cords Product description: Roll of thread is used as a ligature thread to tie the umbilical cord. Non-absorbable. Braided and flat. Material: 100% cotton. Width: approximately 3mm. Length: approximately 50m. Non-sterile and single use. Hygenically packed for medical use Instructions for use: Umbilical tape is used as a ligature thread to tie the umbilical cord. To prevent cutting or damage to the umbilical cord: do not select monofilament, round or thin ligature threads. Select cotton to prevent untying of knots. This item can be subjected to steam sterilising. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) umbilical tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) umbilical tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( must comply to latest revisions of the standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: thin, non-adhesive cotton tape for the tying of umbilical cords Product description: Roll of thread is used as a ligature thread to tie the umbilical cord. Non-absorbable. Braided and flat. Material: 100% cotton. Width: approximately 3mm. Length: approximately 50m. Non-sterile and single use. Hygenically packed for medical use Instructions for use: Umbilical tape is used as a ligature thread to tie the umbilical cord. To prevent cutting or damage to the umbilical cord: do not select monofilament, round or thin ligature threads. Select cotton to prevent untying of knots. This item can be subjected to steam sterilising. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) umbilical tape in a plastic bag. Secondary packaging: Protected unit. Ten (10) umbilical tapes in a box. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( must comply to latest revisions of the standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO 14001
310 | Medical supplies | TAPEUMBIL15CM_1 | Tape, umbilical, 3mm x 15cm, non-sterile, P1 | Tape, umbilical, 3mm x 15cm, non-sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: thin, non-adhesive cotton tape for the tying of umbilical cords Product description: Umbilical tape is used as a ligature thread to tie the umbilical cord. Non-absorbable Braided and flat, without defects. Material: 100% cotton Thread. Wide: approx. 3 mm length approx. 15 cm Non-sterile and single use. Pack of 3 pieces Hygienically clean, for medical use. Instructions for use: Umbilical tape is used as a ligature thread to tie the umbilical cord. To prevent cutting or damage to the umbilical cord: do not select monofilament, round or thin ligature threads. Select cotton to prevent untying of knots. This item can be subjected to steam sterilising. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging; three ( 3) pieces of umbilical tape in a plastic bag. Labelling : Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( must comply to latest revisions of the standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: thin, non-adhesive cotton tape for the tying of umbilical cords Product description: Umbilical tape is used as a ligature thread to tie the umbilical cord. Non-absorbable Braided and flat, without defects. Material: 100% cotton Thread. Wide: approx. 3 mm length approx. 15 cm Non-sterile and single use. Pack of 3 pieces Hygienically clean, for medical use. Instructions for use: Umbilical tape is used as a ligature thread to tie the umbilical cord. To prevent cutting or damage to the umbilical cord: do not select monofilament, round or thin ligature threads. Select cotton to prevent untying of knots. This item can be subjected to steam sterilising. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging; three ( 3) pieces of umbilical tape in a plastic bag. Labelling : Name and/or trademark and address of the manufacturer. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( must comply to latest revisions of the standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO 14001
311 | Medical supplies | TOURNIQUET | Tourniquet, latex rubber, 75cm | Tourniquet, latex rubber, 75cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Elastic strip used to compress a limb to stop blood flow. May be solid or tubular. Material: Latex rubber Autoclavable at 121ºC Length: 75 - 100cm. Elastic strip solid, width: approx. 18 - 20mm. Elastic strip tubular, inner diameter: approx. 7mm. Elastic strip tubular, outer diameter: approx. 10mm. Reusable. Non-sterile. Instructions for use: A tourniquet is used to facilitate the puncture of veins. Supplied with: Manufacturer's instruction for use. cleaning and disinfection instructions shall be provided Packaging and labelling: One (1) Tourniquet in a plastic bag Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Elastic strip used to compress a limb to stop blood flow. May be solid or tubular. Material: Latex rubber Autoclavable at 121ºC Length: 75 - 100cm. Elastic strip solid, width: approx. 18 - 20mm. Elastic strip tubular, inner diameter: approx. 7mm. Elastic strip tubular, outer diameter: approx. 10mm. Reusable. Non-sterile. Instructions for use: A tourniquet is used to facilitate the puncture of veins. Supplied with: Manufacturer's instruction for use. cleaning and disinfection instructions shall be provided Packaging and labelling: One (1) Tourniquet in a plastic bag Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device
312 | Medical supplies | TUBESUCTICH10_1 | Tube, suction, CH10, 50 cm long, conical tip, sterile, P1 | Tube, suction, CH10, 50 cm long, conical tip, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for periodic aspiration of liquids and/or semi-solids from a patient's upper airway. Product description: Sterile suction catheter for single use. Material: Polyvinyl chloride (PVC) DEHP free Diameter size :Diameter expressed in Charriere, French gauge. Gauge CH10. Length expressed in cm. Length approx. 50cm Tube: Straight distal end with 2 side windows Cup connector: Conical tip. Colour code/ external diameter: Visible on cup connector. The suction tube consists of a single channel tube with open distal end with side windows. Proximal end with cup connector (conical tip) allowing the tube to be connected to devices such as suction pump system Disposable, single use Sterile Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile suction tube device used for aspirating pus, blood, secretions, food or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging. Do not use if the packaging is damaged. Safety process: The suction tube is for single use only. When inserting a suction tube: use clean techniques. Collect and destroy by incineration in a controlled. Supplied with: Manufacturer's instruction for use, indications for use (such as "This tube can be connected: either to a syringe with conical tip for manual aspiration or to a mechanical or electrical suction system"), contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Primary packaging: -One (1) suction tube packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of standards): ISO 8836: 2019: Suction catheters for use in the respiratory tract ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 80369-7, Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for periodic aspiration of liquids and/or semi-solids from a patient's upper airway. Product description: Sterile suction catheter for single use. Material: Polyvinyl chloride (PVC) DEHP free Diameter size :Diameter expressed in Charriere, French gauge. Gauge CH10. Length expressed in cm. Length approx. 50cm Tube: Straight distal end with 2 side windows Cup connector: Conical tip. Colour code/ external diameter: Visible on cup connector. The suction tube consists of a single channel tube with open distal end with side windows. Proximal end with cup connector (conical tip) allowing the tube to be connected to devices such as suction pump system Disposable, single use Sterile Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile suction tube device used for aspirating pus, blood, secretions, food or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging. Do not use if the packaging is damaged. Safety process: The suction tube is for single use only. When inserting a suction tube: use clean techniques. Collect and destroy by incineration in a controlled. Supplied with: Manufacturer's instruction for use, indications for use (such as "This tube can be connected: either to a syringe with conical tip for manual aspiration or to a mechanical or electrical suction system"), contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Primary packaging: -One (1) suction tube packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of standards): ISO 8836: 2019: Suction catheters for use in the respiratory tract ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 80369-7, Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications Environmental requirements: ISO 14001
313 | Medical supplies | TUBESUCTICH14_1 | Tube, suction, CH14, 50 cm long, conical tip, sterile, P1 | Tube, suction, CH14, 50 cm long, conical tip, sterile, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for periodic aspiration of liquids and/or semi-solids from a patient's upper airway. Product description: Sterile suction catheter for single use. Material: Polyvinyl chloride (PVC) DEHP free Diameter size :Diameter expressed in Charriere, French gauge. Gauge CH14. Length expressed in cm. Length approx. 50cm Tube: Straight distal end with 2 side windows Cup connector: Conical tip. Colour code/ external diameter: Visible on cup connector. The suction tube consists of a single channel tube with open distal end with side windows. Proximal end with cup connector (conical tip) allowing the tube to be connected to devices such as suction pump system Disposable, single use Sterile Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile suction tube device used for aspirating pus, blood, secretions, food or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging. Do not use if the packaging is damaged. Safety process: The suction tube is for single use only. When inserting a suction tube: use clean techniques. Collect and destroy by incineration in a controlled. Supplied with: Manufacturer's instruction for use, indications for use (such as "This tube can be connected: either to a syringe with conical tip for manual aspiration or to a mechanical or electrical suction system"), contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Primary packaging: -One (1) suction tube packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of standards): ISO 8836: 2019: Suction catheters for use in the respiratory tract ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 80369-7, Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for periodic aspiration of liquids and/or semi-solids from a patient's upper airway. Product description: Sterile suction catheter for single use. Material: Polyvinyl chloride (PVC) DEHP free Diameter size :Diameter expressed in Charriere, French gauge. Gauge CH14. Length expressed in cm. Length approx. 50cm Tube: Straight distal end with 2 side windows Cup connector: Conical tip. Colour code/ external diameter: Visible on cup connector. The suction tube consists of a single channel tube with open distal end with side windows. Proximal end with cup connector (conical tip) allowing the tube to be connected to devices such as suction pump system Disposable, single use Sterile Initial sterilisation method: Ethylene oxide gas. Shelf life to be included Instructions for use: Sterile suction tube device used for aspirating pus, blood, secretions, food or other liquids obstructing the pharynx or airways. The sterile suction tube device is introduced either via the mouth, nose, through the channel of endotracheal tube or tracheotomy tube. The size has been chosen as being the most commonly used. Conditions for stock: Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging. Do not use if the packaging is damaged. Safety process: The suction tube is for single use only. When inserting a suction tube: use clean techniques. Collect and destroy by incineration in a controlled. Supplied with: Manufacturer's instruction for use, indications for use (such as "This tube can be connected: either to a syringe with conical tip for manual aspiration or to a mechanical or electrical suction system"), contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Primary packaging: -One (1) suction tube packed in an individual sterilised peel-pack made of paper and/or plastic. Labelling on the primary packaging: - Name and/or trade mark, and address of f the manufacturer. - Manufacturer's product reference. - Type of product and main characteristics. - The word "sterile" (or equivalent harmonised symbol). - Sterilisation method (or equivalent harmonised symbol). - Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). - Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol). - The words "or single use" (or equivalent harmonised symbol). - The words "check the integrity of the individual sterilised pack before use" (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (must comply to the latest revisions of standards): ISO 8836: 2019: Suction catheters for use in the respiratory tract ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 80369-7, Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications Environmental requirements: ISO 14001
314 | Medical supplies | URINEBAG2L_10 | Bag, urine, collecting, 2 litres, P1 | Bag, urine, collecting, 2 litres, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: designed to connect to an indwelling urinary catheter to collect discharged urine from a patient Product description: The 2-litre, urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. Material for bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA), flexible. Tube & connector/protective cap: polyvinyl chloride (PVC) Urine bag has two (2) strong eyelets from which to suspend the bag Bag collector fitted with a non retun valve. The non-return valve is located inside the urine collector bag at the upper part or urine entry point, and prevents urine backflow into the indwelling urinary catheter. Tube with universal connector and protective cap to connect to a catheter. Outlet tube fitted with a drain valve with seal cap allowing to maintain sterility. Tube length: approx. 90cm. Capacity: 2000ml. Graduation: 100ml. Sterile, Single use. Shelf life 5 yrs Instructions for use: Used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Safety process: The urine-collecting bag is for single use only. Collect and destroy empty used bags through incineration in a controlled environment. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Ten (10) urine bags in a plastic bag Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of standards): ISO 8669-2:1996 Urine collection bags Requirements and test methods ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001, Alternate materials other than PVC is a preferred option. | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: designed to connect to an indwelling urinary catheter to collect discharged urine from a patient Product description: The 2-litre, urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. Material for bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA), flexible. Tube & connector/protective cap: polyvinyl chloride (PVC) Urine bag has two (2) strong eyelets from which to suspend the bag Bag collector fitted with a non retun valve. The non-return valve is located inside the urine collector bag at the upper part or urine entry point, and prevents urine backflow into the indwelling urinary catheter. Tube with universal connector and protective cap to connect to a catheter. Outlet tube fitted with a drain valve with seal cap allowing to maintain sterility. Tube length: approx. 90cm. Capacity: 2000ml. Graduation: 100ml. Sterile, Single use. Shelf life 5 yrs Instructions for use: Used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Safety process: The urine-collecting bag is for single use only. Collect and destroy empty used bags through incineration in a controlled environment. Supplied with: Manufacturer's instruction for use, indications for use, contraindications, warnings, precautions for use, storage conditions, disposal instructions etc shall be provided. Packaging and labelling: Ten (10) urine bags in a plastic bag Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of standards): ISO 8669-2:1996 Urine collection bags Requirements and test methods ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes Environmental requirements: ISO 14001, Alternate materials other than PVC is a preferred option.
315 | Medical supplies | URINEBAG2L.TAP | Bag, urine, collecting, 2 litres, tap for emptying | Bag, urine, collecting, 2 litres, tap for emptying | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Urine collection bag. Medical grade plastic bag, capacity 2000ml, with graduated scale every 100ml to allow proper reading of the liquid contained in the bag and reinforced eyelets for hanging. Kink-resistant and transparent plastic inlet tube, length +/- 90cm, with universal connector and protective cap. With outlet which permits bag to be emptied without disconnecting. Material: Bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA). Tubing & connector/protective cap: polyvinyl chloride (PVC) Single-use Non sterile Instructions for use: The 2-litre urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Five (5) or Ten (10) urine bags in a plastic bag Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) if applicable. Expiry date by year and month, prefixed by the word "EXP"(or equivalent harmonised symbol) if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 1616:1997: Sterile urethral catheters for single use ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. Environmental requirements: Latex free, it is recommended to use silicone | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Urine collection bag. Medical grade plastic bag, capacity 2000ml, with graduated scale every 100ml to allow proper reading of the liquid contained in the bag and reinforced eyelets for hanging. Kink-resistant and transparent plastic inlet tube, length +/- 90cm, with universal connector and protective cap. With outlet which permits bag to be emptied without disconnecting. Material: Bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA). Tubing & connector/protective cap: polyvinyl chloride (PVC) Single-use Non sterile Instructions for use: The 2-litre urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Five (5) or Ten (10) urine bags in a plastic bag Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) if applicable. Expiry date by year and month, prefixed by the word "EXP"(or equivalent harmonised symbol) if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: EN 1616:1997: Sterile urethral catheters for single use ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. Environmental requirements: Latex free, it is recommended to use silicone
316 | Medical supplies | UTERINESOUNDMART | Sound, uterine, Martin, 32cm | Sound, uterine, Martin, 32cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Uterine sound, Martin, for measuring, via the vagina and cervix , the depth of the uterine cavity. Material: Silver-coated copper with brass handle or silver coated austenitic steel ( non quenched, non magnetic) Shank type: Rigid. Blade shape: Oval form of a spoon. Blade edge type: Sharp Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium. Length: Shaft 320 mm. Width: Spoon 6 mm Supplied with: Manufacturer's instruction for use. Instructions for use: For removal of retained products of conception from inside the uterus This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark "IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Urine collection bag. Medical grade plastic bag, capacity 2000ml, with graduated scale every 100ml to allow proper reading of the liquid contained in the bag and reinforced eyelets for hanging. Kink-resistant and transparent plastic inlet tube, length +/- 90cm, with universal connector and protective cap. With outlet which permits bag to be emptied without disconnecting. Material: Bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA). Tubing & connector/protective cap: polyvinyl chloride (PVC) Single-use Non sterile Instructions for use: The 2-litre urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Five (5) or Ten (10) urine bags in a plastic bag Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word ""LOT"" (or equivalent harmonised symbol) if applicable. Expiry date by year and month, prefixed by the word ""EXP""(or equivalent harmonised symbol) if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Uterine sound, Martin, for measuring, via the vagina and cervix , the depth of the uterine cavity. Material: Silver-coated copper with brass handle or silver coated austenitic steel ( non quenched, non magnetic) Shank type: Rigid. Blade shape: Oval form of a spoon. Blade edge type: Sharp Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium. Length: Shaft 320 mm. Width: Spoon 6 mm Supplied with: Manufacturer's instruction for use. Instructions for use: For removal of retained products of conception from inside the uterus This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark "IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Urine collection bag. Medical grade plastic bag, capacity 2000ml, with graduated scale every 100ml to allow proper reading of the liquid contained in the bag and reinforced eyelets for hanging. Kink-resistant and transparent plastic inlet tube, length +/- 90cm, with universal connector and protective cap. With outlet which permits bag to be emptied without disconnecting. Material: Bag: polyvinyl chloride (PVC) or ethylene vinyl acetate (EVA). Tubing & connector/protective cap: polyvinyl chloride (PVC) Single-use Non sterile Instructions for use: The 2-litre urine collecting bag is used for collecting urine from a patient via a urinary (e.g. Foley) catheter, which makes the bag part of a closed system. The size has been chosen as being the most commonly used. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Five (5) or Ten (10) urine bags in a plastic bag Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word ""LOT"" (or equivalent harmonised symbol) if applicable. Expiry date by year and month, prefixed by the word ""EXP""(or equivalent harmonised symbol) if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device
317 | Medical utensils | BASINKIDNEY1400ML | Basin, kidney, stainless steel, 1400ml | Basin, kidney, stainless steel, 1400ml | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: Kidney dish, stainless steel, large, 32 cm Container, kidney shaped, kidney dish, stainless steel, smooth surface. Material: Austenitic stainless steel. Austenitic stainless steel composition: chromium 18-20%, molybdenum 2-3%, nickel 8-10%, silicon 0.5-1%, manganese 0.4-2% nickel. Length: approximately 32 cm Capacity: 1400ml (±5% tolerance limit) Thickness: 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Primary packaging: Unit of use One (1) kidney basin protected by adhesive plastic film or similar Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicatingvthe position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Chlorinated solutions should not be left in the dish/basin/pot. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: Kidney dish, stainless steel, large, 32 cm Container, kidney shaped, kidney dish, stainless steel, smooth surface. Material: Austenitic stainless steel. Austenitic stainless steel composition: chromium 18-20%, molybdenum 2-3%, nickel 8-10%, silicon 0.5-1%, manganese 0.4-2% nickel. Length: approximately 32 cm Capacity: 1400ml (±5% tolerance limit) Thickness: 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Primary packaging: Unit of use One (1) kidney basin protected by adhesive plastic film or similar Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicatingvthe position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Chlorinated solutions should not be left in the dish/basin/pot. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
318 | Medical utensils | BASINKIDNEY475ML | Basin, kidney, polypropylene, 475ml | Basin, kidney, polypropylene, 475ml | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: Kidney shaped container made of polypropylene used for general ward use. Material: Polypropylene, autoclavable Length: approx. 205mm Width: approx. 95mm Height: approx. 45mm Capacity: approx. 475ml (±5% tolerance limit) Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Over packaging: Packaging unit Labelling on the packaging unit: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: Kidney shaped container made of polypropylene used for general ward use. Material: Polypropylene, autoclavable Length: approx. 205mm Width: approx. 95mm Height: approx. 45mm Capacity: approx. 475ml (±5% tolerance limit) Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Over packaging: Packaging unit Labelling on the packaging unit: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
319 | Medical utensils | BASINKIDNEY825ML | Basin, kidney, stainless steel, 825ml | Basin, kidney, stainless steel, 825ml | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: For receipt of solid material (dressings, swabs, soiled tubes, etc); used in medical and surgical departments. Kidney shaped container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel Thickness: min. 0.8mm. Capacity: approx. 825ml. (±5% tolerance limit) Length: approx. 250mm. Width: approx. 140mm. Height: approx. 40mm. Thickness: 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) kidney basin protected by adhesive plastic film or similar Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Kidney basins in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: commonly used to hold and transport medical consumables or instruments for a procedure Product description: For receipt of solid material (dressings, swabs, soiled tubes, etc); used in medical and surgical departments. Kidney shaped container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel Thickness: min. 0.8mm. Capacity: approx. 825ml. (±5% tolerance limit) Length: approx. 250mm. Width: approx. 140mm. Height: approx. 40mm. Thickness: 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) kidney basin protected by adhesive plastic film or similar Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Kidney basins in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
320 | Medical utensils | BOWL180ML | Bowl, stainless steel, 180ml | Bowl, stainless steel, 180ml | Intended purpose of use: Also known as gallipot is a round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Receptacle for liquids and other materials, used before and during operations. Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) composition: 18 to 20% chromium and 8 to 10% nickel. Thickness: min. 0.8mm. Capacity: 150 to 180ml. Inner diameter: approx. 80mm. Height: approx. 40mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowls in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: Also known as gallipot is a round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Receptacle for liquids and other materials, used before and during operations. Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) composition: 18 to 20% chromium and 8 to 10% nickel. Thickness: min. 0.8mm. Capacity: 150 to 180ml. Inner diameter: approx. 80mm. Height: approx. 40mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowls in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
321 | Medical utensils | BOWL4L | Bowl, stainless steel, 4L | Bowl, stainless steel, 4L | Intended purpose of use: A round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel, composition: 18 to 20% chromium and 8 to 10% nickel. Capacity: 4000ml. Thickness: min. 0.8mm. Inner diameter: 300mm. Height: approximately 100mm. shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowl in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: A round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel, composition: 18 to 20% chromium and 8 to 10% nickel. Capacity: 4000ml. Thickness: min. 0.8mm. Inner diameter: 300mm. Height: approximately 100mm. shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowl in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
322 | Medical utensils | BOWL500ML | Bowl, stainless steel, 500 ml | Bowl, stainless steel, 500 ml | Intended purpose of use: Also known as gallipot is a round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Receptacle for liquids and other materials, used before and during operations. Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) Product Description: Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) Smooth walls without serrations. Thickness: min. 0.8mm. Capacity: approx. 500 to 600ml Inner diameter: approx. 120mm. Height: approx. 60mm. shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowl in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards ( as per latest revisions of product standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: Also known as gallipot is a round shaped container that is used to pour and hold antiseptic fluids or to mix them during the procedure or in nursing area. Product description: Receptacle for liquids and other materials, used before and during operations. Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) Product Description: Round container made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel (non-quenched, non-magnetic) Smooth walls without serrations. Thickness: min. 0.8mm. Capacity: approx. 500 to 600ml Inner diameter: approx. 120mm. Height: approx. 60mm. shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X Round bowl in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: to be provided in English/French/Arabic/Spanish. Basic equipment for nursing and surgical use. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards ( as per latest revisions of product standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
323 | Medical utensils | BOWL6L | Bowl, polypropylene, 6L | Bowl, polypropylene, 6L | Intended purpose of use: A round shaped container that is for general ward use/nursing area. Product Description: Receptacle for liquids and other materials, for general use. Round container made of polypropylene, smooth, non-glare surface. Material: Polypropylene, autoclavable. Inner diameter: approximately 345mm. Height: approximately 120mm. Capacity: approximately 6000 ml shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Over packaging: Packaging unit Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: For generalward use to be provided in English/French/Arabic/Spanish. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards ( as per latest revisions of product standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: A round shaped container that is for general ward use/nursing area. Product Description: Receptacle for liquids and other materials, for general use. Round container made of polypropylene, smooth, non-glare surface. Material: Polypropylene, autoclavable. Inner diameter: approximately 345mm. Height: approximately 120mm. Capacity: approximately 6000 ml shall provide the weight of the item (5% tolerance limit allowed) Packaging and labelling: Over packaging: Packaging unit Labelling on the primary packaging: Name and address of the manufacturer. Provide brand name/trade mark details if any. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Instructions for use: For generalward use to be provided in English/French/Arabic/Spanish. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance Classification: Classified under EU MDR 2017/745 as Class I device
324 | Medical utensils | JARFORCEPS | Jar, forceps, 180 mm, no cover | Jar, forceps, 180 mm, no cover | Intended purpose of use: For holding various types of forceps prior to and during a surgical intervention or an examination procedure Product description: Plastic jar for containing forceps, smooth surface without cover Material: Polypropylene Autoclavable at 121 °C Without cover. Internal diameter: approx. 54mm. Depth: approx. 180 mm. Overall height: approx. 200mm. Shall provide the approximate weight of the item (±5% tolerance limit) Instructions for use: to be provided in English/French/Arabic/Spanish. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: plastic film, non-sterile. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards):Must comply with following standards Safety & product Standards: ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices | Intended purpose of use: For holding various types of forceps prior to and during a surgical intervention or an examination procedure Product description: Plastic jar for containing forceps, smooth surface without cover Material: Polypropylene Autoclavable at 121 °C Without cover. Internal diameter: approx. 54mm. Depth: approx. 180 mm. Overall height: approx. 200mm. Shall provide the approximate weight of the item (±5% tolerance limit) Instructions for use: to be provided in English/French/Arabic/Spanish. cleaning, disinfecting and sterilization instructions, list of non comptable disinfectants, if any Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: plastic film, non-sterile. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per latest revisions of the product standards):Must comply with following standards Safety & product Standards: ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices
325 | Medical utensils | JARTHERMOMETER | Jar, thermometer, plastic, 11cm, w/o cover | Jar, thermometer, plastic, 11cm, w/o cover | Intended purpose of use : For holding the thermometer from falling out and breaking. Product description: Plastic jar for containing thermometer, smooth surface without cover. Round container with base. Without cover. Material: Polypropylene Autoclavable Internal diameter: approx. 30mm. Depth: approx. 110mm. Overall height: approx. 120mm. Shall provide the weight of the item Instructions for use: Basic item of medical equipment. Polypropylene jar for holding a thermometer. The jar is provided with a lid to prevent the thermometer from falling out and breaking. The size has been chosen as being the most commonly used. Safety process: This item is used as a "clean" container for "clean" thermometers only. This item must be cleaned, disinfected and sterilized in a steam steriliser as often as necessary Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: plastic film, non-sterile. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer | Intended purpose of use : For holding the thermometer from falling out and breaking. Product description: Plastic jar for containing thermometer, smooth surface without cover. Round container with base. Without cover. Material: Polypropylene Autoclavable Internal diameter: approx. 30mm. Depth: approx. 110mm. Overall height: approx. 120mm. Shall provide the weight of the item Instructions for use: Basic item of medical equipment. Polypropylene jar for holding a thermometer. The jar is provided with a lid to prevent the thermometer from falling out and breaking. The size has been chosen as being the most commonly used. Safety process: This item is used as a "clean" container for "clean" thermometers only. This item must be cleaned, disinfected and sterilized in a steam steriliser as often as necessary Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: plastic film, non-sterile. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): ISO 13485: 2003 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer
326 | Medical utensils | SPONGEBOWL1 | Bowl, sponge, stainless steel, 1.2 L | Bowl, sponge, stainless steel, 1.2 L | Intended purpose of use: Receptacle for liquids and other materials, used before and during operations. Product description: Bowl, round, smooth, non-glare surface. Material: Austenitic stainless steel Inner diameter: approx. 160 mm Height: approx. 75 mm Thickness: min. 0.8 mm. Capacity: approx. 1200 ml. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film or equivalent Symbols used according ISO 15223 Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards):Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: Receptacle for liquids and other materials, used before and during operations. Product description: Bowl, round, smooth, non-glare surface. Material: Austenitic stainless steel Inner diameter: approx. 160 mm Height: approx. 75 mm Thickness: min. 0.8 mm. Capacity: approx. 1200 ml. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel bowl protected by adhesive plastic film or equivalent Symbols used according ISO 15223 Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards):Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
327 | Medical utensils | RULER30CM | Ruler, metal, in centimeters, 30cm | Ruler, metal, in centimeters, 30cm | Product description: 30cm long ruler, Graduated in cm with submarking in mm and inches Length: 30 cm Material: austenitic steel | Product description: 30cm long ruler, Graduated in cm with submarking in mm and inches Length: 30 cm Material: austenitic steel
328 | Medical utensils | TRAYDRESSING300 | Tray, dressing, stainless steel, 30 x 20 x 3cm | Tray, dressing, stainless steel, 30 x 20 x 3cm | Intended purpose of use:used for carrying miscellaneous objects: surgical instruments, dressing materials, injection materials, etc. Product description: Seamless tray, rectangular with rounded corners, stainless steel, non-glare surface Rectangular tray made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel ; composition: 18 to 20% chromium, 8 to 10% nickel. Length: approx.300mm. Width: approx.200mm. Height: approx. 30mm. Thickness: min. 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel dressing tray protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X dressing tray in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use:used for carrying miscellaneous objects: surgical instruments, dressing materials, injection materials, etc. Product description: Seamless tray, rectangular with rounded corners, stainless steel, non-glare surface Rectangular tray made of stainless steel, smooth, non-glare surface. Material: Austenitic stainless steel ; composition: 18 to 20% chromium, 8 to 10% nickel. Length: approx.300mm. Width: approx.200mm. Height: approx. 30mm. Thickness: min. 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) stainless steel dressing tray protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name and address of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X dressing tray in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
329 | Medical utensils | TRAYINSTRUMENT225 | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm, with cover | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm, with cover | Intended purpose of use: Basic equipment for nursing and surgical care Product description: Seamless tray with cover, rectangular with rounded corners, stainless steel, smooth surface. Material: Austenitic stainless steel; composition: 18 to 20% chromium, 8 to 10% nickel With cover Length: approximately 225mm. Width: approximately 125mm. Height: approximately 50mm. Thickness: min. 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) instruments tray with cover protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name andaddress of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X instruments trays with cover in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark. Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: Basic equipment for nursing and surgical care Product description: Seamless tray with cover, rectangular with rounded corners, stainless steel, smooth surface. Material: Austenitic stainless steel; composition: 18 to 20% chromium, 8 to 10% nickel With cover Length: approximately 225mm. Width: approximately 125mm. Height: approximately 50mm. Thickness: min. 0.8mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) instruments tray with cover protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name andaddress of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X instruments trays with cover in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Symbols used according ISO 15223 CE Mark. Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
330 | Medical utensils | TRAYINSTRUMENT320 | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover | Intended purpose of use: Basic equipment for nursing and surgical care Product description: Seamless tray with cover, rectangular with rounded corners, stainless steel, smooth surface. Material: Austenitic stainless steel; composition: 18 to 20% chromium, 8 to 10% nickel Length: approximately 320mm. Width: approximately 200mm. Height: approximately 80mm. Height: 80mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) instruments tray with cover protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name andaddress of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X instruments trays with cover in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box Symbols used according ISO 15223 CE Mark. Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel | Intended purpose of use: Basic equipment for nursing and surgical care Product description: Seamless tray with cover, rectangular with rounded corners, stainless steel, smooth surface. Material: Austenitic stainless steel; composition: 18 to 20% chromium, 8 to 10% nickel Length: approximately 320mm. Width: approximately 200mm. Height: approximately 80mm. Height: 80mm. Reusable, Autoclavable 121Deg C Shall be supplied with cleaning, disinfecting and sterilization instructions. Shall provide the approximate weight of the item (±5% tolerance limit) Packaging and labelling: Primary packaging: Unit of use One (1) instruments tray with cover protected by adhesive plastic film or equivalent Labelling on the primary packaging: Name andaddress of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit X instruments trays with cover in a box Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box Symbols used according ISO 15223 CE Mark. Instructions for use: Basic equipment for nursing and surgical care. The size has been chosen as being the most commonly used. Note: Do not use a stainless steel container to hold a chlorine solution, as this may damage the stainless steel; use a plastic container. This item must be cleaned, disinfected after each use. It can be sterilized in a steam sterilizer when necessary. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (submissions shall be as per latest revisions of the standards): Must comply with following standards ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel
331 | Medical Attire & Linen | APRONREUSABLE | Apron, protection, plastic, reusable | Apron, protection, plastic, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Apron to be used in healthcare facilities by personnel performing medical/obsterical/surigcal procedures with high risk of contamination by body fluids projection. Product description: Protective apron to be worn by surgeon or midwife for protection. Straight apron with bib, back fastening and neckband, should be strong and not detachable. Moisture-proof and blood, water, chemical and stain resistant. Material: Opaque or translucent high quality plastic material withstanding washing, boiling, steam sterilization, resistant to 0.5% chlorine. Size: Adult size. Length: 120cm (+/- 5cm from top of the bib to lower edge of the apron). Total length: 145 cm (+/-5cm) Width: 80cm (+/-10cm). Thickness: 0.15 - 0.30mm. Non-sterile and reusable. Instructions for use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. The size has been chosen as the most commonly used. After use, the apron must be washed, disinfected with chlorine solution (0.05% concentration) and dried. If the apron is damaged (perforation, etc.), it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packaged in a transparent plastic bag. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device The words “Non-sterile, reusable" (or equivalent harmonised symbol). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 22610 Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment - Test method to determine the resistance to wet bacterial penetration ISO 22612 Clothing for protection against infectious agents - Test method for resistance to dry microbial penetration Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Apron to be used in healthcare facilities by personnel performing medical/obsterical/surigcal procedures with high risk of contamination by body fluids projection. Product description: Protective apron to be worn by surgeon or midwife for protection. Straight apron with bib, back fastening and neckband, should be strong and not detachable. Moisture-proof and blood, water, chemical and stain resistant. Material: Opaque or translucent high quality plastic material withstanding washing, boiling, steam sterilization, resistant to 0.5% chlorine. Size: Adult size. Length: 120cm (+/- 5cm from top of the bib to lower edge of the apron). Total length: 145 cm (+/-5cm) Width: 80cm (+/-10cm). Thickness: 0.15 - 0.30mm. Non-sterile and reusable. Instructions for use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. The size has been chosen as the most commonly used. After use, the apron must be washed, disinfected with chlorine solution (0.05% concentration) and dried. If the apron is damaged (perforation, etc.), it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packaged in a transparent plastic bag. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR certified), use-by-date, symbol for Medical device The words “Non-sterile, reusable" (or equivalent harmonised symbol). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 22610 Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment - Test method to determine the resistance to wet bacterial penetration ISO 22612 Clothing for protection against infectious agents - Test method for resistance to dry microbial penetration Environmental requirements: ISO 14001
332 | Medical Attire & Linen | APRONSURGICALRB | Apron, surgical, rubber, washable | Apron, surgical, rubber, washable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. Product description: Protective apron used for protecting gown, clothing or work wear when carrying out certain medical surgical procedures, non-sterile and multiple uses. Material: 100% nitril rubber, autoclavable. Moisture-proof and resistant to fats, acids, stains and heat. Straight apron with bib. Back fastening and neckband should be strong and not detachable Colour: white. Size: One-size-fit-all (120 x 89 cm) Thickness: 0.15 to 0.30mm Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Instructions for use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. The size has been chosen as the most commonly used. After use, the apron must be washed, disinfected with chlorine solution (0.05% concentration) and dried. If the apron is damaged (perforation, etc.), it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packaged in a transparent plastic bag. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 22610 Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment - Test method to determine the resistance to wet bacterial penetration ISO 22612 Clothing for protection against infectious agents - Test method for resistance to dry microbial penetration. Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. Product description: Protective apron used for protecting gown, clothing or work wear when carrying out certain medical surgical procedures, non-sterile and multiple uses. Material: 100% nitril rubber, autoclavable. Moisture-proof and resistant to fats, acids, stains and heat. Straight apron with bib. Back fastening and neckband should be strong and not detachable Colour: white. Size: One-size-fit-all (120 x 89 cm) Thickness: 0.15 to 0.30mm Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Instructions for use: Apron to be used in healthcare facilities by personnel performing medical / obstetrical / surgical procedures with high risk of contamination by body fluids projection. The size has been chosen as the most commonly used. After use, the apron must be washed, disinfected with chlorine solution (0.05% concentration) and dried. If the apron is damaged (perforation, etc.), it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Individually packaged in a transparent plastic bag. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 22610 Surgical drapes, gowns and clean air suits, used as medical devices, for patients, clinical staff and equipment - Test method to determine the resistance to wet bacterial penetration ISO 22612 Clothing for protection against infectious agents - Test method for resistance to dry microbial penetration. Environmental requirements: ISO 14001
333 | Medical Attire & Linen | BEDSHEET | Sheet, bed, 289 x 178 cm, washable | Sheet, bed, 289 x 178 cm, washable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Hemmed rectangular textile used for making up a hospital bed. Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g Size: 289 x 178 cm for a mattress hospital bed of 200 x 85 x 12 cm.(l x w x h) 4 hemmed borders. Colour: white. Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Packaging and labelling: Unit presentation: 1 sheet, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Hemmed rectangular textile used for making up a hospital bed. Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g Size: 289 x 178 cm for a mattress hospital bed of 200 x 85 x 12 cm.(l x w x h) 4 hemmed borders. Colour: white. Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Packaging and labelling: Unit presentation: 1 sheet, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc
334 | Medical Attire & Linen | BLANKETSURVIVAL | Blanket, survival, 220 x 140cm | Blanket, survival, 220 x 140cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Wraparound body blanket protects against cold, heat, wind and rain Product description: Low-weight blanket made of heat-reflective thin plastic sheeting, intended to reduce the heat loss in a person's body due to thermal radiation, water evaporation, or convection. Used to protect the body against cold, heat, wind and rain. Thin sheet of plastic , one side coated with a metallic reflecting agent (silver or gold), reflects up to 90% of radiated heat Lightweight flexible wrap-around blanket Easy to unfold, strong, tear resistant One side silver coated, reflects up to 90% of radiated heat Wind and waterproof, and rot proof Stays flexible in freezing temperatures Single use, non-sterile Materials: Non-stretch polyester film, or polyethylene terephthalate (PET), Rectangular size, approx.: 220 x 140 cm Dimensions: Length, approximately: 220cm Width, approximately: 140cm Thickness, approximately: 12µm Weight: approx. 0.075kg Volume: approx. 5cdm Instructions for use: Wraparound body blanket protects against cold, heat, wind and rain Arrange the silver coated surface according the desired effect: - External - reflects heat away, isolates content against heat - Internal - preserves (body) heat, isolates from wind and humidity If necessary the blanket can be cut into smaller pieces for children and new-borns. Safety process: In case of absolute necessity, the blanket may be reused after cleaning and disinfecting with chlorine solution. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) survival blanket in a plastic bag. Labelling on the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device The words “Non-sterile, reusable" (or equivalent harmonised symbol). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.Name and/or trademark and address of the manufacturer. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 5085-1Textiles -- Determination of thermal resistance -- Part 1: Low thermal resistance ISO 9237 Textiles -- Determination of the permeability of fabrics to air ISO 12952-1&2&3&4 Fire resistant ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Wraparound body blanket protects against cold, heat, wind and rain Product description: Low-weight blanket made of heat-reflective thin plastic sheeting, intended to reduce the heat loss in a person's body due to thermal radiation, water evaporation, or convection. Used to protect the body against cold, heat, wind and rain. Thin sheet of plastic , one side coated with a metallic reflecting agent (silver or gold), reflects up to 90% of radiated heat Lightweight flexible wrap-around blanket Easy to unfold, strong, tear resistant One side silver coated, reflects up to 90% of radiated heat Wind and waterproof, and rot proof Stays flexible in freezing temperatures Single use, non-sterile Materials: Non-stretch polyester film, or polyethylene terephthalate (PET), Rectangular size, approx.: 220 x 140 cm Dimensions: Length, approximately: 220cm Width, approximately: 140cm Thickness, approximately: 12µm Weight: approx. 0.075kg Volume: approx. 5cdm Instructions for use: Wraparound body blanket protects against cold, heat, wind and rain Arrange the silver coated surface according the desired effect: - External - reflects heat away, isolates content against heat - Internal - preserves (body) heat, isolates from wind and humidity If necessary the blanket can be cut into smaller pieces for children and new-borns. Safety process: In case of absolute necessity, the blanket may be reused after cleaning and disinfecting with chlorine solution. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) survival blanket in a plastic bag. Labelling on the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark or similar, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device The words “Non-sterile, reusable" (or equivalent harmonised symbol). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.Name and/or trademark and address of the manufacturer. Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 5085-1Textiles -- Determination of thermal resistance -- Part 1: Low thermal resistance ISO 9237 Textiles -- Determination of the permeability of fabrics to air ISO 12952-1&2&3&4 Fire resistant ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001
335 | Medical Attire & Linen | COTTONCLOTH | Cotton cloth, “tetra”, 100 x 100cm | Cotton cloth, “tetra”, 100 x 100cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Cotton towel, 100 x 100 cm size Material: Hydrophilic “tetra”cotton. Density approx 90 g/m2 Washable in chlorine Size: 100 x 100cm. Non-sterile. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packed in polythene bags with the following options: Option 1 - Individual towel packed in a bag with each towel (folded 4 times) and placed in an export carton (288 bags per carton) Option 2 - 6 towel packed in a bag with each towel folded (4 times) placed in an export carton (48 bags per carton) Option 3 - 288 towel packed in a bag with each towel (folded 4 times) placed in an export carton (1 bag per carton). Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc Must comply with following standards | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Cotton towel, 100 x 100 cm size Material: Hydrophilic “tetra”cotton. Density approx 90 g/m2 Washable in chlorine Size: 100 x 100cm. Non-sterile. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packed in polythene bags with the following options: Option 1 - Individual towel packed in a bag with each towel (folded 4 times) and placed in an export carton (288 bags per carton) Option 2 - 6 towel packed in a bag with each towel folded (4 times) placed in an export carton (48 bags per carton) Option 3 - 288 towel packed in a bag with each towel (folded 4 times) placed in an export carton (1 bag per carton). Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc Must comply with following standards
336 | Medical Attire & Linen | DRAPESURG100X100 | Drape, surgical, woven, 100x100 cm, reusable | Drape, surgical, woven, 100x100 cm, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Specific drape, square of material used to delimit the zone for medical/surgical procedures on the patient. Material: 100% cotton or 50% polyester - 50% cotton. Reusable woven fabrics: Fabric: 100% cotton - Weight per m²: 196 g/m² - Number of threads per 3 cm: Warp: 75; Weft: 75 - Metric count: Warp: 28 / Weft: 27 or Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g or Fabric : 50%polyesther-50% cotton fabric - Weight per m²: +/- 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 Size: 100 x 100 cm. Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations Wahing: normal; withstands boiling and autoclaving; resists to chlorine 0.1% Multiple use Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. All new drapes must be washed prior to sterilization. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 drape, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc Must comply with following standards BS EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels. ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Specific drape, square of material used to delimit the zone for medical/surgical procedures on the patient. Material: 100% cotton or 50% polyester - 50% cotton. Reusable woven fabrics: Fabric: 100% cotton - Weight per m²: 196 g/m² - Number of threads per 3 cm: Warp: 75; Weft: 75 - Metric count: Warp: 28 / Weft: 27 or Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g or Fabric : 50%polyesther-50% cotton fabric - Weight per m²: +/- 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 Size: 100 x 100 cm. Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations Wahing: normal; withstands boiling and autoclaving; resists to chlorine 0.1% Multiple use Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. All new drapes must be washed prior to sterilization. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 drape, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc Must comply with following standards BS EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels. ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc
337 | Medical Attire & Linen | DRAPESURG100X150 | Drape, surgical woven, 100 x 150cm | Drape, surgical woven, 100 x 150cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Material: 100% cotton or 50% polyester - 50% cotton. Fabric : 50%polyesther-50% cotton fabric - Weight per m²: approx. 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 or Fabric: cretonne, cotton 100%: - Weight per m²: approx. 180g - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 or Fabric: 100% cotton or 50/50 cotton polyester blend weight: 196g/m2 number of threads: warp 75, weft 75 75 (per 3cm) metric count: warf 28, weft 27 (per 3cm) Size: approx.100 x 150 cm Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations. All fabrics used must be suitable for normal washing - boiling and autoclaving- and be resistant to a 0.1% chlorine solution. Non-sterile; reusable. washable Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. All new drapes must be washed prior to sterilization. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: Unit presentation: 1 drape, non-sterile, in a plastic bag with instructions for use Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI, use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Material: 100% cotton or 50% polyester - 50% cotton. Fabric : 50%polyesther-50% cotton fabric - Weight per m²: approx. 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 or Fabric: cretonne, cotton 100%: - Weight per m²: approx. 180g - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 or Fabric: 100% cotton or 50/50 cotton polyester blend weight: 196g/m2 number of threads: warp 75, weft 75 75 (per 3cm) metric count: warf 28, weft 27 (per 3cm) Size: approx.100 x 150 cm Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations. All fabrics used must be suitable for normal washing - boiling and autoclaving- and be resistant to a 0.1% chlorine solution. Non-sterile; reusable. washable Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. All new drapes must be washed prior to sterilization. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: Unit presentation: 1 drape, non-sterile, in a plastic bag with instructions for use Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI, use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc
338 | Medical Attire & Linen | DRAPESURGWOVEN | Drape, surgical, woven, 90x150 cm, reusable | Drape, surgical, woven, 90x150 cm, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Material: 100% cotton or 50% polyester - 50% cotton. Reusable woven fabrics: Fabric: 100% cotton - Weight per m²: 196 g/m² - Number of threads per 3 cm: Warp: 75; Weft: 75 - Metric count: Warp: 28 / Weft: 27 or Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g or Fabric : 50%polyesther-50% cotton fabric - Weight per m²: +/- 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 Size: approx. 90 x 150 cm. Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations. All fabrics used must be suitable for normal washing - boiling and autoclaving- and be resistant to a 0.1% chlorine solution. Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 drape, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes BS EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Product description: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Material: 100% cotton or 50% polyester - 50% cotton. Reusable woven fabrics: Fabric: 100% cotton - Weight per m²: 196 g/m² - Number of threads per 3 cm: Warp: 75; Weft: 75 - Metric count: Warp: 28 / Weft: 27 or Fabric: cretonne, cotton 100%: - Number of threads/cm: warp: 24, weft: 24 - Metric count: warp: 28, weft: 28 - Weight per m²: 180g or Fabric : 50%polyesther-50% cotton fabric - Weight per m²: +/- 175gr/m² - Number of threads: Warp: 24; Weft 22 - Metric count: Warp: 28; Weft 28 Size: approx. 90 x 150 cm. Colour: blue or green, but not khaki-green. Colour quality must be able to withstand repeated washings and sterilizations. All fabrics used must be suitable for normal washing - boiling and autoclaving- and be resistant to a 0.1% chlorine solution. Instructions for use: To cover patients on the operating table and to maintain aseptic conditions in the operation area. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 drape, non-sterile, multiple-use, size and manufacture's name. Packaging: Protected in a plastic bag, single or multiple units. Labelling of the primary packaging displays, at least: Product name, size, reference number, manufacturer's name and address, CE mark, lot number or batch number, UDI (if MDR followed), use-by-date, symbol for Medical device. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Instruction for sterilization Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards: ISO 13485 Medical devices -- Quality management systems -- Requirements for regulatory purposes BS EN 13795 Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment. General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO13934-1 and ISO 6330 Tensile strength loss after washing Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc
339 | Medical Attire & Linen | DRAWSHEET | Draw sheet, waterproof, 80 x 120cm, single use | Draw sheet, waterproof, 80 x 120cm, single use | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Used under the patient to cover the operating table to protect against soaking. Used to cover beds of patients with extensive wounds to protect against soaking. Product description: Rectangular plastic sheet, used to protect equipment from soiling. Material: Non-woven fabric. 1st layer polyethylene, shiny side. 2nd layer non-woven absorbent material, viscose type. The non-woven fabric must be reinforced lengthwise with blue polyether fibres. - Absorbent and impervious protection. - Good resistance to tears. - Individually folded for swift fitting. - Non-allergic. - Hygienically clean for hospital use. Size: approx. 80 x 120cm. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. After the first use it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: 1 single-use draw sheet, folded. Secondary packaging: 75 units per carton. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Used under the patient to cover the operating table to protect against soaking. Used to cover beds of patients with extensive wounds to protect against soaking. Product description: Rectangular plastic sheet, used to protect equipment from soiling. Material: Non-woven fabric. 1st layer polyethylene, shiny side. 2nd layer non-woven absorbent material, viscose type. The non-woven fabric must be reinforced lengthwise with blue polyether fibres. - Absorbent and impervious protection. - Good resistance to tears. - Individually folded for swift fitting. - Non-allergic. - Hygienically clean for hospital use. Size: approx. 80 x 120cm. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. After the first use it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: 1 single-use draw sheet, folded. Secondary packaging: 75 units per carton. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 Environmental requirements: ISO 14001 free from heavy metals, substances of very high concern (SVHC), allergy-causing chemicals, azo dyes, formaldehyde etc
340 | Medical Attire & Linen | DRAWSHEETPL | Draw sheet, plastic, 90 x 180cm, reusable | Draw sheet, plastic, 90 x 180cm, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Rectangular plastic sheet, used to protect equipment from soiling. Material: Opaque or translucent plastic (PVC) Width: 90-110cm Length: 150-180cm Thickness: 10-20 microns Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Reusable and non-sterile. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. The draw sheet must be washed, can be disinfected with chlorine solution (0.05% concentration) and dried after use. If the draw sheet is damaged (perforation, etc.) it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) plastic drawsheet in a plastic bag. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Rectangular plastic sheet, used to protect equipment from soiling. Material: Opaque or translucent plastic (PVC) Width: 90-110cm Length: 150-180cm Thickness: 100-150 microns Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Reusable and non-sterile. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. The draw sheet must be washed, can be disinfected with chlorine solution (0.05% concentration) and dried after use. If the draw sheet is damaged (perforation, etc.) it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) plastic drawsheet in a plastic bag. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001
341 | Medical Attire & Linen | DRAWSHEETPL2 | Draw sheet, plastic, 100 x 100 cm | Draw sheet, plastic, 100 x 100 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Product description: Rectangular plastic sheet, used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Material: Opaque or translucent plastic (PVC) Size: approx. 100 x 100cm. Thickness: 100-200 microns.c Waterproof. Single use and non-sterile. Hygienically clean for hospital use. Instructions for use: Used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Material: Opaque or translucent plastic (PVC) After the first use it must be disposed in waste container, collected and destroyed. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: One (1) plastic drawsheet in a plastic bag. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Product description: Rectangular plastic sheet, used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Material: Opaque or translucent plastic (PVC) Size: approx. 100 x 100cm. Thickness: 100-200 microns.c Waterproof. Single use and non-sterile. Hygienically clean for hospital use. Instructions for use: Used to protect equipment from liquid soiling, non-sterile and multiple use. Also, is used tod to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Material: Opaque or translucent plastic (PVC) After the first use it must be disposed in waste container, collected and destroyed. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Packaging: One (1) plastic drawsheet in a plastic bag. Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001
342 | Medical Attire & Linen | DRAWSHEETRB | Draw sheet, rubber, 90x150 cm, reusable | Draw sheet, rubber, 90x150 cm, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Product description: Rectangular rubber sheet, used to protect operating tables and patient beds. Material: Natural or synthetic rubber. Size: length 150cm, width between 90 to 100 cm. Waterproof, resistant to urine and acids. Colour: white cream. Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. The draw sheet must be washed, can be disinfected with chlorine solution (0.05% concentration) and dried after use. If the draw sheet is damaged (perforation, etc.) it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 draw sheet, hygienically clean, for medical use, non-sterile, multiple uses. Packaging: Can be packaged as one unit, or Per roll: 90cmx25m (cut to size, according to requirements). Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE mark or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. Product description: Rectangular rubber sheet, used to protect operating tables and patient beds. Material: Natural or synthetic rubber. Size: length 150cm, width between 90 to 100 cm. Waterproof, resistant to urine and acids. Colour: white cream. Washing: normal, withstands boiling and steam sterilization, resists to 0.5% chlorine. Instructions for use: The draw sheet is used to protect the examination, delivery, operating tables and/or bed mattresses from soiling. The draw sheet must be washed, can be disinfected with chlorine solution (0.05% concentration) and dried after use. If the draw sheet is damaged (perforation, etc.) it must be disposed in waste container, collected and destroyed. Please refer to WHO publication "Safe management of waste for Health Care". Supplied with: Manufacturer's instruction for use. Packaging and labelling: Unit presentation: 1 draw sheet, hygienically clean, for medical use, non-sterile, multiple uses. Packaging: Can be packaged as one unit, or Per roll: 90cmx25m (cut to size, according to requirements). Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE mark or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485 Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO 14001
343 | Medical Attire & Linen | SAFETY_GLASSES | Glasses, safety, regular size, disposable | Glasses, safety, regular size, disposable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Product description: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Wraparound protective eyewear with wide field view, fits over prescription eyewear Lightweight and comfortable. Easy to combine with the wearing of protective masks (surgical/respiratory masks) and over eyeglasses. Materials: Clear Polycarbonate. latex free 99.9% UV protection. Distortion-free Anti-fog Abrasion resistant Adapted to the shape of the face.Iintegral side-shields and contoured nose bridge. Flat side arms offering good sideways protection, temple length preferably adjustable. Weight: approx. 40-60 g Standard size Disposable Instructions for use: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Disposable glasses: to be destroyed if they are soiled or damaged, otherwise they could be reused after cleaning with water and soap. Never use solvent. Supplied with: Manufacturer's instruction for use ( English, French and Spanish) Packaging and labelling: Primary packaging: unit of use One (1) Pair of protective glasses in a sealed plastic bag with instructions for use Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows) or equivalent harmonised symbol. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Secondary packaging: Protected unit. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices EN 166 Personal eye-protection - Specification (optic class 1) EN 167 Personal eye-protection - Optical test methods EN 170 Personal eye-protection - Ultraviolet filters - Transmittance requirements and recommended use Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Product description: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Wraparound protective eyewear with wide field view, fits over prescription eyewear Lightweight and comfortable. Easy to combine with the wearing of protective masks (surgical/respiratory masks) and over eyeglasses. Materials: Clear Polycarbonate. latex free 99.9% UV protection. Distortion-free Anti-fog Abrasion resistant Adapted to the shape of the face.Iintegral side-shields and contoured nose bridge. Flat side arms offering good sideways protection, temple length preferably adjustable. Weight: approx. 40-60 g Standard size Disposable Instructions for use: Security glasses to protect against blood and exposure to other human fluids during surgery, midwifery and other medical procedures. Disposable glasses: to be destroyed if they are soiled or damaged, otherwise they could be reused after cleaning with water and soap. Never use solvent. Supplied with: Manufacturer's instruction for use ( English, French and Spanish) Packaging and labelling: Primary packaging: unit of use One (1) Pair of protective glasses in a sealed plastic bag with instructions for use Labelling: - Product name, size, reference number. - Manufacturer's name and address. - CE marked or similar - LOT or batch number - UDI (if MDR followed) - use-by-date - symbol for medical device - Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "destroy after use" (if space allows) or equivalent harmonised symbol. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable. Symbols used according ISO 15223 Secondary packaging: Protected unit. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices EN 166 Personal eye-protection - Specification (optic class 1) EN 167 Personal eye-protection - Optical test methods EN 170 Personal eye-protection - Ultraviolet filters - Transmittance requirements and recommended use Environmental requirements: ISO 14001
344 | In Vitro Diagnostics | BLOODGROUPINGTILE | Tile, porcelain, with depressions for blood grouping | Tile, porcelain, with depressions for blood grouping | Product name BLOOD GROUPING TILE Type of Immunoassay/ Intended Used: Blood type identification Description: White acrylic or ceramic tile or plate for multiple blood samples, include columns for Anti-A, Anti-B, Anti-AB, Auto Control, A Cells, B Cells, O Cells, Resus Grouping for Anti-D, and Resus Grouping for Control. Features: Glazed surface. Acid and alkali resistant Easy to clean Autoclavable with 12 +- well depressions. Reusable Packaging/labelling Primary packaging: Unit presentation: 1(one) Plate, cross matching/grouping, sets. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Batch number, Information for storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol) and Information for handling. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note: For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Classification: Classified under EU MDR 2017/745 as Class II device Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance ISO 14001 Material of batteries environmentally friendly Not use of PVC | Product name BLOOD GROUPING TILE Type of Immunoassay/ Intended Used: Blood type identification Description: White acrylic or ceramic tile or plate for multiple blood samples, include columns for Anti-A, Anti-B, Anti-AB, Auto Control, A Cells, B Cells, O Cells, Resus Grouping for Anti-D, and Resus Grouping for Control. Features: Glazed surface. Acid and alkali resistant Easy to clean Autoclavable with 12 +- well depressions. Reusable Packaging/labelling Primary packaging: Unit presentation: 1(one) Plate, cross matching/grouping, sets. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Batch number, Information for storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol) and Information for handling. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note: For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Classification: Classified under EU MDR 2017/745 as IVDD Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use Environmental requirements: Rohs compliance ISO 14001 Material of batteries environmentally friendly Not use of PVC
345 | Med. Diagnostic Equip&Supplies | New | Dopplerfetalheart | Dopplerfetalheart | GENERAL DESCRIPTION Doppler, fetal heart rate detector, monitor, w/accessories Handheld doppler based foetal heart rate detector with amplifier loudspeaker. INTENDED USE Ultrasonic fetal heart detectors are low-cost devices used in a variety of healthcare settings to provide audible and visual information about the fetus. Certain conditions, such as vaginal bleeding or lack of fetal movement, can cause anxiety in the mother, making her question the health of her unborn baby. The primary purpose of the fetal heart detector is to provide quick reassurance of fetal well-being to both the mother and the healthcare worker. The fetal heartbeat cannot be heard with an obstetric stethoscope until 24 weeks after conception. Ultrasonic fetal heart detectors can easily detect fetal heart motion throughout the pregnancy, starting as early as 8 weeks. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Doppler, foetal heart rate (FHR) detector, with accessories To be used in basic health infrastructures for routine examination of foetal life, from about 10-12 weeks gestation through to delivery. Self-test is performed each time the device is switchedCapable of detecting foetal heart rates (FHRs) in the range of 50 - 210 bpm, with 1 bpm resolution and 2 bpm accuracy. System reports, with audio-visual alert: operational status, malfunctions and low battery. Advanced noise/disturbance suppression system assures quality diagnostic sound Probe (transducer) with 2 MHz frequency with a probe attached via a cable. Regular size, approx: 18 to 20 cm. The cable can be extended to a 25 cm length when stretched. Probe detector head diameter at least 39 mm. Handheld, weights less than 0.45 kg with probe and batteries. Built-in speaker with volume adjustment. LCD screen displays FHR, pulse indicator and battery status Visual alerts for status, system error and low battery Operates on 3 x AA batteries. Battery life is sufficient for 10 hours continuous use. SUPPLIED WITH 1 x Soft carry bag easy to clean 1 x Instructions manual (instructions for assembly, use and maintenance in English, French, and Spanish) OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - The probe is interchangeable and probes with other frequencies are available: 4Mhz, 5Mhz and 8Mhz." INSTRUCTION FOR USE To be used in basic health infrastructures for routine examination of fetal life, from about 10-12 weeks gestation through to delivery. Doppler fetal heart rate detector should be operated by an adequately trained person only. Clean and disinfect the device after each use. Doppler fetal heart detector must be used and maintained according to manufacturer's instructions" ACCESSORIES/ SPARE PARTS / CONSUMABLES Ultrasound gel, batteries PACKAGING & LABELLING Unit presentation: 1 (one) Doppler, FHR detector, with accessories and Instructions for use. Symbols used according ISO 15223 CE with notified body number REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-1-6:2010 Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - IEC 60601-2-37:2007+AMD1:2015 CSV Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. - IEC 61157:2007+AMD1:2013 CSV Standard means for the reporting of the acoustic output of medical diagnostic ultrasonic equipment. - IEC 61266:1994 Ultrasonics - Hand-held probe Doppler foetal heartbeat detectors - Performance requirements and methods of measurement and reporting. - IEC 62366 : 2007" "ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 40°C / 25% – 80% RH. - Storage conditions: -20°C - 55°C / 25% – 93% RH. - Atmospheric pressure: 860 ~ 1060 hPa. - Ingress protection rating: Probe only IPX8. ENVIRONMENTAL REQUIREMENT Rohs compliant ESTIMATED DELIVERY LEAD TIME 120 days. INSTALLATION REQUIREMENTS There is no assembly or commisioning required for this device, other than instaling the batteries. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. COMPONENT OF A KIT No part of a kit. WEIGHT AND VOLUME (packaged) - estimated ESTIMATED LIFE SPAN Seven years WARRANTY Two years from shipping date." | GENERAL DESCRIPTION Doppler, fetal heart rate detector, monitor, w/accessories Handheld doppler based foetal heart rate detector with amplifier loudspeaker. INTENDED USE Ultrasonic fetal heart detectors are low-cost devices used in a variety of healthcare settings to provide audible and visual information about the fetus. Certain conditions, such as vaginal bleeding or lack of fetal movement, can cause anxiety in the mother, making her question the health of her unborn baby. The primary purpose of the fetal heart detector is to provide quick reassurance of fetal well-being to both the mother and the healthcare worker. The fetal heartbeat cannot be heard with an obstetric stethoscope until 24 weeks after conception. Ultrasonic fetal heart detectors can easily detect fetal heart motion throughout the pregnancy, starting as early as 8 weeks. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Doppler, foetal heart rate (FHR) detector, with accessories To be used in basic health infrastructures for routine examination of foetal life, from about 10-12 weeks gestation through to delivery. Self-test is performed each time the device is switchedCapable of detecting foetal heart rates (FHRs) in the range of 50 - 210 bpm, with 1 bpm resolution and 2 bpm accuracy. System reports, with audio-visual alert: operational status, malfunctions and low battery. Advanced noise/disturbance suppression system assures quality diagnostic sound Probe (transducer) with 2 MHz frequency with a probe attached via a cable. Regular size, approx: 18 to 20 cm. The cable can be extended to a 25 cm length when stretched. Probe detector head diameter at least 39 mm. Handheld, weights less than 0.45 kg with probe and batteries. Built-in speaker with volume adjustment. LCD screen displays FHR, pulse indicator and battery status Visual alerts for status, system error and low battery Operates on 3 x AA batteries. Battery life is sufficient for 10 hours continuous use. SUPPLIED WITH 1 x Soft carry bag easy to clean 1 x Instructions manual (instructions for assembly, use and maintenance in English, French, and Spanish) OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - The probe is interchangeable and probes with other frequencies are available: 4Mhz, 5Mhz and 8Mhz." INSTRUCTION FOR USE To be used in basic health infrastructures for routine examination of fetal life, from about 10-12 weeks gestation through to delivery. Doppler fetal heart rate detector should be operated by an adequately trained person only. Clean and disinfect the device after each use. Doppler fetal heart detector must be used and maintained according to manufacturer's instructions" ACCESSORIES/ SPARE PARTS / CONSUMABLES Ultrasound gel, batteries PACKAGING & LABELLING Unit presentation: 1 (one) Doppler, FHR detector, with accessories and Instructions for use. Symbols used according ISO 15223 CE with notified body number REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-1-6:2010 Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - IEC 60601-2-37:2007+AMD1:2015 CSV Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. - IEC 61157:2007+AMD1:2013 CSV Standard means for the reporting of the acoustic output of medical diagnostic ultrasonic equipment. - IEC 61266:1994 Ultrasonics - Hand-held probe Doppler foetal heartbeat detectors - Performance requirements and methods of measurement and reporting. - IEC 62366 : 2007" "ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C - 40°C / 25% – 80% RH. - Storage conditions: -20°C - 55°C / 25% – 93% RH. - Atmospheric pressure: 860 ~ 1060 hPa. - Ingress protection rating: Probe only IPX8. ENVIRONMENTAL REQUIREMENT Rohs compliant ESTIMATED DELIVERY LEAD TIME 120 days. INSTALLATION REQUIREMENTS There is no assembly or commisioning required for this device, other than instaling the batteries. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. COMPONENT OF A KIT No part of a kit. WEIGHT AND VOLUME (packaged) - estimated ESTIMATED LIFE SPAN Seven years WARRANTY Two years from shipping date."
346 | Med. Diagnostic Equip&Supplies | HEMOGLOBINOMETER | Haemoglobin photometer handheld w/accessories | Haemoglobin photometer handheld w/accessories | Type of Immunoassay/ Intended Used: Quantitative determination of hemoglobin in capillary, venous, or arterial whole blood Description: small, portable or handheld device; device-free/disposable preferred. Specimen or sample type to be used: Venous or arterial whole blood / capillary blood Accessories: None Note: If any Accessories applicable, the requirements must be specified by Suppliers whether they are included or not in the package. Sensitivity: More than 80% Population: All population Packaging and Labelling: Individually packaged including Instruction for use. (French, English and Spanish) Product name, size, reference number, manufacture´s name and address, CE mark. Harmonised symbol or equivalent for "sterile or reusable" conditions. Information of storage conditions. All applicable for CE marking or US FDA Clearance. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Kit Stability & Storage: Stable for >12 months with harsh ambient conditions (temperature 5-45°C, humidity 15% to 95%, dusty air, elevation >=2000 meters) and transport stress (48h with fluctuations up to 50°C and down to 0°C) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 Calibration: No calibration/minimal Power Requirement: None/Minimal(110-220V AC current; DC power with rechargeable battery lasting up to 8 hours of testing) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Type of Immunoassay/ Intended Used: Quantitative determination of hemoglobin in capillary, venous, or arterial whole blood Description: small, portable or handheld device; device-free/disposable preferred. Specimen or sample type to be used: Venous or arterial whole blood / capillary blood Accessories: None Note: If any Accessories applicable, the requirements must be specified by Suppliers whether they are included or not in the package. Sensitivity: More than 80% Population: All population Packaging and Labelling: Individually packaged including Instruction for use. (French, English and Spanish) Product name, size, reference number, manufacture´s name and address, CE mark. Harmonised symbol or equivalent for "sterile or reusable" conditions. Information of storage conditions. All applicable for CE marking or US FDA Clearance. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Kit Stability & Storage: Stable for >12 months with harsh ambient conditions (temperature 5-45°C, humidity 15% to 95%, dusty air, elevation >=2000 meters) and transport stress (48h with fluctuations up to 50°C and down to 0°C) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 Calibration: No calibration/minimal Power Requirement: None/Minimal(110-220V AC current; DC power with rechargeable battery lasting up to 8 hours of testing) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
347 | Med. Diagnostic Equip&Supplies | MEASURINGTAPE | Height measuring instrument (0-2.2m) | Height measuring instrument (0-2.2m) | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Device used for measuring the height Product description: Growth monitoring. Device used for measuring the height of adult or child over 2 years when standing upright. Roller-type. Tape material, coated or painted metal Wall-mounted. Measuring range: 0 up to 2 - 2.2 m. Accuracy: 1 mm. Direct reading in cm. Easy readable in low light working situations With wall stop allowing easy fixation Auto roll back in storage case via spring Storage case is fixed to the extension that is placed on head of the patient. Design allows rough handling Wall stop and storage case, plastic and/or metal Sturdy design allows heavy duty use. Supplied with: Text and illustrated user instructions in English/French/Spanish. Instructions for use: Device used for measuring the height of adult or child over 2 years when standing upright. Packaging & Labelling: One (1) Height measurement + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IS0 13485 Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Device used for measuring the height Product description: Growth monitoring. Device used for measuring the height of adult or child over 2 years when standing upright. Roller-type. Tape material, coated or painted metal Wall-mounted. Measuring range: 0 up to 2 - 2.2 m. Accuracy: 1 mm. Direct reading in cm. Easy readable in low light working situations With wall stop allowing easy fixation Auto roll back in storage case via spring Storage case is fixed to the extension that is placed on head of the patient. Design allows rough handling Wall stop and storage case, plastic and/or metal Sturdy design allows heavy duty use. Supplied with: Text and illustrated user instructions in English/French/Spanish. Instructions for use: Device used for measuring the height of adult or child over 2 years when standing upright. Packaging & Labelling: One (1) Height measurement + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IS0 13485 Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
348 | Med. Diagnostic Equip&Supplies | MICROCUVETTE_200 | Capillary tubes for HemoCue Hb 301, P200 | Capillary tubes for HemoCue Hb 301, P200 | Type of Immunoassay/ Intended Used: Quantitative determination of hemoglobin in capillary, venous and arterial whole blood, using a specially designed photometric analyzer. Description: The microcuvette serves both as a pipette and as a measuring cuvette and is for single use only. Specimen or sample type to be used: Venous or arterial whole blood / capillary blood Accessories: None Note: If any Accessories applicable, the requirements must be specified by Suppliers whether they are included or not in the package. Packaging and Labelling: 50, 100, 200/Package Container Instruction for use. (French, English and Spanish) Product name, size, reference number, manufacture´s name and address, CE mark. Harmonized symbol or equivalent for "sterile or reusable" conditions. Information of storage conditions. All applicable for CE marking or US FDA Clearance. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Kit Stability & Storage: The microcuvettes are to be stored at a temperature of 15–30 °C (59–86 °F) and in a dry place. The reagents in the micro-cuvette are moisture sensitive. Screw cap to be tightly closed after use. Do not refrigerate. Shelf life: 24 months from date of manufacture. Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Type of Immunoassay/ Intended Used: Quantitative determination of hemoglobin in capillary, venous and arterial whole blood, using a specially designed photometric analyzer. Description: The microcuvette serves both as a pipette and as a measuring cuvette and is for single use only. Specimen or sample type to be used: Venous or arterial whole blood / capillary blood Accessories: None Note: If any Accessories applicable, the requirements must be specified by Suppliers whether they are included or not in the package. Packaging and Labelling: 50, 100, 200/Package Container Instruction for use. (French, English and Spanish) Product name, size, reference number, manufacture´s name and address, CE mark. Harmonized symbol or equivalent for "sterile or reusable" conditions. Information of storage conditions. All applicable for CE marking or US FDA Clearance. Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Kit Stability & Storage: The microcuvettes are to be stored at a temperature of 15–30 °C (59–86 °F) and in a dry place. The reagents in the micro-cuvette are moisture sensitive. Screw cap to be tightly closed after use. Do not refrigerate. Shelf life: 24 months from date of manufacture. Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
349 | Med. Diagnostic Equip&Supplies | PELVIMETERCOLLYER | Pelvimeter, Collyer type, external | Pelvimeter, Collyer type, external | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to measure the pelvis externally. Product description: Collyer Pelvimeter, use to measure the pelvis externally. Material: Austenitic stainless steel composition: 16 - 18% chromium, 10 - 14% nickel and 2 -3% molybdenum. Graduation in centimeters and inches, maximum width Width of jaws approx.: 300 - 500 mm. Instructions for use: Device used for measuring the pelvis externally. Packaging & Labelling: One (1) Pelvimeter + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IS0 13485: 2016 ISO 10993-1:2009 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to measure the pelvis externally. Product description: Collyer Pelvimeter, use to measure the pelvis externally. Material: Austenitic stainless steel composition: 16 - 18% chromium, 10 - 14% nickel and 2 -3% molybdenum. Graduation in centimeters and inches, maximum width Width of jaws approx.: 300 - 500 mm. Instructions for use: Device used for measuring the pelvis externally. Packaging & Labelling: One (1) Pelvimeter + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IS0 13485: 2016 ISO 10993-1:2009 Environmental requirements: ISO 14001
350 | Med. Diagnostic Equip&Supplies | PREGNCALC | Pregnancy wheel | Pregnancy wheel | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For use during antenatal care and consultations. Product description: Pregnancy Due Date Calculator/ pregnancy wheel / gestational calendar based on last menstrual period. Besides the expected date of delivery, the pregnancy wheel may also assist in defining other important happenings such as the appearance of fetal movement, the detection of the heartbeat, general information, prenatal screening parameters etc. Size: From 4" diameter to 6.75" diameter, approx. cardboard or plastic. PVC is not recommended. Easy to use and helpful Instructions for use: It is a good promotional ruler for midwives, birthing practitioners, nurses or doctors to calculate pregnancy due date Packaging and labelling: 1 Pregnancy Calculator in a bag Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards IS0 13485: 2016 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For use during antenatal care and consultations. Product description: Pregnancy Due Date Calculator/ pregnancy wheel / gestational calendar based on last menstrual period. Besides the expected date of delivery, the pregnancy wheel may also assist in defining other important happenings such as the appearance of fetal movement, the detection of the heartbeat, general information, prenatal screening parameters etc. Size: From 4" diameter to 6.75" diameter, approx. cardboard or plastic. PVC is not recommended. Easy to use and helpful Instructions for use: It is a good promotional ruler for midwives, birthing practitioners, nurses or doctors to calculate pregnancy due date Packaging and labelling: 1 Pregnancy Calculator in a bag Symbols used according ISO 15223 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards IS0 13485: 2016
351 | Med. Diagnostic Equip&Supplies | SCALEINFANT | Scale, infant, beam type, 16kg x 5g | Scale, infant, beam type, 16kg x 5g | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Scale for weighing babies to monitor growth. Product description: Mechanical infant clinical scale, beam type, with tray, for use up to 16kg with 10g precision, with splash cover Material: Outside made of plastic or metal coated with thermo hardened white paint, very strong. Body made of stainless steel, metal (white baked enamel) or plastic. non corroding in hot and humid climatic conditions Measuring range: 0 to approx. 16kg, with precision of 10g. Graduation: minimum every 5g. Readout in kg and in lbs. Easy readable in low light working situations. Reading time: maximum 15 seconds. Includes removable tray for infant. Stable and non-tilting tray. With easy zero adjustment possibility. Adjustable feet allow for horizontal levelling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection. All vital parts made of proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Sturdy design allows heavy duty use. Warranty: minimum 24 months from the date of shipping Supplied with: - Text and pictorial user instructions for operation, preventive maintenance, calibration and troubleshooting in English/French/Spanish - Splash proof carry bag -Certificate of calibration to be provided. Instructions for use: Scale used for weighing babies to monitor growth. Scale must be used and maintained according manufacturer´s instructions. Packaging & Labelling: One (1) scale + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Scale for weighing babies to monitor growth. Product description: Mechanical infant clinical scale, beam type, with tray, for use up to 16kg with 10g precision, with splash cover Material: Outside made of plastic or metal coated with thermo hardened white paint, very strong. Body made of stainless steel, metal (white baked enamel) or plastic. non corroding in hot and humid climatic conditions Measuring range: 0 to approx. 16kg, with precision of 10g. Graduation: minimum every 5g. Readout in kg and in lbs. Easy readable in low light working situations. Reading time: maximum 15 seconds. Includes removable tray for infant. Stable and non-tilting tray. With easy zero adjustment possibility. Adjustable feet allow for horizontal levelling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection. All vital parts made of proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Sturdy design allows heavy duty use. Warranty: minimum 24 months from the date of shipping Supplied with: - Text and pictorial user instructions for operation, preventive maintenance, calibration and troubleshooting in English/French/Spanish - Splash proof carry bag -Certificate of calibration to be provided. Instructions for use: Scale used for weighing babies to monitor growth. Scale must be used and maintained according manufacturer´s instructions. Packaging & Labelling: One (1) scale + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
352 | Med. Diagnostic Equip&Supplies | SCALEINFANTSLING | Sling, for use with infant scale | Sling, for use with infant scale | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Scale for weighing babies to monitor growth. Product description: Sling for use with spring-type infant scale weighing babies to monitor growth. Material: quality cotton fabric, strong and non-elastic. Weight: approx. 110g (+/- 20g). Size: square shape, 80 x 80cm, prior to washing. Edges are double seam finished. Easy to iron. Corners can be knotted together to hold baby securely when suspended from birth weight indicator or scale. Can hold baby upto 5kg. Machine washable up to 40°C with domestic type detergent. Material does not shrink more than 10% with the above indicated washing. Supplied with: Text and pictorial user instructions for usage and maintenance in English/French/Spanish Bag for carrying and/or storage. Instructions for use: Sling for use with spring-type infant scale weighing babies to monitor growth. Packaging & Labelling: One (1) sling in a plastic bag + Information for user Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards IEC 60601-1:2005; IEC 60601-1-1:2000 IEC 60601-1-2:2007; IEC 62366:2007 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Scale for weighing babies to monitor growth. Product description: Sling for use with spring-type infant scale weighing babies to monitor growth. Material: quality cotton fabric, strong and non-elastic. Weight: approx. 110g (+/- 20g). Size: square shape, 80 x 80cm, prior to washing. Edges are double seam finished. Easy to iron. Corners can be knotted together to hold baby securely when suspended from birth weight indicator or scale. Can hold baby upto 5kg. Machine washable up to 40°C with domestic type detergent. Material does not shrink more than 10% with the above indicated washing. Supplied with: Text and pictorial user instructions for usage and maintenance in English/French/Spanish Bag for carrying and/or storage. Instructions for use: Sling for use with spring-type infant scale weighing babies to monitor growth. Packaging & Labelling: One (1) sling in a plastic bag + Information for user Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards IEC 60601-1:2005; IEC 60601-1-1:2000 IEC 60601-1-2:2007; IEC 62366:2007
353 | Med. Diagnostic Equip&Supplies | SCALEMOTHERCHILD | Scale, mother/child, electronic, 200kg | Scale, mother/child, electronic, 200kg | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: allows a child's weight to be measured while being held by an adult. Product description: Measuring range: 0 to 200 kg. Minimum graduation: 100g Accuracy: ± 100g Precision: ± 100g. With taring button, so called "mother-and-child" function. Foot switch on/off button, rubber covered to eliminate accidental turning on during transportation. Auto switch off 180 sec. After weighing process. Easily readable in low light working situations. Reading time: approx. 5 seconds. With footprint marks on the surface of the scale. All vital parts made of rust proof materials. Splash proof and shock resistant body. Powered with 4 or 6 alkaline AA batteries; customer replaceable. Desirable: additional power supply: solar power supply and/or power plug supply. Approx 3.000 measurements per battery set. All metal construction with plastic surface. Display cover: durable plastic. Use of PVC is not recommended for any plastic parts. WARRANTY: Two years Supplied with: - Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish, appropriately illustrated with pictograms and graphics. - Separately packed batteries. - Splash proof carry bag. Instructions for use: Scale used for weighing mother and child. Scale must be used and maintained according manufacturer´s instructions Accessories/ spare parts/consumables: 1x packed of batteries Packaging & Labelling: - One piece per carton + Information for user Symbols used according ISO 15223 CE mark with notified body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IEC 60601-1:2005; IEC 60601-1-1:2000 IEC 60601-1-2:2007; IEC 62366:2007 Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: allows a child's weight to be measured while being held by an adult. Product description: Measuring range: 0 to 200 kg. Minimum graduation: 100g Accuracy: ± 100g Precision: ± 100g. With taring button, so called "mother-and-child" function. Foot switch on/off button, rubber covered to eliminate accidental turning on during transportation. Auto switch off 180 sec. After weighing process. Easily readable in low light working situations. Reading time: approx. 5 seconds. With footprint marks on the surface of the scale. All vital parts made of rust proof materials. Splash proof and shock resistant body. Powered with 4 or 6 alkaline AA batteries; customer replaceable. Desirable: additional power supply: solar power supply and/or power plug supply. Approx 3.000 measurements per battery set. All metal construction with plastic surface. Display cover: durable plastic. Use of PVC is not recommended for any plastic parts. WARRANTY: Two years Supplied with: - Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish, appropriately illustrated with pictograms and graphics. - Separately packed batteries. - Splash proof carry bag. Instructions for use: Scale used for weighing mother and child. Scale must be used and maintained according manufacturer´s instructions Accessories/ spare parts/consumables: 1x packed of batteries Packaging & Labelling: - One piece per carton + Information for user Symbols used according ISO 15223 CE mark with notified body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards IEC 60601-1:2005; IEC 60601-1-1:2000 IEC 60601-1-2:2007; IEC 62366:2007 Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
354 | Med. Diagnostic Equip&Supplies | SCALEPHYSICIAN | Scale, physician, adult beam balance, 0-200kg | Scale, physician, adult beam balance, 0-200kg | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: All metal construction Leaf spring beneath poise to facilitate zeroing. Knob on handpost (support pole) to be glued and screwed in place. Beam mechanical scale with 2 sliding weights, one for kg and one for gram. Beam to be secured so that steelyard rods cannot become disconnected. Calibration screw of stainless steel. Strong and stable. Paint thickness 50 microns or more. All screw heads used for assembly to be cross and/or hexagonal. All exposed parts treated with a rust inhibitor. All moving parts must be secured and retained during shipment. Non-skidding surface. Attached measuring apparatus: approximately 60 - 230 cm. Graduation: every 50 g. With 100 g precision. Measuring range: 0 to 200 kg. Colour: White. Warranty: 2 yrs Supplied with: - Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish, appropriately illustrated with pictograms and graphics. Instructions for use: Physician scale beam balance used for growth monitoring of adults. Must be used and maintained according manufacturer´s instructions. Packaging & Labelling: Each scale must be individually packed in a triple wall carton with sufficient insert wrappings to protect the scale during transit. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: IEC 62366:2007 where possible Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: All metal construction Leaf spring beneath poise to facilitate zeroing. Knob on handpost (support pole) to be glued and screwed in place. Beam mechanical scale with 2 sliding weights, one for kg and one for gram. Beam to be secured so that steelyard rods cannot become disconnected. Calibration screw of stainless steel. Strong and stable. Paint thickness 50 microns or more. All screw heads used for assembly to be cross and/or hexagonal. All exposed parts treated with a rust inhibitor. All moving parts must be secured and retained during shipment. Non-skidding surface. Attached measuring apparatus: approximately 60 - 230 cm. Graduation: every 50 g. With 100 g precision. Measuring range: 0 to 200 kg. Colour: White. Warranty: 2 yrs Supplied with: - Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish, appropriately illustrated with pictograms and graphics. Instructions for use: Physician scale beam balance used for growth monitoring of adults. Must be used and maintained according manufacturer´s instructions. Packaging & Labelling: Each scale must be individually packed in a triple wall carton with sufficient insert wrappings to protect the scale during transit. Symbols used according ISO 15223 CE mark with notified body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: IEC 62366:2007 where possible Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
355 | Med. Diagnostic Equip&Supplies | SCALESPRING | Scale, infant, spring-type, 5kg x 25g | Scale, infant, spring-type, 5kg x 25g | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to measure the weight Product description: Scale for weighing babies to monitor growth. Materials: hooks - metal; dial - plastic or metal ( stainless steel preferred); body - metal ( stainless steel, preferred) or plastic. PVC is not recommended Measuring range: 0 to 5kg Graduation: minimum every 25g. Enhanced readability via reflecting surface behind needle. Readout in kg and in lbs. Easy readable in low light working situations. Reading time: maximum 15 seconds. With easy zero adjustment possibility. Fit with 2 hooks: upper for fixation and lower for attaching weighing trouser or sling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection. All vital parts made of rust proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Sturdy design allows heavy duty use. Warranty 2 yrs Supplied with: - Text and pictorial user instructions for operation, preventive maintenance and troubleshooting in English/French/Spanish - Splash proof carry bag Instructions for use: Scale used for weighing babies to monitor growth. Scale must be used and maintained according manufacturer´s instructions. Packaging & Labelling: One (1) scale + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to measure the weight Product description: Scale for weighing babies to monitor growth. Materials: hooks - metal; dial - plastic or metal ( stainless steel preferred); body - metal ( stainless steel, preferred) or plastic. PVC is not recommended Measuring range: 0 to 5kg Graduation: minimum every 25g. Enhanced readability via reflecting surface behind needle. Readout in kg and in lbs. Easy readable in low light working situations. Reading time: maximum 15 seconds. With easy zero adjustment possibility. Fit with 2 hooks: upper for fixation and lower for attaching weighing trouser or sling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection. All vital parts made of rust proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Sturdy design allows heavy duty use. Warranty 2 yrs Supplied with: - Text and pictorial user instructions for operation, preventive maintenance and troubleshooting in English/French/Spanish - Splash proof carry bag Instructions for use: Scale used for weighing babies to monitor growth. Scale must be used and maintained according manufacturer´s instructions. Packaging & Labelling: One (1) scale + Information for user Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
356 | Med. Diagnostic Equip&Supplies | SPHYGMOANEROID | Sphygmomanometer, aneroid, adult | Sphygmomanometer, aneroid, adult | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For physical examination, Diagnosis hypertention, Monitoring by measuring blood pressure. Product description: Device used for the indirect (non-invasive) measurement of arterial blood pressure. The systolic and diastolic pressure is usually assessed by listening to Korotkoff sounds generated by arterial blood flow using a stethoscope simultaneously. The sphygmomanometer is composed of cloth cuff containing an inflatable bag. Connected via a tube to a flexible bulb with valve and integrated manometer needle gauge. Material for cuff: durable nylon, non-deformable, washable at 30ºC. Material for tube: rubber. Very strong cuff with double velcro fastening, enabling it to be adjusted to fit tightly around the arm. Cuff reinforced at both ends. Dimensions of cuff: approx. 570 x 145mm. The bag is inflated by means of a flexible bulb connected via a tube (length 60cm flexible and reliable quick connector). Rubber tubes to be detachable from other parts, allowing periodic cutting of decayed ends The quick connector can easily be connected to all types of inflation bulbs. Rubber inflation bulb with integrated manometer needle gauge. Glass and metal aneroid pressure gauge with needle. Dial graduation: 0 to 300mmHG. Pressure gauge to allow reading of pressure to 2mmHg accuracy With pressure release valve. Easily to grip bulb. Warranty 2 years Instructions for use: For the measurement of arterial blood pressure. Operating conditions: Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Packaging and labelling: as per regulations applied to the device One (1) sphygmomanometer (adult) in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark with Notified Number Body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards (as per the latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 1060-3:1997+A2:2009 Non-invasive sphygmomanometers - Part 3: Supplementary requirements for electro-mechanical blood pressure measuring systems EN 1060-4:2004 Non-invasive sphygmomanometers - Part 4: Test procedures to determine the overall system accuracy of automated non-invasive sphygmomanometers ISO 81060-1:2007 Non-invasive sphygmomanometers - Part 1: Requirements and test methods for non-automated measurement type ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For physical examination, Diagnosis hypertention, Monitoring by measuring blood pressure. Product description: Device used for the indirect (non-invasive) measurement of arterial blood pressure. The systolic and diastolic pressure is usually assessed by listening to Korotkoff sounds generated by arterial blood flow using a stethoscope simultaneously. The sphygmomanometer is composed of cloth cuff containing an inflatable bag. Connected via a tube to a flexible bulb with valve and integrated manometer needle gauge. Material for cuff: durable nylon, non-deformable, washable at 30ºC. Material for tube: rubber. Very strong cuff with double velcro fastening, enabling it to be adjusted to fit tightly around the arm. Cuff reinforced at both ends. Dimensions of cuff: approx. 570 x 145mm. The bag is inflated by means of a flexible bulb connected via a tube (length 60cm flexible and reliable quick connector). Rubber tubes to be detachable from other parts, allowing periodic cutting of decayed ends The quick connector can easily be connected to all types of inflation bulbs. Rubber inflation bulb with integrated manometer needle gauge. Glass and metal aneroid pressure gauge with needle. Dial graduation: 0 to 300mmHG. Pressure gauge to allow reading of pressure to 2mmHg accuracy With pressure release valve. Easily to grip bulb. Warranty 2 years Instructions for use: For the measurement of arterial blood pressure. Operating conditions: Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Packaging and labelling: as per regulations applied to the device One (1) sphygmomanometer (adult) in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark with Notified Number Body Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards (as per the latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 1060-3:1997+A2:2009 Non-invasive sphygmomanometers - Part 3: Supplementary requirements for electro-mechanical blood pressure measuring systems EN 1060-4:2004 Non-invasive sphygmomanometers - Part 4: Test procedures to determine the overall system accuracy of automated non-invasive sphygmomanometers ISO 81060-1:2007 Non-invasive sphygmomanometers - Part 1: Requirements and test methods for non-automated measurement type ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types".
357 | Med. Diagnostic Equip&Supplies | New | Stethoscope, Binaural, Complete | Stethoscope, Binaural, Complete | GENERAL DESCRIPTION: Stethoscope, binaural, complete INTENDED USE A mechanical listening device designed for listening to sounds from the heart, lungs, and/or gastrointestinal tract from adult and paediatric patients. It comprises a membrane at the listening head connected by a dual tube to the headgear with ear olives that are placed into the users’ ears. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Stethoscope of either ""Littman"" or ""Rappaport"" type, for listening to sounds within the body walls, used for pulmonary and cardiac auscultation. Double cup, dual-use (adult and paediatric auscultation) chest piece in stainless steel or chromed brass or Zinc alloy. Adult diaphragm: 43mm to 45.5 mm Paediatric diaphragm: 28mm to 31.5 mm Adjustable arms in stainless steel or chrome brass with flexible spring treated to give lasting spring and maximum reliability and comfort. Sensitivity 3.2 dB to 26dB in a range from 50Hz to 500Hz for cardiology. Sensitivity 8.1 dB in a range from 600Hz to 1,500Hz for pneumology." "Y tube treated rubber with large diameter: 10 mm. Tube made of PVC and is crack resistant. Tube impervious to outside noises, guaranteeing full transmission of sound, good auditive quality. Tube diameter: outer diameter 10mm, inner diameter 4.8mm. Tube length 400 - 560mm. Removable plastic ear-pieces. Easy to dismantle, and therefore to clean and disinfect. Designed for frequent and easy disassembly and disinfection with hospital-grade products. Easy to use for training health technicians to auscultation. Latex-free. Supplied with: 1 spare adult diaphragm, 1 spare paediatric diaphragm and 1 spare pair of removable plastic ear-pieces." "SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x spare adult diaphragm. 1 x spare paediatric diaphragm. 1 x set spare earpieces. ESTIMATED LIFE SPAN 5 years WARRANTY Two years PACKAGING & LABELLING One (1) binaural stethoscope in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark" REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class Im device SAFETY AND PRODUCT STANDARDS ISO 10993-1:2009" ESTIMATED LIFE SPAN 5 years WARRANTY Two years | GENERAL DESCRIPTION: Stethoscope, binaural, complete INTENDED USE A mechanical listening device designed for listening to sounds from the heart, lungs, and/or gastrointestinal tract from adult and paediatric patients. It comprises a membrane at the listening head connected by a dual tube to the headgear with ear olives that are placed into the users’ ears. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Stethoscope of either ""Littman"" or ""Rappaport"" type, for listening to sounds within the body walls, used for pulmonary and cardiac auscultation. Double cup, dual-use (adult and paediatric auscultation) chest piece in stainless steel or chromed brass or Zinc alloy. Adult diaphragm: 43mm to 45.5 mm Paediatric diaphragm: 28mm to 31.5 mm Adjustable arms in stainless steel or chrome brass with flexible spring treated to give lasting spring and maximum reliability and comfort. Sensitivity 3.2 dB to 26dB in a range from 50Hz to 500Hz for cardiology. Sensitivity 8.1 dB in a range from 600Hz to 1,500Hz for pneumology." "Y tube treated rubber with large diameter: 10 mm. Tube made of PVC and is crack resistant. Tube impervious to outside noises, guaranteeing full transmission of sound, good auditive quality. Tube diameter: outer diameter 10mm, inner diameter 4.8mm. Tube length 400 - 560mm. Removable plastic ear-pieces. Easy to dismantle, and therefore to clean and disinfect. Designed for frequent and easy disassembly and disinfection with hospital-grade products. Easy to use for training health technicians to auscultation. Latex-free. Supplied with: 1 spare adult diaphragm, 1 spare paediatric diaphragm and 1 spare pair of removable plastic ear-pieces." "SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x spare adult diaphragm. 1 x spare paediatric diaphragm. 1 x set spare earpieces. ESTIMATED LIFE SPAN 5 years WARRANTY Two years PACKAGING & LABELLING One (1) binaural stethoscope in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark" REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class Im device SAFETY AND PRODUCT STANDARDS ISO 10993-1:2009" ESTIMATED LIFE SPAN 5 years WARRANTY Two years
358 | Med. Diagnostic Equip&Supplies | STETHOSCOPEPINARD | Stethoscope, obstetrical, Pinard, monoaural | Stethoscope, obstetrical, Pinard, monoaural | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: To listen to the foetal heart beat Product description: Foetal Stethoscope Pinard Monoaural, for diagnosis and prenatal care of pregnant women to listen to the foetal heart. Material: unbreakable plastic or aluminum Monoaural 1 piece Length: approx. 150mm Instructions for use: For diagnosis and prenatal care of pregnant women to listen to the foetal heart. Packaging and labelling: One (1) foetal stethoscope in a plastic bag with manufacturer's instruction for use in English, French and Spanish. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". PVC is not recommended | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: To listen to the foetal heart beat Product description: Foetal Stethoscope Pinard Monoaural, for diagnosis and prenatal care of pregnant women to listen to the foetal heart. Material: unbreakable plastic or aluminum Monoaural 1 piece Length: approx. 150mm Instructions for use: For diagnosis and prenatal care of pregnant women to listen to the foetal heart. Packaging and labelling: One (1) foetal stethoscope in a plastic bag with manufacturer's instruction for use in English, French and Spanish. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards ISO 10993-1 Biological Evaluation of Medical Devices Part 1:Evaluation and Testing. Environmental requirements: ISO 14001 Chemicals used shall not be "Carcinogenic, Mutagenic or Toxic to Reproduction, Persistent & Bio accumulative types". PVC is not recommended
359 | Med. Diagnostic Equip&Supplies | THERMOMETER | Thermometer, clinical, digital 32°C-43°C | Thermometer, clinical, digital 32°C-43°C | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Electronic device to measure a patient's body temperature Product description: Electronic device (without mercury) to measure the body temperature of children and adults in Celsius degrees. Safe to use, no glass, no mercury. Digital thermometer Celsius scale. Liquid crystal display (LCD), easy to read. Measurement range: 32°C to 43°C Accuracy: +/- 0.1°C between 35°C to 41°C Measurement time: 30 to 40 seconds. Long-life battery: minimum 200 hours (about 1000 measurements). Automatic shut-off after some minutes. Beep sound and switch off. Water proof for ease of cleaning Resistant to chlorine ("high level" disinfection). Operating ambient temperature: from +10º C to 35º C Battery powered. Low battery indicator. Battery powered (min. 4000 measurement for full battery) Supplied with clear instructions for use/preventive maintenance. Instructions for use: To measure temperature of children and adults. It is recommended to follow the manufacturer's instructions for use / preventive maintenance. Operating conditions: 10 - 40°C / 85% RH. Packaging and labelling: One (1) thermometer in storage case with manufacturer's instruction for use in English, French and Spanish. Symbols used according ISO 15223 CE mark with Notified Body number Accessories/ spare parts/consumables: Pack of batteries Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im device Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 12470-3:2000+A1:2009 Clinical thermometers - Part 3: Performance of compact electrical thermometers (non-predictive and predictive) with maximum device ISO 80601-2-56:2009. Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement. IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests Environmental requirements: ISO 14001 Material of batteries environmental friendly Not use of PVC | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Electronic device to measure a patient's body temperature Product description: Electronic device (without mercury) to measure the body temperature of children and adults in Celsius degrees. Safe to use, no glass, no mercury. Digital thermometer Celsius scale. Liquid crystal display (LCD), easy to read. Measurement range: 32°C to 43°C Accuracy: +/- 0.1°C between 35°C to 41°C Measurement time: 30 to 40 seconds. Long-life battery: minimum 200 hours (about 1000 measurements). Automatic shut-off after some minutes. Beep sound and switch off. Water proof for ease of cleaning Resistant to chlorine ("high level" disinfection). Operating ambient temperature: from +10º C to 35º C Battery powered. Low battery indicator. Battery powered (min. 4000 measurement for full battery) Supplied with clear instructions for use/preventive maintenance. Instructions for use: To measure temperature of children and adults. It is recommended to follow the manufacturer's instructions for use / preventive maintenance. Operating conditions: 10 - 40°C / 85% RH. Packaging and labelling: One (1) thermometer in storage case with manufacturer's instruction for use in English, French and Spanish. Symbols used according ISO 15223 CE mark with Notified Body number Accessories/ spare parts/consumables: Pack of batteries Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 12470-3:2000+A1:2009 Clinical thermometers - Part 3: Performance of compact electrical thermometers (non-predictive and predictive) with maximum device ISO 80601-2-56:2009. Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement. IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests Environmental requirements: ISO 14001 Material of batteries environmental friendly Not use of PVC
360 | Med. Diagnostic Equip&Supplies |  | Timer, for respiration, ARI | Timer, for respiration, ARI | Type of Immunoassay/ Intended Used: to count the breath of children to detect acute respiratory infections Description: ARI Timer, timer comes with a fixed 60 seconds' countdown to determine the respiration rate of newborns and young children for suspected acute respiratory infection diagnosis. Easily accessible button located at the center of the device and activation by pressing the button. One audible beep at the start and two audible beeps at the end of the cycle, 60 sec. The unit does not make any sounds during the countdown, flashing red signal during countdown Accuracy = ± 0.1 second. The unit is battery powered and full batteries allow the unit to support 10 ,000 cycles of 60 seconds as a minimum. Not replaceable battery At the press of the start button: 1. the device produces an audible beep to indicate the countdown has started. 2. Provides red visual indicator informing the user that the device has been activated and is counting down (For example a flashing LED could indicate the unit is currently counting down.) 3. After 60 seconds is elapsed, a double audible beep produced, and the visual indicator stops indicting the unit has stopped counting down. 4. Pressing the button during an ongoing countdown reset the unit which means stop the countdown and the visual indicator. A provision included to ensure that the battery is not connected to the internal circuit at the time of delivery. This to avoid unwanted activation, leakage or accidental short-circuit of the unit. Operating conditions: 10 - 40°C / 85% RH. Supplied with 1 x cord (lanyard) for wearing timer around the neck Instruction for use and maintenance in English, French and Spanish. Precautions for Use such as "To ensure maximum battery life, store in cool dry place" etc. Packaging/labelling Primary packaging: Unit of use. One (1) Timer, respiration for ARI + instructions for use in English French and Spanish. Symbols used according to ISO 15223 Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Batch number, Information for storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol) and Information for handling. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note: For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Classification: Classified under EU MDR 2017/745 as Class II device Warranty: 2 yrs minimum from the date of shipping Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance ISO 14001 Material of batteries environmentally friendly Not use of PVC | Type of Immunoassay/ Intended Used: to count the breath of children to detect acute respiratory infections Description: ARI Timer, timer comes with a fixed 60 seconds' countdown to determine the respiration rate of newborns and young children for suspected acute respiratory infection diagnosis. Easily accessible button located at the center of the device and activation by pressing the button. One audible beep at the start and two audible beeps at the end of the cycle, 60 sec. The unit does not make any sounds during the countdown, flashing red signal during countdown Accuracy = ± 0.1 second. The unit is battery powered and full batteries allow the unit to support 10 ,000 cycles of 60 seconds as a minimum. Not replaceable battery At the press of the start button: 1. the device produces an audible beep to indicate the countdown has started. 2. Provides red visual indicator informing the user that the device has been activated and is counting down (For example a flashing LED could indicate the unit is currently counting down.) 3. After 60 seconds is elapsed, a double audible beep produced, and the visual indicator stops indicting the unit has stopped counting down. 4. Pressing the button during an ongoing countdown reset the unit which means stop the countdown and the visual indicator. A provision included to ensure that the battery is not connected to the internal circuit at the time of delivery. This to avoid unwanted activation, leakage or accidental short-circuit of the unit. Operating conditions: 10 - 40°C / 85% RH. Supplied with 1 x cord (lanyard) for wearing timer around the neck Instruction for use and maintenance in English, French and Spanish. Precautions for Use such as "To ensure maximum battery life, store in cool dry place" etc. Packaging/labelling Primary packaging: Unit of use. One (1) Timer, respiration for ARI + instructions for use in English French and Spanish. Symbols used according to ISO 15223 Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Batch number, Information for storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol) and Information for handling. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note: For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485: 2016 Classification: Classified under EU MDR 2017/745 as Class II device Warranty: 2 yrs minimum from the date of shipping Safety & product Standards: Must comply with following standards (as per latest revision of standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices Environmental requirements: Rohs compliance ISO 14001 Material of batteries environmentally friendly Not use of PVC
361 | Med. Diagnostic Equip&Supplies | ULTRASOUNDSCANNER | Ultrasound, scanner, mobile w/ printer and accessories | Ultrasound, scanner, mobile w/ printer and accessories | IMPORTANT NOTE: These technical specifications are generic & may vary depending on the selected Vendor Intended purpose of use: Portable type for diagnostic imaging applications Product description: High performance ultrasound scanner To be used in health infrastructures for diagnostic imaging applications: general abdomen, obstetrics & gynaecology, small parts, neonatal head & hip, cardio-vascular, urology & intra-operative System integrates: scanner, 2 probes, trolley & video-printer Compact & lightweight, easy to operate Alphanumeric keyboard with trackball & time gain control Piezoelectric probes, electronically scanned: convex & linear Imaging display modes:B, dual B,M,B & M Adjustable field-of-view, 6 level zoom Imaging technologies: dynamic frequency imaging, multi-stage focusing, aperture control Depth range selection: convex sector image & linear image,3 steps Image orientation: lateral & vertical inversion (in B mode) Freeze function with storage of approx 50 images Self-test is performed each time the machine is switched on System reports, with audio-visual alert on: operational status & malfunctions Measurements&Analysis: Calibre control: trackball B-mode image: distance, area & circumference by ellipse & trace method, volume, ratio, gestational age, foetal weight, angle Gestational table: user programmable M-mode: velocity, time interval, depth, heart rate, LV function Alphanumerics&graphics: Text annotations & body markers Automatic display: date & time, focal point setting, image orientation indicator, image, scrolled position, distance scale mark, M-mode time mark, gray scale for calibration Automatic display adjustment control, on/off for: gain setting, near gain setting, far gain setting, frequency probe, puncture guide line High resolution B/W monitor 10-inch/25cm (diagonal) provided with reflection filter Image gray-scale: 256 levels Video output: 625 lines/frame Two transducer ports leaves 2 probes permanently available, electronic switch between probes Data communication interface: RS232, BNC, IEEE, USB or equivalent Power consumption system approx: 1000W Resolution approx: 400 dpi Gray level gradations: 256 Image size, approx: 90x70mm Print time per image, approx: 3 sec Automatic paper roll load Audio-visual indication on loss of signal and paper empty Aspect ratio, user selectable: 4:3 or 1:1 Data communication interface: BNC, USB, IEEE or equivalent Side handle facilitates positioning Surface fits ultrasound scanner as well as folded-open keyboard Under surface shelf fits video printer Side support provides for safe storage of 2 probes Dimensions of trolley: approx: 60x50x1.20cm (LxWxH) Sturdy & stable construction on 4 antistatic ball-bearing swivel castors, 2 with breaks. Trolley to include shelf space for image printer and documentation Unit layout to enable easy cleaning and sterilization of all surfaces Warranty: 2 yrs minimum from the date of shipping Power supply: the right voltage, plug & socket type that fit the country it is delivering the goods to as per UNFPA PO. Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure. Mains supply cable to be at least 3m in length. operating conditions: Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Documentation requirements: User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Supplied with: 1 x scanner, ultrasound portable system, main unit with accessories. 1 x standard electronic convex sector probe: frequency 3.5MHz, scanning angle 60°, radius 60mm 1 x ultrasonic gel 5-litre (supplied as one bottle or multiple). 1 x Instructions for assembly, use and maintenance in English, French and Spanish. Additional accessories must be made available at extra costs for orders from UNFPA: 1 convex abdominal probe, frequency range: 2.5/3.5/5.0MHz 1 convex transvaginal probe, frequency range: 4.5/6.5/8.0MHz 2 tubes of ultrasound gel, approx 350ml 1 set of spare fuses 1 digital B/W video printer with data connecting cable to ultrasound scanner: Thermal printing head 1xPlastic protective dustcover 1 printer head cleaning sheet 1 set of 10 video printer paper rolls, length approx 20m 1 matching trolley: 1xInstructions for assembly, use & maintenance in English, French & Spanish List of accessories/parts Instructions for use: To be used in health infrastructures for diagnostic imaging applications: general abdomen, obstetrics & gynaecology, small parts, neonatal head & hip, cardio-vascular, urology & intra-operative Clean & disinfect probes after each use The scanner should be operated by adequately trained staff only System must be used & maintained strictly according manufacturer's instructions Recommendation: wall plug shall be made available locally Accessories/spare parts/consumables: Spare fuses, B/W printer, ultrasound gel tube Packaging&Labelling: shall be as per regulations applied to the device Unit presentation: one scanner, ultrasound, mobile, with accessories & with manufacturer's instruction for use & service in English, French & Spanish Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety&Product Standards (must comply to the latest revisions of the standards): ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-37:2007 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment Environmental: RoSH compliance | IMPORTANT NOTE: These technical specifications are generic & may vary depending on the selected Vendor Intended purpose of use: Portable type for diagnostic imaging applications Product description: High performance ultrasound scanner To be used in health infrastructures for diagnostic imaging applications: general abdomen, obstetrics & gynaecology, small parts, neonatal head & hip, cardio-vascular, urology & intra-operative System integrates: scanner, 2 probes, trolley & video-printer (note : trolley and printer are considered as additional accessories, please refer to descriptions under “Additional Accessories”) Compact & lightweight, easy to operate Alphanumeric keyboard with trackball & time gain control Piezoelectric probes, electronically scanned: convex & linear Imaging display modes: B, dual B,M,B & M Adjustable field-of-view, 6 level zoom Imaging technologies: dynamic frequency imaging, multi-stage focusing, aperture control Depth range selection: convex sector image & linear image,3 steps Image orientation: lateral & vertical inversion (in B mode) Freeze function with storage of approx 50 images Self-test is performed each time the machine is switched on System reports, with audio-visual alert on: operational status & malfunctions Measurements & Analysis: Calibre control: trackball B-mode image: distance, area & circumference by ellipse & trace method, volume, ratio, gestational age, foetal weight, angle Gestational table: user programmable M-mode: velocity, time interval, depth, heart rate, LV function Alphanumerics & graphics: Text annotations & body markers Automatic display: date & time, focal point setting, image orientation indicator, image, scrolled position, distance scale mark, M-mode time mark, gray scale for calibration Automatic display adjustment control, on/off for: gain setting, near gain setting, far gain setting, frequency probe, puncture guide line High resolution B/W monitor 10-inch/25cm (diagonal) provided with reflection filter Image gray-scale: 256 levels Video output: 625 lines/frame Two transducer ports leaves 2 probes permanently available, electronic switch between probes Data communication interface: RS232, BNC, IEEE, USB or equivalent Power consumption system approx: 1000W Resolution approx: 400 dpi Gray level gradations: 256 Image size, approx: 90x70mm Print time per image, approx: 3 sec Automatic paper roll load Audio-visual indication on loss of signal and paper empty Aspect ratio, user selectable: 4:3 or 1:1 Data communication interface: BNC, USB, IEEE or equivalent Side handle facilitates positioning Surface fits ultrasound scanner as well as folded-open keyboard Under surface shelf fits video printer Side support provides for safe storage of 2 probes Dimensions of trolley: approx: 60x50x1.20cm (LxWxH) Sturdy & stable construction on 4 antistatic ball-bearing swivel castors, 2 with breaks. Trolley to include shelf space for image printer and documentation Unit layout to enable easy cleaning and sterilization of all surfaces Warranty: 2 yrs minimum from the date of shipping Power supply: the right voltage, plug & socket type that fit the country it is delivering the goods to as per UNFPA PO. Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure. Mains supply cable to be at least 3m in length. Operating conditions: Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Documentation requirements: User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Supplied with: 1 x scanner, ultrasound portable system, main unit with accessories. 1 x standard electronic convex sector probe: frequency 3.5MHz, scanning angle 60°, radius 60mm 1x Endocavitary (transvaginal) probe, frequency range: 4.5/6.5/8.0MHz 1 x ultrasonic gel 5-litre (supplied as one bottle or multiple). 1 x Instructions for assembly, use and maintenance in English, French and Spanish. 1x Accessories / parts list Additional accessories must be made available at extra costs for orders from UNFPA: 1 convex abdominal probe, frequency range: 2.5/3.5/5.0MHz 1 convex transvaginal probe, frequency range: 4.5/6.5/8.0MHz 2 tubes of ultrasound gel, approx 350ml 1 set of spare fuses 1 digital B/W video printer with data connecting cable to ultrasound scanner: Thermal printing head 1xPlastic protective dustcover 1 printer head cleaning sheet 1 set of 10 video printer paper rolls, length approx 20m 1 matching trolley
362 | Medical & Surgical Instruments | SPATULA_RETRACTOR | Spatula, Ribbon retractor, malleable, 27 mm x 250 mm | Spatula, Ribbon retractor, malleable, 27 mm x 250 mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Used to retract tissues, muscles or viscera in abdominal incisions. Materials: Austenitic steel (non-quenched, non-magnetic steel). Specifications: Depressor spatula, rectangular with rounded ends. Dimensions: approx. 27 x 250mm. Reusable Instructions for use: Used to retract tissues, muscles or viscera in abdominal incisions. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacture's instruction for use. Packaging and labelling: One (1) spatula in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Used to retract tissues, muscles or viscera in abdominal incisions. Materials: Austenitic steel (non-quenched, non-magnetic steel). Specifications: Depressor spatula, rectangular with rounded ends. Dimensions: approx. 27 x 250mm. Reusable Instructions for use: Used to retract tissues, muscles or viscera in abdominal incisions. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacture's instruction for use. Packaging and labelling: One (1) spatula in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer
363 | Medical & Surgical Instruments | BLADESURGIC10 | Blade, scalpel, sterile, single use, no.10 | Blade, scalpel, sterile, single use, no.10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 10) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Curved cutting edge blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, r+D5eference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 10) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Curved cutting edge blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, r+D5eference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001
364 | Medical & Surgical Instruments | BLADESURGIC11_1 | Blade, scalpel, sterile, single use, no. 11, P1 | Blade, scalpel, sterile, single use, no. 11, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 11) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Number 11 engraved on the blade, Elongated triangular blade, sharpened along the hypotenuse edge, with a strong pointed tip Blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 11) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Number 11 engraved on the blade, Elongated triangular blade, sharpened along the hypotenuse edge, with a strong pointed tip Blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001
365 | Medical & Surgical Instruments | BLADESURGIC11_100 | Blade, scalpel, sterile, single use, no. 11, P100 | Blade, scalpel, sterile, single use, no. 11, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 11) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Number 11 engraved on the blade, Elongated triangular blade, sharpened along the hypotenuse edge, with a strong pointed tip Blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 11) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Number 11 engraved on the blade, Elongated triangular blade, sharpened along the hypotenuse edge, with a strong pointed tip Blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001
366 | Medical & Surgical Instruments | BLADESURGIC15_1 | Blade, scalpel, sterile, single use, no. 15, P1 | Blade, scalpel, sterile, single use, no. 15, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 15) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Small curved cutting edge blde. Standard scalpel handle n°3 and 7 Number 15 engraved on the blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 15) for use in minor surgery with standard scalpel handle no.3 and 7 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Small curved cutting edge blde. Standard scalpel handle n°3 and 7 Number 15 engraved on the blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.3 or 7. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001
367 | Medical & Surgical Instruments | BLADESURGIC22_100 | Blade, scalpel, sterile, single use, no. 22, P100 | Blade, scalpel, sterile, single use, no. 22, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 22) for use in minor surgery with standard scalpel handle no.4 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Curved cutting edge blde. Standard scalpel handle n°4 Number 22 engraved on the blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.4. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Mounted on a handle and used as surgical instrument to cut and dissect tissues Product description: Basic cutting instrument used for surgical incisions. Surgical blade (no. 22) for use in minor surgery with standard scalpel handle no.4 Material: Martensitic stainless steel (quenched, magnetic steel) Hardness: 50 HRC to 58 HRC. Sterile and single use. The number indicates the shape of the blade and 'Number' engraved on the blade. Curved cutting edge blde. Standard scalpel handle n°4 Number 22 engraved on the blade Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Scalpel handle, no.4. Instructions for use: Single use oly To expose the vaginal cavity The surgical blade is single use only. Use aseptic techniques. Wear gloves, always apply and remove the blade from the scalpel handle with forceps. Put the used blades into a sealed container to prevent injuries. After use, collect and destroy the sealed containers by incineration in a controlled environment. Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging, may only be used if packaging is undamaged. Packaging & Labelling: shall be as per regulations applied to the device Individually wrapped in a reinforced laminated foil peel pack, sterile, single-use. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark with Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes (ISO 11607-2:2006) ISO 7153-1 Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions Instructions to CO: CO to order the handle separately. Environmental: ISO 14001
368 | Medical & Surgical Instruments | CLAMPINTESTINAL23 | Forceps, intestinal clamp, Doyen, 23cm, curved | Forceps, intestinal clamp, Doyen, 23cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for atraumatic clamping/occlusion of a section intestine while suturing, resection or anastomosis. Product description: Doyen Forceps are a lightweight, non-crushing clamp. Commonly used to grasp large organs such as lungs, or intestines, the Doyen forceps have long jaws and shorter shanks. Intestinal clamp, Doyen, curved, used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. U​sed for transverse non-crushing occlusion of a section of intestine (while suturing, resection or anastomosis). Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Jaws very springy and soft. Features: Non traumatic Occlusion: The form and mechanical action must permit precise closing. Multiple ratchet Box lock. Highly impact resistant. Length: approx 230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. The jaws must be meticulously cleaned. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revisions of the standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental: ISO 14001 Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for atraumatic clamping/occlusion of a section intestine while suturing, resection or anastomosis. Product description: Doyen Forceps are a lightweight, non-crushing clamp. Commonly used to grasp large organs such as lungs, or intestines, the Doyen forceps have long jaws and shorter shanks. Intestinal clamp, Doyen, curved, used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. U​sed for transverse non-crushing occlusion of a section of intestine (while suturing, resection or anastomosis). Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Jaws very springy and soft. Features: Non traumatic Occlusion: The form and mechanical action must permit precise closing. Multiple ratchet Box lock. Highly impact resistant. Length: approx 230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. The jaws must be meticulously cleaned. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revisions of the standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental: ISO 14001 Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices
369 | Medical & Surgical Instruments | CLAMPINTESTINALST | Forceps, intestinal, clamp, Doyen, 23cm, straight | Forceps, intestinal, clamp, Doyen, 23cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for atraumatic clamping/occlusion of a section intestine while suturing, resection or anastomosis. Product description: Doyen Forceps are a lightweight, non-crushing clamp. Commonly used to grasp large organs such as lungs, or intestines, the Doyen forceps have long jaws and shorter shanks. Intestinal clamp, Doyen, straight, used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Jaws very springy and soft. Features: Non traumatic Occlusion: The form and mechanical action must permit precise closing. Multiple ratchet Box lock. Highly impact resistant. Length: approx 230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. The jaws must be meticulously cleaned. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revisions of the standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental: ISO 14001 Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices Supplied with: Manufacturer's instruction for use. Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used for atraumatic clamping/occlusion of a section intestine while suturing, resection or anastomosis. Product description: Doyen Forceps are a lightweight, non-crushing clamp. Commonly used to grasp large organs such as lungs, or intestines, the Doyen forceps have long jaws and shorter shanks. Intestinal clamp, Doyen, straight, used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Jaws very springy and soft. Features: Non traumatic Occlusion: The form and mechanical action must permit precise closing. Multiple ratchet Box lock. Highly impact resistant. Length: approx 230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. The jaws must be meticulously cleaned. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revisions of the standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental: ISO 14001 Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices Supplied with: Manufacturer's instruction for use. Instructions for use: Used for transverse, non-crushing occlusion of a section intestine while suturing an anastomosis. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Special instructions: Refer WHO guidance document 'Decontamination and reprocessing of medical devices for health-care facilities World Health Organization (‎2016)'‎ for reprocessing of reusable medical devices
370 | Medical & Surgical Instruments | CLAMPUMBILIC_100 | Clamp, umbilical, 5.2cm, sterile, single use, P100 | Clamp, umbilical, 5.2cm, sterile, single use, P100 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Box of 100 Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Box of 100 Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001
371 | Medical & Surgical Instruments | CLAMPUMBILICAL_1 | Clamp, umbilical, 5.2cm, sterile, single use, P1 | Clamp, umbilical, 5.2cm, sterile, single use, P1 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class Is Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Is device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001
372 | Medical & Surgical Instruments | CLAMPUMBILICAL_50 | Clamp, umbilical, 5.2cm, sterile, single use, P50 | Clamp, umbilical, 5.2cm, sterile, single use, P50 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Box of 50 Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to temporarily compress the umbilical cord immediately after birth. It is used before cutting or ligating the cord, and is intended to enable aseptic haemostasis. Product description: Disposable sterile umbilical cord clamp, safe, effective and easy to apply. Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. Grooves all along the length to prevent slip of the umbilical cord and to retain it in the same position. Finger grip to ensure safe and easy handling. Measurements: approx. 52 mm Material: Plastic (polymer) medical grade. DEHP free & latex free Single-use. Sterile Supplied with: Manufacturer's instruction for use Instructions for use: Disposable sterile umbilical cord clamp, safe, effective and easy to apply, designed for clamping the umbilical cord soon after the birt CE mark and reference number of notifying body. Packaging & Labelling: One (1) umbilical cord clamp in an individual sterilised peel pack. Box of 50 Symbols used according ISO 15223 and EN 15986:2011 CE mark and reference number of notifying body. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per the latest revisions of the standard): ISO 11135 - Sterilization of health-care products ISO 11607-1:2018, Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2019 Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process ISO 10993-7 Environmental requirements: ISO 14001
373 | Medical & Surgical Instruments | CRANIOCLASTBRAUN | Cranioclast, Braun, 420 mm | Cranioclast, Braun, 420 mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is used for crushing the foetal head, permitting vaginal extraction of the foetal parts of a dead or anomalous (abnormal) foetus Product description: Cranioclast, Braun, for gripping and decreasing the foetal skull in case of death in uterus, in order to extract it via the vagina. Pressure force instrument (crushing forceps) with adjustable screw clamp. Materials: Martensitic (quenched, magnetic steel). Length: 42cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to crush a foetal skull, permitting vaginal extraction of the foetal parts following a death in utero (embryotomy). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser.Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) cranioclast in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) Symbols used according ISO 15223 an EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per the latets revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is used for crushing the foetal head, permitting vaginal extraction of the foetal parts of a dead or anomalous (abnormal) foetus Product description: Cranioclast, Braun, for gripping and decreasing the foetal skull in case of death in uterus, in order to extract it via the vagina. Pressure force instrument (crushing forceps) with adjustable screw clamp. Materials: Martensitic (quenched, magnetic steel). Length: 42cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to crush a foetal skull, permitting vaginal extraction of the foetal parts following a death in utero (embryotomy). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser.Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) cranioclast in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) Symbols used according ISO 15223 an EN 980 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (as per the latets revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
374 | Medical & Surgical Instruments | CURETTE12MM | Curette, uterine, Sims, 12 mm, sharp | Curette, uterine, Sims, 12 mm, sharp | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 12mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 12mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
375 | Medical & Surgical Instruments | CURETTE7MM | Curette, uterine, Sims, 7 mm, sharp | Curette, uterine, Sims, 7 mm, sharp | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 7mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 7mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
376 | Medical & Surgical Instruments | CURETTE8MM | Curette, uterine, Sims, 8 mm, blunt | Curette, uterine, Sims, 8 mm, blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Blunt. Length: Shaft 260 mm. Width: Spoon 8mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus. Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Blunt. Length: Shaft 260 mm. Width: Spoon 8mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
377 | Medical & Surgical Instruments | CURETTE9MM | Curette, uterine, Sims, 9 mm, sharp | Curette, uterine, Sims, 9 mm, sharp | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 9mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Uterine curette for scrapping and removal of retained products of conception from inside the uterus Product description: Material: Martensitic steel (quenched, magnetic steel). Shank type: Rigid. Blade shape: Oval, fenestrated. Blade edge type: Sharp. Length: shaft approx. 260mm. Width: spoon approx. 9mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) curette in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
378 | Medical & Surgical Instruments | DECAPITATIONHOOK | Hook, decapitation, Braun, 31cm | Hook, decapitation, Braun, 31cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Obstetrical hook for the vaginal extraction of a dead foetus in utero (embryotomy) meant for use by experienced doctors. Product description: Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Length: approx.310 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to extract the foetus through the vagina in case of death in uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 . Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): Must comply with the following standards: ISO 7153-1:1991 ISO 17664:2004 ISO 13402:1995 ISO 7741:1986 BS 5194-4:1985 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Obstetrical hook for the vaginal extraction of a dead foetus in utero (embryotomy) meant for use by experienced doctors. Product description: Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Length: approx.310 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to extract the foetus through the vagina in case of death in uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 . Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards (must comply to the latest revision of standards): Must comply with the following standards: ISO 7153-1:1991 ISO 17664:2004 ISO 13402:1995 ISO 7741:1986 BS 5194-4:1985 Environmental requirements: ISO 14001
379 | Medical & Surgical Instruments | DILATORHEGAR | Dilator, uterine, Hegar, set of 16 dilators, size 3-18 | Dilator, uterine, Hegar, set of 16 dilators, size 3-18 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to dilate the cervical canal after its insertion through the cervical os to allow the passage of a curette or other instruments into the uterus and/or enlarge the uterus prior to surgical intervention by pressing on the inner walls of the uterus Product description: Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: From 17% to19 % chromium and From 8% to 10% nickel. Double ended. The dilator's distal end must be rounded, smooth and atraumatic Set of 8, size 3-18, Provides 16 sizes 3-18: 3/4; 5/6; 7/8; 9/10; 11/12; 13/14; 15/16; 17/18 Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: dilator single unit or a set of 8 (16 pieces), individually, in a plastic bag. Symbols used according ISO 15223 Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety and product standards: Must comply with the following standards: ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to dilate the cervical canal after its insertion through the cervical os to allow the passage of a curette or other instruments into the uterus and/or enlarge the uterus prior to surgical intervention by pressing on the inner walls of the uterus Product description: Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: From 17% to19 % chromium and From 8% to 10% nickel. Double ended. The dilator's distal end must be rounded, smooth and atraumatic Set of 8, size 3-18, Provides 16 sizes 3-18: 3/4; 5/6; 7/8; 9/10; 11/12; 13/14; 15/16; 17/18 Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For removal of retained products of conception from inside the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: dilator single unit or a set of 8 (16 pieces), individually, in a plastic bag. Symbols used according ISO 15223 Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety and product standards: Must comply with the following standards: ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental requirements: ISO 14001
380 | Medical & Surgical Instruments | FORCEPCHERON | Forceps, dressing, Cheron, 25cm | Forceps, dressing, Cheron, 25cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: primarily for non-dedicated grasping of devices, sponges and/or dressings during a procedure; it is not intended for tissue manipulation. Product description: Vaginal dressing forceps, Cheron, for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) Material: Martensitic stainless steel (quenched, magnetic steel. composition: 0.20% carbon; 13% chromium. Spring-type Flexible arms. Variable setting of ratchet, lockable. Adjustment of jaws. Highly impact resistant. Length: approx 250 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: primarily for non-dedicated grasping of devices, sponges and/or dressings during a procedure; it is not intended for tissue manipulation. Product description: Vaginal dressing forceps, Cheron, for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) Material: Martensitic stainless steel (quenched, magnetic steel. composition: 0.20% carbon; 13% chromium. Spring-type Flexible arms. Variable setting of ratchet, lockable. Adjustment of jaws. Highly impact resistant. Length: approx 250 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 7153-1:1991 ISO 7151:1988 ISO 17664:2004 ISO 13402:1995 Environmental requirements: ISO 14001
381 | Medical & Surgical Instruments | FORCEPDRESSING14.5 | Forceps, dressing, standard, 145 mm, straight | Forceps, dressing, standard, 145 mm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Dissecting forceps used in surgery for gripping and dissecting tissues, as well as for the coagulation of vessels, also used in nursing for making a compress into a tampon for sponging. Product description: Dissecting forceps, spring-type. Straight. Without teeth. Flexible arms. Serrated jaws Good jaw grips. Material: Martensitic steel (quenched, magnetic steel). Highly impact resistant. Length: approx. 145 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to hold tissues during dissections and for coagulation of vessels. Used in surgery and nursing. Used to change dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) forceps in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Dissecting forceps used in surgery for gripping and dissecting tissues, as well as for the coagulation of vessels, also used in nursing for making a compress into a tampon for sponging. Product description: Dissecting forceps, spring-type. Straight. Without teeth. Flexible arms. Serrated jaws Good jaw grips. Material: Martensitic steel (quenched, magnetic steel). Highly impact resistant. Length: approx. 145 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to hold tissues during dissections and for coagulation of vessels. Used in surgery and nursing. Used to change dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: One (1) forceps in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
382 | Medical & Surgical Instruments | FORCEPDUPLAY28CM | Forceps, uterine, Duplay, 28cm, curved | Forceps, uterine, Duplay, 28cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes Product description: Uterine traction/tenaculum/vulsellum forceps, Duplay, for gripping and immobilisation of thick tissue (uterine traction forceps). Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Double curved. Hard ratchet, lockable. Highly impact resistant Length: approx. 280 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: For gripping and immobilisation of thick tissue (uterine traction forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes Product description: Uterine traction/tenaculum/vulsellum forceps, Duplay, for gripping and immobilisation of thick tissue (uterine traction forceps). Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Double curved. Hard ratchet, lockable. Highly impact resistant Length: approx. 280 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: For gripping and immobilisation of thick tissue (uterine traction forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
383 | Medical & Surgical Instruments | FORCEPMAGILLCHILD | Forceps, catheter introducing, Magill, child | Forceps, catheter introducing, Magill, child | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: angled forceps used to guide the endotracheal tube into the laryngeal inlet during nasotracheal intubation or a nasogastric tube into the esophagus under direct vision. It is also used to place pharyngeal packs and removal of foreign bodies. Product description: Catheter forceps, Magill, scissors-like design with ring handles. The distal working end has grasping blades that have small ringed loops. The blades are serrated to provide extra grip. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Flexible arms. With two non-detachable legs curved through an angle of about 30 degrees, ending in loops. The inside of each loop (spatula) is ribbed. Length: 190 – 200 approx.mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: For guiding an endotracheal tube into the trachea or an aspirating/feeding tube into the oesophagus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps catheter in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: angled forceps used to guide the endotracheal tube into the laryngeal inlet during nasotracheal intubation or a nasogastric tube into the esophagus under direct vision. It is also used to place pharyngeal packs and removal of foreign bodies. Product description: Catheter forceps, Magill, scissors-like design with ring handles. The distal working end has grasping blades that have small ringed loops. The blades are serrated to provide extra grip. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Flexible arms. With two non-detachable legs curved through an angle of about 30 degrees, ending in loops. The inside of each loop (spatula) is ribbed. Length: 190 – 200 approx.mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: For guiding an endotracheal tube into the trachea or an aspirating/feeding tube into the oesophagus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps catheter in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
384 | Medical & Surgical Instruments | FORCEPOBSTETRIC | Forceps, obstetric, Simpson-Braun, 35 cm | Forceps, obstetric, Simpson-Braun, 35 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Obstetric forceps, Simpson-Braun, to grasp and apply traction to the foetal head in the birth passage and facilitate delivery. Product description: Obstetrical forceps with long, curved blades with fenestrations Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Highly impact resistant. Length: approx 350 mm Fenestrated blades for minimal trauma. Lock System to achieve Self-Retaining clamping. Grooved handles for secure grip. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used to grasp and apply traction to the foetal head in the birth passage and facilitate delivery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Obstetric forceps, Simpson-Braun, to grasp and apply traction to the foetal head in the birth passage and facilitate delivery. Product description: Obstetrical forceps with long, curved blades with fenestrations Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Highly impact resistant. Length: approx 350 mm Fenestrated blades for minimal trauma. Lock System to achieve Self-Retaining clamping. Grooved handles for secure grip. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used to grasp and apply traction to the foetal head in the birth passage and facilitate delivery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
385 | Medical & Surgical Instruments | FORCEPPHANEUF | Forceps, uterine, Phaneuf, 21.5 cm, curved | Forceps, uterine, Phaneuf, 21.5 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For haemostasis of the arteries, especially uterine arteries Product description: Hemostatic forceps for haemostasis of the arteries in the uterine, curved. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium Double curved. Slightly springy. Flexible arms. Variable setting of ratchet (lockable). Adjustment of the jaws. Curved, angled on flat forceps with 1 x 2 teeth. Length: 215mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For haemostasis of the arteries in the uterine,. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For haemostasis of the arteries, especially uterine arteries Product description: Hemostatic forceps for haemostasis of the arteries in the uterine, curved. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium Double curved. Slightly springy. Flexible arms. Variable setting of ratchet (lockable). Adjustment of the jaws. Curved, angled on flat forceps with 1 x 2 teeth. Length: 215mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For haemostasis of the arteries in the uterine,. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
386 | Medical & Surgical Instruments | FORCEPROCHEST20 | Forceps, artery, Rochester-Pean, 20 cm, curved | Forceps, artery, Rochester-Pean, 20 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.200 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.200 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
387 | Medical & Surgical Instruments | FORCEPROCHEST22 | Forceps, artery, Rochester-Pean, 22cm, straight | Forceps, artery, Rochester-Pean, 22cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Haemostatic, forceps, staight, used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. Straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic with serrated jaws. Adjustment of jaws Without teeth. Highly impact resistant. Length: aprprox.220 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Haemostatic, forceps, staight, used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. Straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic with serrated jaws. Adjustment of jaws Without teeth. Highly impact resistant. Length: aprprox.220 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
388 | Medical & Surgical Instruments | FORCEPROCHEST24C | Forceps, artery, Rochester-Pean, 24 cm, curved | Forceps, artery, Rochester-Pean, 24 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.240 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.240 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
389 | Medical & Surgical Instruments | FORCEPROCHESTER | Forceps, artery, Rochester-Pean, 16 cm, curved | Forceps, artery, Rochester-Pean, 16 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.160 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Hemostatic forceps, Rochester-Pean, curved for haemostasis, gripping, dissection and tampon holder. The full horizontal serrations and availability in multiple lengths make these forceps a versatile instrument used in multiple procedures. Curved Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic jaws with full horizontal jaw serrations Adjustment of jaws Without teeth. Highly impact resistant. Length: approx.160 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for for haemostasis, gripping, dissection and tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
390 | Medical & Surgical Instruments | FORCEPROCHESTERST | Forceps, artery, Rochester-Pean, 16 cm, straight | Forceps, artery, Rochester-Pean, 16 cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Haemostatic, forceps, staight, used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. Straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic with serrated jaws. Adjustment of jaws Without teeth. Highly impact resistant. Length: aprprox.160 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Rochester-Pean haemostatic forceps are used for clamping larger tissue and vessels for haemostasis. Product description: Haemostatic, forceps, staight, used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. Straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy locking Variable setting of ratchet, lockable Non traumatic with serrated jaws. Adjustment of jaws Without teeth. Highly impact resistant. Length: aprprox.160 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Used for gripping, haemostasis and the dissection of deep structures: abdominal and thoracic surgeries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
391 | Medical & Surgical Instruments | FORCEPSBENGOLEA20 | Forceps, artery, Bengolea, 20cm, curved | Forceps, artery, Bengolea, 20cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Hemostatic forceps used in deep surgery. Product description: Hemostatic forceps, Bengolea, used in deep surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved and slim. Spring-type Serrated or cross jaws, gripping the thread well. Flexible arms. Progressive setting and locking the ratchet. Highly impact resistant. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Bengolea, used in deep surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark ,Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Hemostatic forceps used in deep surgery. Product description: Hemostatic forceps, Bengolea, used in deep surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved and slim. Spring-type Serrated or cross jaws, gripping the thread well. Flexible arms. Progressive setting and locking the ratchet. Highly impact resistant. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Bengolea, used in deep surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark ,Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
392 | Medical & Surgical Instruments | FORCEPSBIERER16MM | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 16mm, S | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 16mm, S | Intended purpose of use: for properly removing pieces of the placenta or products of conception from the uterus to prevent infection Product Description: Bierer ovum forceps are for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. Serrated and fenestrated blades provide a firm grasp. Material: Martensitic stainless steel (quenched magnetic steel) 0.20% carbon, 13% Chromium. Gripping forceps with fully serrated jaws (with large 'teeth'). Fenestrated oval jaws. Slightly curved. Ring handles. Non-ratcheted, not lockable. Highly impact resistant. Length: 330 mm (+/- 5%) Size: Small Jaws: 16 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instructions for use, cleaning, disinfection and sterilization in English, French and Spanish. Intended use: Used for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. This item must be cleaned and disinfected after each use, and sterilised in a steam steriliser. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | Intended purpose of use: for properly removing pieces of the placenta or products of conception from the uterus to prevent infection Product Description: Bierer ovum forceps are for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. Serrated and fenestrated blades provide a firm grasp. Material: Martensitic stainless steel (quenched magnetic steel) 0.20% carbon, 13% Chromium. Gripping forceps with fully serrated jaws (with large 'teeth'). Fenestrated oval jaws. Slightly curved. Ring handles. Non-ratcheted, not lockable. Highly impact resistant. Length: 330 mm (+/- 5%) Size: Small Jaws: 16 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instructions for use, cleaning, disinfection and sterilization in English, French and Spanish. Intended use: Used for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. This item must be cleaned and disinfected after each use, and sterilised in a steam steriliser. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
393 | Medical & Surgical Instruments | FORCEPSBIERER19MM | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 19mm, L | Forceps, uterine evacuation, ovum, Bierer, 33cm, jaws 19mm, L | Intended purpose of use: for properly removing pieces of the placenta or products of conception from the uterus to prevent infection Product Description: Bierer ovum forceps are for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. Serrated and fenestrated blades provide a firm grasp. Material: Martensitic stainless steel (quenched magnetic steel) 0.20% carbon, 13% Chromium. Gripping forceps with fully serrated jaws (with large 'teeth'). Fenestrated oval jaws. Slightly curved. Ring handles. Non-ratcheted, not lockable. Highly impact resistant. Length: 330 mm (+/- 5%) Size: Small Jaws: 19 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instructions for use, cleaning, disinfection and sterilization in English, French and Spanish. Intended use: Used for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. This item must be cleaned and disinfected after each use, and sterilised in a steam steriliser. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | Intended purpose of use: for properly removing pieces of the placenta or products of conception from the uterus to prevent infection Product Description: Bierer ovum forceps are for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. Serrated and fenestrated blades provide a firm grasp. Material: Martensitic stainless steel (quenched magnetic steel) 0.20% carbon, 13% Chromium. Gripping forceps with fully serrated jaws (with large 'teeth'). Fenestrated oval jaws. Slightly curved. Ring handles. Non-ratcheted, not lockable. Highly impact resistant. Length: 330 mm (+/- 5%) Size: Small Jaws: 19 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instructions for use, cleaning, disinfection and sterilization in English, French and Spanish. Intended use: Used for obstetrics / gynecological procedures such as removal of placenta fragments from the uterus. This item must be cleaned and disinfected after each use, and sterilised in a steam steriliser. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel SO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
394 | Medical & Surgical Instruments | FORCEPSCHEATLE | Forceps, sterilizer, Cheatle, 27 cm, curved | Forceps, sterilizer, Cheatle, 27 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used to transfer either sterilized instruments or packing materials to the surgeon’s hands. Product description: Grasping forceps that the operator uses for clamping, holding and mobilizing sterilized devices during surgical procedures in all areas and also used to take a sterile items out of drums. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Length: 270 approx. mm Slightly curved Jaws for cecurely carrying instruments. Finger Rings Ensuring Maximum Control Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place or remove instruments from a sterilizer. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used to transfer either sterilized instruments or packing materials to the surgeon’s hands. Product description: Grasping forceps that the operator uses for clamping, holding and mobilizing sterilized devices during surgical procedures in all areas and also used to take a sterile items out of drums. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Length: 270 approx. mm Slightly curved Jaws for cecurely carrying instruments. Finger Rings Ensuring Maximum Control Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place or remove instruments from a sterilizer. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
395 | Medical & Surgical Instruments | FORCEPSCRAFOORD24 | Forceps, artery, Crafoord, 24cm, curved | Forceps, artery, Crafoord, 24cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to provide an atraumatic way to clamp blood vessels, hold needles and grasp tissues. Used for haemostasis in deep surgery Product description: Hemostatic forceps, Crafoord, also be known as a Crafoord Coarctation Clamp, may also be identified by the name Coller. Not intended for nursing care. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Spring-type slightly curved jaws with inner serrations Flexible arms with finger ring handle Progressive setting of the ratchet. Variable settings of the ratchet, lockable. Highly impact resistant. Length: approx. 240mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Hemostatic forceps, Craford, used for deep haemostasis, not intended for nursing care This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to provide an atraumatic way to clamp blood vessels, hold needles and grasp tissues. Used for haemostasis in deep surgery Product description: Hemostatic forceps, Crafoord, also be known as a Crafoord Coarctation Clamp, may also be identified by the name Coller. Not intended for nursing care. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Spring-type slightly curved jaws with inner serrations Flexible arms with finger ring handle Progressive setting of the ratchet. Variable settings of the ratchet, lockable. Highly impact resistant. Length: approx. 240mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Hemostatic forceps, Craford, used for deep haemostasis, not intended for nursing care This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
396 | Medical & Surgical Instruments | FORCEPSDRESSING20 | Forceps, dressing, standard, fine serrated jaws, 20cm | Forceps, dressing, standard, fine serrated jaws, 20cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping and dissecting tissues, sometimes used for the coagulation of vessels, also used in nursing for dressing change. Product description: Dissecting forceps, springy. Fine serrated jaw Flexible arms. Good gripping of the jaws. Without teeth Length: 18 or 20cm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Dissecting forceps used to hold tissues during dissections and for coagualation of vessels This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping and dissecting tissues, sometimes used for the coagulation of vessels, also used in nursing for dressing change. Product description: Dissecting forceps, springy. Fine serrated jaw Flexible arms. Good gripping of the jaws. Without teeth Length: 18 or 20cm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Dissecting forceps used to hold tissues during dissections and for coagualation of vessels This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
397 | Medical & Surgical Instruments | FORCEPSDRESSING25 | Forceps, dressing, standard, 250 mm, straight | Forceps, dressing, standard, 250 mm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping and dissecting tissues, sometimes used for the coagulation of vessels, also used in nursing for dressing change. Product description: Dissecting forceps, springy. Fine serrated jaw Flexible arms. Good gripping of the jaws. Without teeth Length: 25cm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Dissecting forceps used to hold tissues during dissections and for coagualation of vessels This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping and dissecting tissues, sometimes used for the coagulation of vessels, also used in nursing for dressing change. Product description: Dissecting forceps, springy. Fine serrated jaw Flexible arms. Good gripping of the jaws. Without teeth Length: 25cm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Dissecting forceps used to hold tissues during dissections and for coagualation of vessels This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
398 | Medical & Surgical Instruments | FORCEPSKELLY | Forceps, artery, Kelly, 14cm, curved | Forceps, artery, Kelly, 14cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for clamping large blood vessels or manipulating heavy tissues. They may also be used for soft tissue dissection. Not to be used for nusrsing care to handle compress and make dressings. Product description: Small curved haemostatic forceps Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Flexible arms. Variable settings of the ratchet, lockable. Inner serration with atraumatic jaws Adjustment of the jaws. Length: approx.140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kelly , used for haemostasis. Do not use these forceps for nursing care. To grip compresses and make dressings, use the appropriate forceps (standard dressing forceps or Cheron dressing forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for clamping large blood vessels or manipulating heavy tissues. They may also be used for soft tissue dissection. Not to be used for nusrsing care to handle compress and make dressings. Product description: Small curved haemostatic forceps Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Curved. Flexible arms. Variable settings of the ratchet, lockable. Inner serration with atraumatic jaws Adjustment of the jaws. Length: approx.140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kelly , used for haemostasis. Do not use these forceps for nursing care. To grip compresses and make dressings, use the appropriate forceps (standard dressing forceps or Cheron dressing forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
399 | Medical & Surgical Instruments | FORCEPSKOCHER140 | Forceps, artery, Kocher, 140 mm, 1x2 teeth, curved | Forceps, artery, Kocher, 140 mm, 1x2 teeth, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: provide an atraumatic way to control haemostasis during surgery. Multi-purpose in surgery: haemostasis, dissection, gripping of vessels Product description:Haemostatic forceps, Kocher Curved. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly springy. Flexible arms. Variable setting of the ratchet. Adjustment of the jaws. 1x2 teeth. Heavily serrated jaws. Box lock. Highly impact resistant. Length: approx.140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kocher, used for for general use: haemostasis, gripping, dissection, tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: provide an atraumatic way to control haemostasis during surgery. Multi-purpose in surgery: haemostasis, dissection, gripping of vessels Product description:Haemostatic forceps, Kocher Curved. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly springy. Flexible arms. Variable setting of the ratchet. Adjustment of the jaws. 1x2 teeth. Heavily serrated jaws. Box lock. Highly impact resistant. Length: approx.140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kocher, used for for general use: haemostasis, gripping, dissection, tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
400 | Medical & Surgical Instruments | FORCEPSKOCHERSTR | Forceps, artery, Kocher, 14cm, straight | Forceps, artery, Kocher, 14cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: provide an atraumatic way to control haemostasis during surgery. Multi-purpose in surgery: haemostasis, dissection, gripping of vessels Product description: Haemostatic forceps, Kocher, straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly springy. Flexible arms. Variable setting of the ratchet. Adjustment of the jaws. 1x2 teeth. Heavily serrated jaws. Box lock. Highly impact resistant. Length: approx. 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kocher, used for for general use: haemostasis, gripping, dissection, tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: provide an atraumatic way to control haemostasis during surgery. Multi-purpose in surgery: haemostasis, dissection, gripping of vessels Product description: Haemostatic forceps, Kocher, straight Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly springy. Flexible arms. Variable setting of the ratchet. Adjustment of the jaws. 1x2 teeth. Heavily serrated jaws. Box lock. Highly impact resistant. Length: approx. 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Kocher, used for for general use: haemostasis, gripping, dissection, tampon holder. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
401 | Medical & Surgical Instruments | FORCEPSMIXTER14 | Forceps, artery, Mixter, 14cm | Forceps, artery, Mixter, 14cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for clamping, dissection, or grasping tissue. ​Used for dissecting and for passing a thread round a vein or an artery. Product description: Very delicate threading forceps with serrations on the entire length of the jaw Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy Flexible ratchet, lockable. Jaws which grip the thread well Highly impact resistant. Length: approx. 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for clamping, dissection, or grasping tissue. ​Used for dissecting and for passing a thread round a vein or an artery. Product description: Very delicate threading forceps with serrations on the entire length of the jaw Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy Flexible ratchet, lockable. Jaws which grip the thread well Highly impact resistant. Length: approx. 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
402 | Medical & Surgical Instruments | FORCEPSMIXTER23 | Forceps, artery, Mixter, 23cm | Forceps, artery, Mixter, 23cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for clamping, dissection, or grasping tissue. ​Used for dissecting and for passing a thread round a vein or an artery. Product description: Very delicate threading forceps with serrations on the entire length of the jaw Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy Flexible ratchet, lockable. Jaws which grip the thread well Highly impact resistant. Length: approx. 230 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for clamping, dissection, or grasping tissue. ​Used for dissecting and for passing a thread round a vein or an artery. Product description: Very delicate threading forceps with serrations on the entire length of the jaw Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Springy Flexible ratchet, lockable. Jaws which grip the thread well Highly impact resistant. Length: approx. 230 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
403 | Medical & Surgical Instruments | FORCEPSMOSQUITO | Forceps, artery, Halsted-Mosquito, 12.5cm, curved | Forceps, artery, Halsted-Mosquito, 12.5cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:To clamp blood vessels, especially small arteries, during a variety of surgical procedures. The instrument is also useful to grasp disposable materials, tissues and sutures. Product description: Haemostatic forceps, springy, non traumatic jaws. Flexible arms. Variable setting of the ratchet, lockable. Slim, curved forceps without teeth. Length: 125mm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Halsted-Mosquito, used for haemostasis. Not intended for nursing care. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:To clamp blood vessels, especially small arteries, during a variety of surgical procedures. The instrument is also useful to grasp disposable materials, tissues and sutures. Product description: Haemostatic forceps, springy, non traumatic jaws. Flexible arms. Variable setting of the ratchet, lockable. Slim, curved forceps without teeth. Length: 125mm Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Hemostatic forceps, Halsted-Mosquito, used for haemostasis. Not intended for nursing care. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: forceps single unit, in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
404 | Medical & Surgical Instruments | FORCEPSMUSEUX24CM | Forceps, uterine, Museux, 24cm, curved | Forceps, uterine, Museux, 24cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for grasping the uterus during obstetrical/gynaecological procedures. Product description: uterine traction forceps also known as Tenaculum forceps or Forceps uterine vulsellum Museux Uterine forceps for gripping and immobilising thick tissue, curved sideways. Material: Martensitic stainless steel (quenched, magnetic steel) Martensitic stainless steel composition: 0.20% carbon; 13% chromium. Gripping forceps with teeth: 2 x 2 teeth. curved sideways. Precise adjustment of the teeth. Hard ratchet, lockable. Highly impact resistant. Length: 240mm. Jaws: 10mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip and to immobilise thick tissue(uterine traction forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for grasping the uterus during obstetrical/gynaecological procedures. Product description: uterine traction forceps also known as Tenaculum forceps or Forceps uterine vulsellum Museux Uterine forceps for gripping and immobilising thick tissue, curved sideways. Material: Martensitic stainless steel (quenched, magnetic steel) Martensitic stainless steel composition: 0.20% carbon; 13% chromium. Gripping forceps with teeth: 2 x 2 teeth. curved sideways. Precise adjustment of the teeth. Hard ratchet, lockable. Highly impact resistant. Length: 240mm. Jaws: 10mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip and to immobilise thick tissue(uterine traction forceps). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
405 | Medical & Surgical Instruments | FORCEPSPENCERWEL | Forceps, artery, Spencer-Wells, 20.5 cm, curved | Forceps, artery, Spencer-Wells, 20.5 cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to manipulate blood vessels, clamp their lumen and control the blood flow, in order to achieve adequate haemostasis Product description: Spencer-Wells Artery forceps, curved, without teeth, lockable. Measurements: 20.5cm Material: Martensitic stainless steel (quenched, magnetic steel) Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to manipulate blood vessels, clamp their lumen and control the blood flow, in order to achieve adequate haemostasis Product description: Spencer-Wells Artery forceps, curved, without teeth, lockable. Measurements: 20.5cm Material: Martensitic stainless steel (quenched, magnetic steel) Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For dissecting and for passing a thread around a vein or artery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
406 | Medical & Surgical Instruments | FORCEPSPERITONEAL | Forceps, peritoneal, Faure, 21cm, slightly curved | Forceps, peritoneal, Faure, 21cm, slightly curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to grasp, mobilize and statically clamp soft tissues and peritoneum, especially during abdominal surgical procedures. Product description: Haemostatic forceps. Used for haemostasis of uterine arteries. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly curved haemostatic forceps with 1 x 2 teeth Flexible arms Soft ratchet, lockable. Highly impact resistant. Length: approx 210 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for haemostasis of uterine arteries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to grasp, mobilize and statically clamp soft tissues and peritoneum, especially during abdominal surgical procedures. Product description: Haemostatic forceps. Used for haemostasis of uterine arteries. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Slightly curved haemostatic forceps with 1 x 2 teeth Flexible arms Soft ratchet, lockable. Highly impact resistant. Length: approx 210 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for haemostasis of uterine arteries. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
407 | Medical & Surgical Instruments | FORCEPSSPONGE20 | Forceps, sponge-holding, Forester, 20 cm, straight | Forceps, sponge-holding, Foerster, 20 cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: use for grasping, holding and mobilizing surgical dressings Product description: Sponge-holding forceps, Forester, straight, used to hold tampons and gauze compresses. Also uses in gallbladder surgery Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Straight arms Fenestrated round Jaws with Serrations Multiple ratchets, box lock. Length: approx.200 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to hold tampons and gauze compresses. Also uses in gallbladder surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: use for grasping, holding and mobilizing surgical dressings Product description: Sponge-holding forceps, Foerster, straight, used to hold tampons and gauze compresses. Also uses in gallbladder surgery Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Straight arms Fenestrated round Jaws with Serrations Multiple ratchets, box lock. Length: approx.200 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to hold tampons and gauze compresses. Also uses in gallbladder surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
408 | Medical & Surgical Instruments | FORCEPSSPONGE24 | Forceps, sponge-holding, Forester, 24 cm, straight | Forceps, sponge-holding, Foerster, 24 cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: use for grasping, holding and mobilizing surgical dressings Product description: Sponge-holding forceps, Forester, straight, used to hold tampons and gauze compresses. Also uses in gallbladder surgery Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Straight arms Fenestrated round Jaws with Serrations Multiple ratchets, box lock. Length: approx.240 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to hold tampons and gauze compresses. Also uses in gallbladder surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: use for grasping, holding and mobilizing surgical dressings Product description: Sponge-holding forceps, Foerster, straight, used to hold tampons and gauze compresses. Also uses in gallbladder surgery Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Straight arms Fenestrated round Jaws with Serrations Multiple ratchets, box lock. Length: approx.240 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to hold tampons and gauze compresses. Also uses in gallbladder surgery This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
409 | Medical & Surgical Instruments | FORCEPSTENACULUM | Forceps, tenaculum, vulsellum, Schröder, straight, 25cm | Forceps, tenaculum, vulsellum, Schröder, straight, 25cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A specialized obstetrical/gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes. Its principal use is to retract tissues for performing tubal ligation, hysterectomies and other procedures Product description: forceps for manipulation of the cervix or uterus, and vaginal or abdominal hysterectomy. Straight gynaecological Schröder forceps with one hook at the tip of each blade. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. (AISI 410/420) Single flat tip/tooth at the end of each blade (Braun type), also called 1x1 prongs. Prong tips angle inward at 90 degrees. Lockable ratchet with variable settings Straight Length: 250 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use (sterilization, cleaning, maintenance). Instructions for use: Gynaecological forceps used for fistula repair. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A specialized obstetrical/gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes. Its principal use is to retract tissues for performing tubal ligation, hysterectomies and other procedures Product description: forceps for manipulation of the cervix or uterus, and vaginal or abdominal hysterectomy. Straight gynaecological Schröder forceps with one hook at the tip of each blade. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. (AISI 410/420) Single flat tip/tooth at the end of each blade (Braun type), also called 1x1 prongs. Prong tips angle inward at 90 degrees. Lockable ratchet with variable settings Straight Length: 250 mm (+/- 5%) Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use (sterilization, cleaning, maintenance). Instructions for use: Gynaecological forceps used for fistula repair. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
410 | Medical & Surgical Instruments | FORCEPSTISS20_3X4 | Forceps, tissue, Allis, 3x4 teeth, 20 cm | Forceps, tissue, Allis, 3x4 teeth, 20 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 3 x 4 teeth. Atraumatic style jaws Precise adjustment of the teeth. Flexible ratchet, lockable Lockable at different positions Length: 20cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 3 x 4 teeth. Atraumatic style jaws Precise adjustment of the teeth. Flexible ratchet, lockable Lockable at different positions Length: 20cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
411 | Medical & Surgical Instruments | FORCEPSTISSUE14.5 | Forceps, tissue, standard, 14.5 cm, straight | Forceps, tissue, standard, 14.5 cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, spring-type. Straight. With 1 x 2 teeth. Flexible arms. Good adjustment of the teeth. Good jaw grips. Material: Martensitic steel (quenched, magnetic steel). Highly impact resistant. Length: approx. 145 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used for gripping, dissecting tissue and coagulation of vessels. Used in surgery and nursing. Forceps with teeth are used for dissecting thick tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use Packaging and labelling: One (1) forceps in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, spring-type. Straight. With 1 x 2 teeth. Flexible arms. Good adjustment of the teeth. Good jaw grips. Material: Martensitic steel (quenched, magnetic steel). Highly impact resistant. Length: approx. 145 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used for gripping, dissecting tissue and coagulation of vessels. Used in surgery and nursing. Forceps with teeth are used for dissecting thick tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use Packaging and labelling: One (1) forceps in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Symbols used according ISO 15223 and EN 980 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
412 | Medical & Surgical Instruments | FORCEPSTISSUE20 | Forceps, tissue, Allis, 4x5 teeth, 20 cm | Forceps, tissue, Allis, 4x5 teeth, 20 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 4X5 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 200mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 4X5 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 200mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
413 | Medical & Surgical Instruments | FORCEPSTISSUE20CM | Forceps, tissue, Babcock, 20 cm | Forceps, tissue, Babcock, 20 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping soft and delicate tissues (lungs and intestines) and organs. Product description: Grasping Forceps, Babcock, Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Springy High impact resistant Flexible arms. Soft ratchet, lockable. Pronounced but non traumatic ridges of the grippers Ridged grippers with aperture (triangular). Length: approx. 200 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft and delicate tissues (lungs and intestines) and organs. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for gripping soft and delicate tissues (lungs and intestines) and organs. Product description: Grasping Forceps, Babcock, Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Springy High impact resistant Flexible arms. Soft ratchet, lockable. Pronounced but non traumatic ridges of the grippers Ridged grippers with aperture (triangular). Length: approx. 200 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft and delicate tissues (lungs and intestines) and organs. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark , Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
414 | Medical & Surgical Instruments | FORCEPSTISSUE20S | Forceps,tissue,standard,dissecting,1x2 teeth,20cm, straight | Forceps,tissue,standard,dissecting,1x2 teeth,20cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, springy. Straight with 1 x 2 teeth. With fine teeth Flexible arms. Good adjustment of the teeth Light weight Good gripping of the jaws. Length: 20cm Material: Martensitic steel (quenched, magnetic steel), fine. Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, springy. Straight with 1 x 2 teeth. With fine teeth Flexible arms. Good adjustment of the teeth Light weight Good gripping of the jaws. Length: 20cm Material: Martensitic steel (quenched, magnetic steel), fine. Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
415 | Medical & Surgical Instruments | FORCEPSTISSUE25 | Forceps, tissue, standard, 250 mm, straight | Forceps, tissue, standard, 250 mm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, springy. Straight with 1 x 2 teeth. With fine teeth Flexible arms. Good adjustment of the teeth Light weight Good gripping of the jaws. Length: 250mm Material: Martensitic steel (quenched, magnetic steel), fine. Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in deep surgery for gripping and dissecting tissues, as well as for the coagulation of vessels. Product description: Dissecting forceps, springy. Straight with 1 x 2 teeth. With fine teeth Flexible arms. Good adjustment of the teeth Light weight Good gripping of the jaws. Length: 250mm Material: Martensitic steel (quenched, magnetic steel), fine. Martensitic steel composition: 0.20% carbon; 13% chromium; 1% silicon ; 1% maganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
416 | Medical & Surgical Instruments | FORCEPSVULSELLUM | Forceps, Schroeder, Vulsellum, 25 cm, 2x2 prongs, curved | Forceps, Schroeder, Vulsellum, 25 cm, 2x2 prongs, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A specialized obstetrical/gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes. Its principal use is to retract tissues for performing tubal ligation, hysterectomies and other procedures Product description: Gynaecological forceps with hooks at the tip of each blade. Instrument used for fistula repair. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. 2X2 inward pointed prongs Lockable ratchet with variable settings Length: approx. 250 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use (sterilization, cleaning, maintenance). Instructions for use: Gynaecological forceps used for fistula repair. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: A specialized obstetrical/gynaecological grasping forceps that surgeons use for holding multiple structures, such as the cervix, uterus, blood vessels and the fallopian tubes. Its principal use is to retract tissues for performing tubal ligation, hysterectomies and other procedures Product description: Gynaecological forceps with hooks at the tip of each blade. Instrument used for fistula repair. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. 2X2 inward pointed prongs Lockable ratchet with variable settings Length: approx. 250 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use (sterilization, cleaning, maintenance). Instructions for use: Gynaecological forceps used for fistula repair. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
417 | Medical & Surgical Instruments | FORCEPTISSUE15 | Forceps, tissue, Allis, 4x5 teeth, 15 cm | Forceps, tissue, Allis, 4x5 teeth, 15 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 4X5 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 150mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 4X5 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 150mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
418 | Medical & Surgical Instruments | FORCEPTISSUE15_3X4 | Forceps, tissue, Allis, 3x4 teeth, 15 cm | Forceps, tissue, Allis, 3x4 teeth, 15 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 3x4 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 150mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping organs and slippery or dense tissue. They are commonly used with intestines, vaginal, breast, and thyroid tissue. Product description: Gripping forceps. Straight with 3x4 teeth. Atraumatic style jaws Precise adjustment of the teeth. hard ratchet, lockable Length: 150mm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium ; 1% silicon ; 1% manganese Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
419 | Medical & Surgical Instruments | FORCEPTISSUECOLIN | Forceps, tissue, Colin, 16 cm | Forceps, tissue, Colin, 16 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for gripping soft and delicate tissues (lungs and intestines), as well as tumors. Can also be used for making a compress into a tampon for sponging. Product description: Grasping forceps Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Very large, flexible arms Soft ratchet with locking Jaws: oval shaped fenestrated jaws with strong but atraumatic serrations, Width of the jaws: 21 mm Springy. Length: 160 approx. mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft and delicate tissues (intestines) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for gripping soft and delicate tissues (lungs and intestines), as well as tumors. Can also be used for making a compress into a tampon for sponging. Product description: Grasping forceps Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Very large, flexible arms Soft ratchet with locking Jaws: oval shaped fenestrated jaws with strong but atraumatic serrations, Width of the jaws: 21 mm Springy. Length: 160 approx. mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip soft and delicate tissues (intestines) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
420 | Medical & Surgical Instruments | FORCEPTISSUEDUVAL | Forceps, tissue, Duval, 23cm | Forceps, tissue, Duval, 23cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping lung tissue and other tissues in the stomach or intestines. Can also be used for making a compress into a tampon for sponging. Product description: Grasping forceps, with serrated tips and a ratcheted handle. Material: Martensitic stainless steel (quenched, magnetic steel) Straight with triangular, fenestrated, serrated tips Non-traumatic parallel serrations. Flexible arms. Soft ratchet, lockable. Pronounced but non traumatic ridges of the grippers. Ridged grippers with aperture (triangular). Highly impact resistant Length: approx. 230 mm. Jaw: approximately 27mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip and hold bulky, soft and delicate tissues (lung and intestines). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for grasping lung tissue and other tissues in the stomach or intestines. Can also be used for making a compress into a tampon for sponging. Product description: Grasping forceps, with serrated tips and a ratcheted handle. Material: Martensitic stainless steel (quenched, magnetic steel) Straight with triangular, fenestrated, serrated tips Non-traumatic parallel serrations. Flexible arms. Soft ratchet, lockable. Pronounced but non traumatic ridges of the grippers. Ridged grippers with aperture (triangular). Highly impact resistant Length: approx. 230 mm. Jaw: approximately 27mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to grip and hold bulky, soft and delicate tissues (lung and intestines). This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
421 | Medical & Surgical Instruments | HOOKFISTULA | Hook, fistula, Emmet, 22 cm | Hook, fistula, Emmet, 22 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Grips soft tissue and secures it in place in order to perform different surgeries such as cervical tear repair and cervical biopsy. Product description: A shaft-like instrument with a sharp single hook at the distal end and a handle at the proximal end. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Pointed hook, Sharp Tip For Secure Grip On Tissues Length: approx.220 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: used to retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Grips soft tissue and secures it in place in order to perform different surgeries such as cervical tear repair and cervical biopsy. Product description: A shaft-like instrument with a sharp single hook at the distal end and a handle at the proximal end. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium. Pointed hook, Sharp Tip For Secure Grip On Tissues Length: approx.220 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: used to retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
422 | Medical & Surgical Instruments | MVAACCESSORYKIT | Accesory kit for MVA, including (1) O-ring, (1) collar stop, (1), cap, (1) 2cc silicone | Accesory kit for MVA, including (1) O-ring, (1) collar stop, (1), cap, (1) 2cc silicone | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Accessories for MVA. Product description: A package of items to be used as spare parts and lubricant for MVA. The kit includes: - Plunger O-ring , spare part to replace damaged O-ring - Collar stop, spare part to replace damaged collar stop - Cap, spare part to replace damaged cap - Silicone tube ( 2 cc), used to lubricate the O-ring Material: O-ring: synthetic rubber latex free Collar stop and cap: Latex-free plastic Lubricant: medical grade silicone Non sterile Instructions for use: Caution: Excessive lubrication can cause the aspirator to lose vacuum. Do not over lubricate the plunger O-ring. Do not lubricate other parts of the aspirator. To be sterilized prior to initial use and between patients. Autoclavable at 121ºC Supplied with: - 1x Plunger O-ring - 1x collar Stop - 1x cap - 1x silicone tube ( 2 cc) - 1x Instruction for use Packaging and labelling: One (1) MVA accessory kit in a plastic bag with manufacturer's instruction for user in English. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Accessories for MVA. Product description: A package of items to be used as spare parts and lubricant for MVA. The kit includes: - Plunger O-ring , spare part to replace damaged O-ring - Collar stop, spare part to replace damaged collar stop - Cap, spare part to replace damaged cap - Silicone tube ( 2 cc), used to lubricate the O-ring Material: O-ring: synthetic rubber latex free Collar stop and cap: Latex-free plastic Lubricant: medical grade silicone Non sterile Instructions for use: Caution: Excessive lubrication can cause the aspirator to lose vacuum. Do not over lubricate the plunger O-ring. Do not lubricate other parts of the aspirator. To be sterilized prior to initial use and between patients. Autoclavable at 121ºC Supplied with: - 1x Plunger O-ring - 1x collar Stop - 1x cap - 1x silicone tube ( 2 cc) - 1x Instruction for use Packaging and labelling: One (1) MVA accessory kit in a plastic bag with manufacturer's instruction for user in English. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer. Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Symbols used according ISO 15223 and EN 980 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO14001
423 | Medical & Surgical Instruments | MVACANNULAE10MM | Cannulae, Easygrip, 10mm, integrated base, SR10 | Cannulae, Easygrip, 10mm, integrated base, SR10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications of MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of Translucent polypropylene (PP) or similar ready to use and sterile ( by ETO method) diameter : 10 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Material: medical grade , polypropylene plastic, latex free, translucent If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. autoclavable at 121°C Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: (all standards shall be as per current versions) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications of MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of Translucent polypropylene (PP) or similar ready to use and sterile ( by ETO method) diameter : 10 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Material: medical grade , polypropylene plastic, latex free, translucent If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. autoclavable at 121°C Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: (all standards shall be as per current versions) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
424 | Medical & Surgical Instruments | MVACANNULAE12MM | Cannulae, Easygrip, 12mm, integrated base, SR12 | Cannulae, Easygrip, 12mm, integrated base, SR12 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or similar sterile ( by ETO method) diameter : 12 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. material: medical grade , polypropylene plastic, latex free, translucent If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. autoclavable at 121°C Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (all standards shall be as per current versions) Safety & product Standards: Must comply with the following standards: ISO 13485:2016 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or similar sterile ( by ETO method) diameter : 12 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. material: medical grade , polypropylene plastic, latex free, translucent If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. autoclavable at 121°C Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (all standards shall be as per current versions) Safety & product Standards: Must comply with the following standards: ISO 13485:2016 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1:2007 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2:2006 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: Not use PVC polymer
425 | Medical & Surgical Instruments | MVACANNULAE4MM | Cannulae, Easygrip, 4mm, integrated base, SR4 | Cannulae, Easygrip, 4mm, integrated base, SR4 | Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product Description: EasyGrip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent polypropylene or similar Supplied as initially sterile. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Sterilization method: ethylene oxide or gamma radiation. Diameter: 4 mm (indicated by a number in the cannulae) Length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Intended use: EasyGrip cannulae combined with Ipas MVA Plus Aspirator are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Select appropriate Easy Grip cannulae. Cannulae are supplied as sterile but can be reused after the first use. Reusable items must be mechanically cleaned, disinfected and sterilized between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Cannulae should be stored in dry, covered containers or packages, protected form contaminants. Supplied with: Manufacturer´s instruction for use, assembly/disassembly and maintenance in English/French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Ipas MVA Plus Aspirator with 2cc Silicone tube. Accesory kit for MVA incl. Plunger O-ring, Collar stop, Cap and 2cc Silicone tube. Packaging and labelling: Packaging: One (1) cannula in an individual sterilised peel-off pack made of paper and/or plastic. Label includes: Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word"sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word "don´t use if the packaging is damag" (or equivalent harmonized symbol). CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. ISO 17664: Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO14001 | Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product Description: EasyGrip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent polypropylene or similar Supplied as initially sterile. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Sterilization method: ethylene oxide or gamma radiation. Diameter: 4 mm (indicated by a number in the cannulae) Length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Intended use: EasyGrip cannulae combined with Ipas MVA Plus Aspirator are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Select appropriate Easy Grip cannulae. Cannulae are supplied as sterile but can be reused after the first use. Reusable items must be mechanically cleaned, disinfected and sterilized between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Cannulae should be stored in dry, covered containers or packages, protected form contaminants. Supplied with: Manufacturer´s instruction for use, assembly/disassembly and maintenance in English/French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Ipas MVA Plus Aspirator with 2cc Silicone tube. Accesory kit for MVA incl. Plunger O-ring, Collar stop, Cap and 2cc Silicone tube. Packaging and labelling: Packaging: One (1) cannula in an individual sterilised peel-off pack made of paper and/or plastic. Label includes: Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word"sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word "don´t use if the packaging is damag" (or equivalent harmonized symbol). CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. ISO 17664: Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO14001
426 | Medical & Surgical Instruments | MVACANNULAE5MM | Cannulae, Easygrip, 5mm, integrated base, SR5 | Cannulae, Easygrip, 5mm, integrated base, SR5 | Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product Description: EasyGrip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or similar Supplied as initially sterile. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Sterilization method: ethylene oxide or gamma radiation. Diameter: 5 mm (indicated by a number in the cannulae) Length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Intended use: EasyGrip cannulae combined with Ipas MVA Plus Aspirator are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Select appropriate Easy Grip cannulae. Cannulae are supplied as sterile but can be reused after the first use. Reusable items must be mechanically cleaned, disinfected and sterilized between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Cannulae should be stored in dry, covered containers or packages, protected form contaminants. Supplied with: Manufacturer´s instruction for use, assembly/disassembly and maintenance in English/French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Ipas MVA Plus Aspirator with 2cc Silicone tube. Accesory kit for MVA incl. Plunger O-ring, Collar stop, Cap and 2cc Silicone tube. Packaging and labelling: Packaging: One (1) cannula in an individual sterilised peel-off pack made of paper and/or plastic. Label includes: Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word"sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word "don´t use if the packaging is damag" (or equivalent harmonized symbol). CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. ISO 17664: Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO14001 | Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product Description: EasyGrip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or similar Supplied as initially sterile. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Sterilization method: ethylene oxide or gamma radiation. Diameter: 5 mm (indicated by a number in the cannulae) Length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Intended use: EasyGrip cannulae combined with Ipas MVA Plus Aspirator are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Select appropriate Easy Grip cannulae. Cannulae are supplied as sterile but can be reused after the first use. Reusable items must be mechanically cleaned, disinfected and sterilized between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Cannulae should be stored in dry, covered containers or packages, protected form contaminants. Supplied with: Manufacturer´s instruction for use, assembly/disassembly and maintenance in English/French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Ipas MVA Plus Aspirator with 2cc Silicone tube. Accesory kit for MVA incl. Plunger O-ring, Collar stop, Cap and 2cc Silicone tube. Packaging and labelling: Packaging: One (1) cannula in an individual sterilised peel-off pack made of paper and/or plastic. Label includes: Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word"sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word "don´t use if the packaging is damag" (or equivalent harmonized symbol). CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 11135: Sterilization of health-care products - Ethylene oxide, or equivalent corresponding to sterilization method. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993-1: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. ISO 17664: Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO14001
427 | Medical & Surgical Instruments | MVACANNULAE6MM | Cannulae, Easygrip, 6mm, integrated base, SR6 | Cannulae, Easygrip, 6mm, integrated base, SR6 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or equivalent sterile ( by ETO method) If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. diameter : 6 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Made of translucent PP or equivalent sterile ( by ETO method) If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. diameter : 6 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm Two opposing apertures at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001
428 | Medical & Surgical Instruments | MVACANNULAE7MM | Cannulae, Easygrip, 7mm, integrated base, SR7 | Cannulae, Easygrip, 7mm, integrated base, SR7 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Material: Medical grade, latex-free, polypropylene plastic. Sterile, single use. Diameter: 7 mm (indicated by a number in the cannulae). Length: 24 cm or 9.5 in (+/- 5%). Two opposing apertures at the end of the cannulae (4-8 mm). Printed insertion guides: six marking dots starting at 6 cm from the tip and spaced 1 cm apart. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Material: Medical grade, latex-free, polypropylene plastic. Sterile, single use. Diameter: 7 mm (indicated by a number in the cannulae). Length: 24 cm or 9.5 in (+/- 5%). Two opposing apertures at the end of the cannulae (4-8 mm). Printed insertion guides: six marking dots starting at 6 cm from the tip and spaced 1 cm apart. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001
429 | Medical & Surgical Instruments | MVACANNULAE8MM | Cannulae, Easygrip, 8mm, integrated base, SR8 | Cannulae, Easygrip, 8mm, integrated base, SR8 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Material: Medical grade, latex-free, polypropylene plastic. Sterile, single use. Diameter: 8 mm (indicated by a number in the cannulae). Length: 24 cm or 9.5 in (+/- 5%). Two opposing apertures at the end of the cannulae (4-8 mm). Printed insertion guides: six marking dots starting at 6 cm from the tip and spaced 1 cm apart. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. Material: Medical grade, latex-free, polypropylene plastic. Sterile, single use. Diameter: 8 mm (indicated by a number in the cannulae). Length: 24 cm or 9.5 in (+/- 5%). Two opposing apertures at the end of the cannulae (4-8 mm). Printed insertion guides: six marking dots starting at 6 cm from the tip and spaced 1 cm apart. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001
430 | Medical & Surgical Instruments | MVACANNULAE9MM | Cannulae, Easygrip, 9mm, integrated base, SR9 | Cannulae, Easygrip, 9mm, integrated base, SR9 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. sterile ( by ETO method) diameter : 9 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. material: medical grade , polypropylene plastic, latex free, translucent reusable autoclavable at 121°C If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Tubular device available in different lengths and diameters with integrated winged adapter, intended for insertion into the uterine cavity for evacuation of uterine content. Product description: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Tubular device with integrated winged adapted , colour code by size Semi-rigid. sterile ( by ETO method) diameter : 9 mm ( indicated by a number in the cannulae) length: 22.5 - 24 cm large, single scoop aperture at the end of the cannulae Printed insertion guides: six marking dots , starting 6 cm form the tip, spaced 1 cm. material: medical grade , polypropylene plastic, latex free, translucent reusable autoclavable at 121°C If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Instructions for use: Easy Grip cannulae combined with MVA are used for uterine aspiration/ uterine evacuation in obstetrics and gynaecology patients. Applications o MVA includes: - treatment of incomplete abortion for uterine sizes up to 12 weeks from the last menstrual period (LMP) - first-trimester abortion (menstrual regulation) - endometrial biopsy Select appropriate Easy Grip cannulae. Reusable items must be cleaned, disinfected and sterilized between patients and sterilized. Detailed procedure for the management of Reusable MVA Instruments shall be available Don´t use "flash" setting. Don´t use higher temperatures as damage may occur. Cannulae should be stored in dry, covered containers or packages, protected form contaminants Supplied with: - 1x manufacture´s instruction for use, assembly/disassembly and maintenance in English. Accessories/ spare parts/consumables: - MVA Aspirator Packaging and labelling: Packaging: 1 cannulae bi-packed in an individual sterilised peel-off pack made of paper and/or plastic. Labelling Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. The words "check the integrity of the individual sterilisation protection before use" (if space allows). The word " don´t use if the packaging is damage" (or equivalent harmonized symbol) CE marking and Notified Body number Symbols used according ISO 15223 and EN 980 Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process ISO 11607-1 Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2 Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO14001
431 | Medical & Surgical Instruments | MVADENISTONDILAT | Dilators, Denniston, polymer, P5 | Dilators, Denniston, polymer, P5 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Dilator is a plastic, lightweight instrument used in gynaecological procedures for dilating the cervix. Product description: Set of five double-ended, tapered dilators are accessories for Ipas MVA Plus aspirator. Set of five different size plastic dilators. Graduated, lightweight. Material: Polypropylene plastic (medical grade, latex-free). Length: 29 cm or 11.5 in (+/- 5%). Sizes: 5/6 mm (one end has 5 mm tapered tip and the other end has a 6 mm tapered tip), 7/8 mm, 9/10 mm, 11/12 mm, and 13/14 mm. Non-sterile, reusable. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Hygenically packed for medical use Instruction for use: Prior to use dilators must be sterilized or high-level disinfected before each use as per manufacturer's instructions (e.g. steam autoclave). Packaging & Labelling: Five (5) different size dilators per pack. Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. Symbols used according ISO 15223. Labelling shall be as per the regulations applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Dilator is a plastic, lightweight instrument used in gynaecological procedures for dilating the cervix. Product description: Set of five double-ended, tapered dilators are accessories for Ipas MVA Plus aspirator. Set of five different size plastic dilators. Graduated, lightweight. Material: Polypropylene plastic (medical grade, latex-free). Length: 29 cm or 11.5 in (+/- 5%). Sizes: 5/6 mm (one end has 5 mm tapered tip and the other end has a 6 mm tapered tip), 7/8 mm, 9/10 mm, 11/12 mm, and 13/14 mm. Non-sterile, reusable. If claimed reusable, supplier shall provide a detailed procedure for cleaning, disinfection and sterilization. Hygenically packed for medical use Instruction for use: Prior to use dilators must be sterilized or high-level disinfected before each use as per manufacturer's instructions (e.g. steam autoclave). Packaging & Labelling: Five (5) different size dilators per pack. Name and/ or trade name, and address of the manufacturer. Manufacturer's product reference. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. Symbols used according ISO 15223. Labelling shall be as per the regulations applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO14001
432 | Medical & Surgical Instruments | MVAKITTINGSERVICES | MVA Ipas kitting service | MVA Ipas kitting service | UNFPA suppliers shall provide kitting service to pack more than one item (1-4 specified below) into an individual bag as a kit. Requestors have the ability to choose which, and the amount of items to be included in one kit | UNFPA suppliers shall provide kitting service to pack more than one item (1-4 specified below) into an individual bag as a kit. Requestors have the ability to choose which, and the amount of items to be included in one kit
433 | Medical & Surgical Instruments | MVAPLUSASPIRATOR | MVA with 2cc silicone | MVA with 2cc silicone | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for uterine evacuation (manual vacuum aspiration) including treatment of incomplete abortion and induced abortion in early pregnancy. Product description: Ipas MVA Plus is a plastic, lightweight MVA (manual vacuum aspiration) system Ipas MVA Plus Aspirator components: Plunger, Cylinder, Valve Body, Collar Stop, O-ring, Buttons, Liner Clamp, Liner, Valve Cap. A valve with 2 buttons (double valve) that control the vacuum, a cap and a removable liner. Valve liner is removable by opening hinged valve body. A plunger with a plunger handle and O-ring. A 60cc cylinder for holding evacuated uterine contents, with a retaining clip for the collar stop. A collar stop. Material type: Polypropylene (latex free). Capacity: 60 ml (60 cc). Vacuum: 610-660 mm (24-26 in) Hg. Non-sterile, Autoclavable, reusable. Supplied with: 1 x tube of 2cc silicone. The cap, collar stop and O-ring as components are suitable to be reused a maximum of 25 times. Instruction for use: Ipas MVA Plus aspirator requires high-level disinfection or sterilization (steam) prior to initial use and between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Aspirator must be suitable to be reused a minimum of 25-50 times under normal use and maintenance. To be used with Ipas EasyGrip® cannulae, all sizes, no adapters needed (cannulae are sold separately). Cannulae must be high-level disinfected or sterile at the time of use. Maintenance instructions shall have information about the following: Read carefully the IFU from the manufacturer. Cleaning, disinfection and sterilization as per manufacturer's instruction. Before use, reassemble, lubricate and check vacuum capability. The aspirator should be discarded and replaced for any of the following reasons or any other as applicable: The cylinder is brittle or cracked or mineral deposits inhibit plunger movement. The valve parts are cracked, bent or broken. The buttons are broken. The plunger arms do not lock. The aspirator no longer holds a vacuum. Storage conditions to be specified as below or as per manufactuers instructions: Store between 15 and 30°C Protect from sunlight and humidity Packaging & Labelling: Unit presentation: 1 (one) MVA Ipas Aspirator with accessories. Symbols used according ISO 15223. Labelling shall be as per the regulations applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for uterine evacuation (manual vacuum aspiration) including treatment of incomplete abortion and induced abortion in early pregnancy. Product description: Ipas MVA Plus is a plastic, lightweight MVA (manual vacuum aspiration) system Ipas MVA Plus Aspirator components: Plunger, Cylinder, Valve Body, Collar Stop, O-ring, Buttons, Liner Clamp, Liner, Valve Cap. A valve with 2 buttons (double valve) that control the vacuum, a cap and a removable liner. Valve liner is removable by opening hinged valve body. A plunger with a plunger handle and O-ring. A 60cc cylinder for holding evacuated uterine contents, with a retaining clip for the collar stop. A collar stop. Material type: Polypropylene (latex free). Capacity: 60 ml (60 cc). Vacuum: 610-660 mm (24-26 in) Hg. Non-sterile, Autoclavable, reusable. Supplied with: 1 x tube of 2cc silicone. The cap, collar stop and O-ring as components are suitable to be reused a maximum of 25 times. Instruction for use: Ipas MVA Plus aspirator requires high-level disinfection or sterilization (steam) prior to initial use and between patients. Detailed procedure for the management of Reusable MVA Instruments shall be available Aspirator must be suitable to be reused a minimum of 25-50 times under normal use and maintenance. To be used with Ipas EasyGrip® cannulae, all sizes, no adapters needed (cannulae are sold separately). Cannulae must be high-level disinfected or sterile at the time of use. Maintenance instructions shall have information about the following: Read carefully the IFU from the manufacturer. Cleaning, disinfection and sterilization as per manufacturer's instruction. Before use, reassemble, lubricate and check vacuum capability. The aspirator should be discarded and replaced for any of the following reasons or any other as applicable: The cylinder is brittle or cracked or mineral deposits inhibit plunger movement. The valve parts are cracked, bent or broken. The buttons are broken. The plunger arms do not lock. The aspirator no longer holds a vacuum. Storage conditions to be specified as below or as per manufactuers instructions: Store between 15 and 30°C Protect from sunlight and humidity Packaging & Labelling: Unit presentation: 1 (one) MVA Ipas Aspirator with accessories. Symbols used according ISO 15223. Labelling shall be as per the regulations applied to the device. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: (All standards shall be current) Must comply with the following standards: ISO 13485 Medical devices -Quality management systems -- Requirements for regulatory purposes ISO 14971 Medical Devices- Application of risk management to medical devices ISO 10993-1 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: ISO14001
434 | Medical & Surgical Instruments | MVASILICONETUBE10 | Silicone, 2cc, P10 | Silicone, 2cc, P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for lubrication of the O-ring of Ipas MVA Plus aspirator. Product description: Medical grade silicone Components: Tube (volume 2ml or 2cc) containing lubricant. Lubricant: Silicone. Tube: The top of tube is removed so the silicone can be removed. Material polypropylene. Cap: Part which is pressed onto the top of the tube to prevent silicone from leaking out of the tube. Material polypropylene. Non-sterile. Instructions for use: Lubricant applied to the O-ring to ease Ipas MVA Plus aspirator plunger movement. Packaging & Labelling: Pack of ten (10) silicone 2cc tubes. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for lubrication of the O-ring of Ipas MVA Plus aspirator. Product description: Medical grade silicone Components: Tube (volume 2ml or 2cc) containing lubricant. Lubricant: Silicone. Tube: The top of tube is removed so the silicone can be removed. Material polypropylene. Cap: Part which is pressed onto the top of the tube to prevent silicone from leaking out of the tube. Material polypropylene. Non-sterile. Instructions for use: Lubricant applied to the O-ring to ease Ipas MVA Plus aspirator plunger movement. Packaging & Labelling: Pack of ten (10) silicone 2cc tubes. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485: 2016 Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical Devices- Application of risk management to medical devices Environmental requirements: ISO14001
435 | Medical & Surgical Instruments | NEEDLEDESCHAMPSL | Needle, aneurysm, Deschamps, 20cm, sharp, left hand | Needle, aneurysm, Deschamps, 20cm, sharp, left hand | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to place ligatures around vessels in poorly accessible sites. Use to place, mobilize and tie suture threads around blood vessels or aneurysms, in order to achieve adequate haemostasis. Product description: Long thin instrument with a palm-held handle and a thin needle-like extension which, at its tip, bends laterally at right angles and then describes a semicircle in the same plane. The tip has a needle eye and a pointed but not sharp end. Designed to place ligatures around vessels in poorly accessible sites. Made of stainless steel ( medical grade) Curved right, for left hand. Needle, pointed 20cm Dimensions: width – 1.5cm; height: 21cm; diameter: 1cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place ligatures around vessels in poorly accessible sites. retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark with Notified Body number Labelling shall be as per regulations applied to the device Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to place ligatures around vessels in poorly accessible sites. Use to place, mobilize and tie suture threads around blood vessels or aneurysms, in order to achieve adequate haemostasis. Product description: Long thin instrument with a palm-held handle and a thin needle-like extension which, at its tip, bends laterally at right angles and then describes a semicircle in the same plane. The tip has a needle eye and a pointed but not sharp end. Designed to place ligatures around vessels in poorly accessible sites. Made of stainless steel ( medical grade) Curved right, for left hand. Needle, pointed 20cm Dimensions: width – 1.5cm; height: 21cm; diameter: 1cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place ligatures around vessels in poorly accessible sites. retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark with Notified Body number Labelling shall be as per regulations applied to the device Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
436 | Medical & Surgical Instruments | NEEDLEDESCHAMPSR | Needle, aneurysm, Deschamps, 20cm, sharp, right hand | Needle, aneurysm, Deschamps, 20cm, sharp, right hand | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to place ligatures around vessels in poorly accessible sites. Use to place, mobilize and tie suture threads around blood vessels or aneurysms, in order to achieve adequate haemostasis. Product description: Designed to place ligatures around vessels in poorly accessible sites. Long thin instrument with a palm-held handle and a thin needle-like extension which, at its tip, bends laterally at right angles and then describes a semicircle in the same plane. The tip has a needle eye and a pointed but not sharp end. Designed to place ligatures around vessels in poorly accessible sites. Made of stainless steel ( medical grade) Curved left, for right hand. Needle, pointed 20cm Dimensions: width – 1.5cm; height: 21cm; diameter: 1cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place ligatures around vessels in poorly accessible sites. retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark with Notified Body number Labelling shall be as per regulations applied to the device Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to place ligatures around vessels in poorly accessible sites. Use to place, mobilize and tie suture threads around blood vessels or aneurysms, in order to achieve adequate haemostasis. Product description: Designed to place ligatures around vessels in poorly accessible sites. Long thin instrument with a palm-held handle and a thin needle-like extension which, at its tip, bends laterally at right angles and then describes a semicircle in the same plane. The tip has a needle eye and a pointed but not sharp end. Designed to place ligatures around vessels in poorly accessible sites. Made of stainless steel ( medical grade) Curved left, for right hand. Needle, pointed 20cm Dimensions: width – 1.5cm; height: 21cm; diameter: 1cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to place ligatures around vessels in poorly accessible sites. retract the lips of an incision or a deep vaginal wound This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 CE mark with Notified Body number Labelling shall be as per regulations applied to the device Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
437 | Medical & Surgical Instruments | NEEDLEHOLDER11.5 | Needle holder, Derf, 11.5 cm, straight | Needle holder, Derf, 11.5 cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to handle and insert suture needles during wound closures. Product description: Needle holder, Derf, for holding suture needles while stitching, straight. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. Springy Highly impact resistant Jaws with well-defined longitudinal groove. Tungsten carbide jaws and rachet that enables the needle to be gripped with varying tightness. Multiple ratchet Box lock Length: approx..115 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Use to hold suture needles while stitching This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to handle and insert suture needles during wound closures. Product description: Needle holder, Derf, for holding suture needles while stitching, straight. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. Springy Highly impact resistant Jaws with well-defined longitudinal groove. Tungsten carbide jaws and rachet that enables the needle to be gripped with varying tightness. Multiple ratchet Box lock Length: approx..115 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Use to hold suture needles while stitching This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
438 | Medical & Surgical Instruments | NEEDLEHOLDER15 | Needle holder, Baby-Crile-Wood, 15cm, delicate | Needle holder, Baby-Crile-Wood, 15cm, delicate | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Baby-Crile-Wood, used for holding tiny suture needles while stitching. For vascular and fine surgeries For fine needles and threads gauge 4/0, 5/0 and 6/0. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Serrated jaws, holds delicate needles, length 150mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for vascular and fine surgeries This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Baby-Crile-Wood, used for holding tiny suture needles while stitching. For vascular and fine surgeries For fine needles and threads gauge 4/0, 5/0 and 6/0. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. Serrated jaws, holds delicate needles, length 150mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for vascular and fine surgeries This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
439 | Medical & Surgical Instruments | NEEDLEHOLDER18C | Needle holder, Mayo-Hegar, 18cm, curved | Needle holder, Mayo-Hegar, 18cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, curved. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. curved A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 180mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, curved. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. curved A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 180mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
440 | Medical & Surgical Instruments | NEEDLEHOLDER18S | Needle holder, Mayo-Hegar, 18cm, straight | Needle holder, Mayo-Hegar, 18cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, straight. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. straight A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 180mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, straight. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. straight A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 180mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
441 | Medical & Surgical Instruments | NEEDLEHOLDER20C | Needle holder, Mayo-Hegar, 20cm, curved | Needle holder, Mayo-Hegar, 20cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, curved. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. curved A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, curved. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.20% carbon; 13% chromium. curved A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
442 | Medical & Surgical Instruments | NEEDLEHOLDER20S | Needle holder, Mayo-Hegar, 20cm, straight | Needle holder, Mayo-Hegar, 20cm, straight | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, used for holding tiny suture needles while stitching. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. Straight A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Designed to grasp a fine suture needle and thread during suturing. Product description: Needle holder, Mayo-Hegar, used for holding tiny suture needles while stitching. Used for general surgery. Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16 to 0.25% carbon; From 12 to 14% chromium. Straight A ratchet that enables the needle to be gripped with varying tightness. A well-defined longitudinal groove to prevent deterioration of the needle. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For holding suture needles while stitching. Used for general surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) needle holder in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151:1988 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
443 | Medical & Surgical Instruments | NEEDLELUER25G_100 | Needle, luer, 25G, sterile, single use, P100 | Needle, luer, 25G, sterile, single use, P100 | Intended purpose of use: intended to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient Product Description: Needle with base (Luer type fitting) and protective cap. Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 25 G (0.5 x 25mm). Hub color Orange (ISO 6009) Non-toxic. Pyrogen free. Cannula and tube: Made of stainless steel. Normal and thin walled. Long bevelled (angle inferior or equal to 13°) or short bevelled (angle between 13° and 19°). Silicone coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropene (PP). Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For subcutaneous or intramuscular (SC,IM) injection: adults For intravenous (IV) injection: children For intramuscular (IM) injection: children, infants For intra-arterial: adults, children Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) needle packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) Box of hundred (100) blister packed needles. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the dice should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: Gel requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices- Part : Evaluation and testing within a risk management process. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management s | Intended purpose of use: intended to be used in conjunction with syringes or giving sets to prepare and administer fluids/medications/drugs to a patient and/or to withdraw (aspirate) fluids from a patient Product Description: Needle with base (Luer type fitting) and protective cap. Material: Needle - stainless steel. Base and protective cap - Medical grade plastic. Sterile, single use (ethylene oxide sterilisation). External diameter expressed in gauge and millimeters. Length expressed preferably in millimetres. Size: 25 G (0.5 x 25mm). Hub color Orange (ISO 6009) Non-toxic. Pyrogen free. Cannula and tube: Made of stainless steel. Normal and thin walled. Long bevelled (angle inferior or equal to 13°) or short bevelled (angle between 13° and 19°). Silicone coated to facilitate insertion. Base: Allows the needle to be connected to a syringe or other injection devices. Formed by a conical female Luer connection. Designed in such a way that it prevents the needle from rolling on a flat surface. Made of polypropene (PP). Internationally recognized colour code system allows easy identification of the hypodermic needle's external diameter. Cap: The cap protects the cannula and its bevel, and preserves the sterility of the needle. Made polypropylene (PP). Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: For subcutaneous or intramuscular (SC,IM) injection: adults For intravenous (IV) injection: children For intramuscular (IM) injection: children, infants For intra-arterial: adults, children Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Packaging and Labelling: Primary packaging: Unit of use. One (1) needle packed in an individual sterilised peel pack made of paper and/or plastic. Secondary packaging: Protected unit. One (1) Box of hundred (100) blister packed needles. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the dice should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "dispose after use", if space allows. Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liquids and gases in healthcare applications — Part 1: Gel requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10993-1: Biological evaluation of medical devices- Part : Evaluation and testing within a risk management process. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management s
444 | Medical & Surgical Instruments | NEEDLESCLPV25G_100 | Needle, scalp vein, butterfly, 25G, sterile, single use,P100 | Needle, scalp vein, butterfly, 25G, sterile, single use,P100 | Intended purpose of use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use. Product Description: Components of one unit: needle with protective cap, butterfly wings, tube including a Luer connector with a cap (all parts fit together to form one unit). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Needle: Material: stainless steel. Sharp and silicone-coated for atraumatic insertion. Mounted on a polymer base with two flexible fins that are used for anchoring and attaching the needle. Thin needle walls. Size selected: 25G (0.5x19 mm). With a protective cap. Butterfly wings: For better fixations and handling. Color code as per ISO 6009: orange Material: medical grade plastic or PVC. The external diameter (mm and gauge): Visible on the wings. Tube: Material: DEHP free polyurethane (PUR) Soft, flexible. Length: approx. 150-300 mm ending in a female Luer connector fitted with a stopper allowing an infusion set to be connected Connector: Female Luer. Usually interlocking that permits the attachment of an infusion set. Intended use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 19 mm) green, for adults. - 25G (0.50 x 19 mm) orange, for children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Packaging and Labelling: Primary packaging: One (1) scalp vein needle unit packed in individual sterilised peel-pack. Secondary packaging: One (1) carton box of fifty (50) scalp vein needle units to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. The words “DEHP free” (or equivalent harmonised symbol). Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liqud gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10555-1: Sterile, single-use intravascular catheters - Part 1: General requirements. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management | Intended purpose of use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. For short duration use. Product Description: Components of one unit: needle with protective cap, butterfly wings, tube including a Luer connector with a cap (all parts fit together to form one unit). Sterile and disposable. Single use. Sterilization method: Ethylene oxide. Needle: Material: stainless steel. Sharp and silicone-coated for atraumatic insertion. Mounted on a polymer base with two flexible fins that are used for anchoring and attaching the needle. Thin needle walls. Size selected: 25G (0.5x19 mm). With a protective cap. Butterfly wings: For better fixations and handling. Color code as per ISO 6009: orange Material: medical grade plastic or PVC. The external diameter (mm and gauge): Visible on the wings. Tube: Material: DEHP free polyurethane (PUR) Soft, flexible. Length: approx. 150-300 mm ending in a female Luer connector fitted with a stopper allowing an infusion set to be connected Connector: Female Luer. Usually interlocking that permits the attachment of an infusion set. Intended use: Sterile scalp vein needle to be introduced into a peripheral vein for infusing fluids and administering IV drugs. Scalp vein needle are for short duration use, for longer duration use, it is preferable to use IV Cannula, which causes less irritation to the veins. The needle should be held in the correct position by suitable dressing, allowing the observation of the site of insertion, and providing efficient barrier against infection. The selected sizes are the most commonly used: - 21G (0.80 x 19 mm) green, for adults. - 25G (0.50 x 19 mm) orange, for children. Safety process: The needle is for single use only. When inserting a needle: use aseptic techniques. Supplied with: Manufacturer’s instructions for use may be printed onto the primary packaging or provided on a separate insert. Languages: English, Spanish and French. Accessories / Spare Parts / Consumables (available but not supplied): N/A Packaging and Labelling: Primary packaging: One (1) scalp vein needle unit packed in individual sterilised peel-pack. Secondary packaging: One (1) carton box of one hundred (100) scalp vein needle units to blister packs. Primary package label includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol). The words "check the package integrity before use", or similar warning. The words "destroy after use", if space allows. The words “DEHP free” (or equivalent harmonised symbol). Secondary package label includes the number of units. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety and Product Standards (current versions of the following standards): ISO 11135: Sterilization of health-care products - Ethylene oxide. ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices — Part 2: Validation requirements for forming, sealing and assembly processes. ISO 6009: Hypodermic needles for single use - Colour coding for identification. ISO 7864: Sterile hypodermic needles for single use. ISO 80369-1: Small-bore connectors for liqud gases in healthcare applications — Part 1: General requirements. ISO 80369-7: Small-bore connectors for liquids and gases in healthcare applications — Part 7: Connectors for intravascular or hypodermic applications. ISO 23908: Sharps injury protection - Requirements and test methods — Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling. ISO 9626-1: Stainless steel needle tubing for the manufacture of medical devices — Requirements and test methods. ISO 10555-1: Sterile, single-use intravascular catheters - Part 1: General requirements. ISO 10993: Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process. Environmental Requirements: Use of other polymer instead of PVC. ISO 14001: Environmental management
445 | Medical & Surgical Instruments | PERFORATOROBSTET | Perforator, Smellie, 25cm | Perforator, Smellie, 25cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for piercing foetal skull and reduce its diameter in case of death in uterus, in order to permit vaginal extraction. Product description: Perforator, Smellie is a destructive tool used to pierce the skull of a dead foetus Material: Martensitic steel (quenched, magnetic steel). Cuts by perforating shears. composition: 0.40 % carbon; 14% chromium. Length: approx. 250mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To pierce foetal skull and reduce its diameter in case of death in uterus, in order to permit vaginal extraction. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) perforwtor in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( as per the latest revision of standards): ISO 7153-1:1991 ISO 17664:2004 ISO 13402:1995 ISO 7741:1986 BS 5194-4:1985 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for piercing foetal skull and reduce its diameter in case of death in uterus, in order to permit vaginal extraction. Product description: Perforator, Smellie is a destructive tool used to pierce the skull of a dead foetus Material: Martensitic steel (quenched, magnetic steel). Cuts by perforating shears. composition: 0.40 % carbon; 14% chromium. Length: approx. 250mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To pierce foetal skull and reduce its diameter in case of death in uterus, in order to permit vaginal extraction. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) perforwtor in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards ( as per the latest revision of standards): ISO 7153-1:1991 ISO 17664:2004 ISO 13402:1995 ISO 7741:1986 BS 5194-4:1985
446 | Medical & Surgical Instruments | PROBE14.5CM | Probe, surgical, double-ended, 14.5 cm | Probe, surgical, double-ended, 14.5 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:for exploring a surgical wound or to follow the path of a fistula. Product description: Probe , a slender, rod-like, hand-held manual surgical instrument Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 16% to 18% chromium; 10% to 14% nickel; 2 % to 3% molybdenum. A bulbous shape at both ends. Diameter: approx. 2mm. Length: approx. 145mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: to explore a surgical wound or to follow the path of a fistula. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) probe in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: (as per the latest revisions of the standards) ISO 7153-1 ISO 17664 ISO 13402 ISO 7741 BS 5194-4 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:for exploring a surgical wound or to follow the path of a fistula. Product description: Probe , a slender, rod-like, hand-held manual surgical instrument Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 16% to 18% chromium; 10% to 14% nickel; 2 % to 3% molybdenum. A bulbous shape at both ends. Diameter: approx. 2mm. Length: approx. 145mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: to explore a surgical wound or to follow the path of a fistula. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) probe in a plastic bag. Symbols used according ISO 15223 CE mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: (as per the latest revisions of the standards) ISO 7153-1 ISO 17664 ISO 13402 ISO 7741 BS 5194-4
447 | Medical & Surgical Instruments | PROBEEYE20CM | Probe, surgical, with eye, malleable, 20 cm | Probe, surgical, with eye, malleable, 20 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:for exploring a surgical wound or to follow the path of a fistula. Product description: Probe with eye, malleable, 20 cm Probe and dilator. Bulbous shape, at both ends. Malleable: length 20cm, thickness 2mm Length: approximately 20cm. Materials: Austenitic steel (non-quenched, non-magnetic steel) Material composition: 16 -18% chromium; 10-12% nickel: 1% silicon; 2% manganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to explore a surgical wound or to follow the path of a fistula.This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) probe in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:for exploring a surgical wound or to follow the path of a fistula. Product description: Probe with eye, malleable, 20 cm Probe and dilator. Bulbous shape, at both ends. Malleable: length 20cm, thickness 2mm Length: approximately 20cm. Materials: Austenitic steel (non-quenched, non-magnetic steel) Material composition: 16 -18% chromium; 10-12% nickel: 1% silicon; 2% manganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used to explore a surgical wound or to follow the path of a fistula.This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) probe in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
448 | Medical & Surgical Instruments | RETRACTFARABEUF15 | Retractor, double-ended, Farabeuf, 15 cm, pair | Retractor, double-ended, Farabeuf, 15 cm, pair | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used in superficial surgery to retract skin, fatty tissues and muscles after the incision Product description: A handheld retractor that is a solid piece of stainless steel with one angled blade at each end. Each blade is smooth, blunt, and also has a slight upward lip at the end Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Retractor with handle, double-ended with angled blade. One end 24 x 16mm, the other end 28 x 16mm. Supplied in pairs. Length: approx. 150mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used in superficial surgery to retract skin, fatty tissues and muscles after the incision Product description: A handheld retractor that is a solid piece of stainless steel with one angled blade at each end. Each blade is smooth, blunt, and also has a slight upward lip at the end Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Retractor with handle, double-ended with angled blade. One end 24 x 16mm, the other end 28 x 16mm. Supplied in pairs. Length: approx. 150mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
449 | Medical & Surgical Instruments | RETRACTLANGENBECK | Retractor, Langenbeck, single-ended, medium, 21cm | Retractor, Langenbeck, single-ended, medium, 21cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to retract the margins of a full-thickness or near-full-thickness incision or injury (i.e., the skin and underlying connective tissue, skeletal muscle) during a surgical intervention. Product description: also called Kocher - Langenbeck retractor. Handheld retractor with fenestrated handle. Single angled (90°) spatula with curved end Handeld retractors used to retract muscle and soft tissues. Spatula deep: 44mm. Spatula wide: 13mm. Handle length: 21cm Materials: Austenitic steel. Material compositiion: 16 -18% chromium; 2-3% molybdenum; 8 -10% nickel ; 1% silicon ; 2% maganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For retracting muscle and soft tissue. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to retract the margins of a full-thickness or near-full-thickness incision or injury (i.e., the skin and underlying connective tissue, skeletal muscle) during a surgical intervention. Product description: also called Kocher - Langenbeck retractor. Handheld retractor with fenestrated handle. Single angled (90°) spatula with curved end Handeld retractors used to retract muscle and soft tissues. Spatula deep: 44mm. Spatula wide: 13mm. Handle length: 21cm Materials: Austenitic steel. Material compositiion: 16 -18% chromium; 2-3% molybdenum; 8 -10% nickel ; 1% silicon ; 2% maganese. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For retracting muscle and soft tissue. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
450 | Medical & Surgical Instruments | RETRACTORBALFOUR | Retractor, abdominal, Balfour, 3 blades | Retractor, abdominal, Balfour, 3 blades | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used during urological, gynecological and general surgeries to retract internal organs and tissues to enhance the view of the surgical field Product description: Self retaining retractor, three Blade Retractor For Increased Surgical View. Used to retract the abdominal wall and viscera to provide good exposure during a laparotomy Suprapubic abdominal retractor with locking mechanism and compass (self-holding). Consists of three parts: 1 compass (frame) + 2 lateral blades + 1 central blade The compass is a support with hinged arms. Fenestrated lateral blades: approx. 35 x 70 mm. Solid center blade: approx. 45 x 80 mm. Maximum wound opening: approx. 180 mm. Material: Martensitic steel (quenched, magnetic steel). composition: From 0.42% to 0.55% carbon; From 12% to 15% chromium; From 0.45% to 0.9% molybdenum. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used during urological, gynecological and general surgeries to retract internal organs and tissues to enhance the view of the surgical field Product description: Self retaining retractor, three Blade Retractor For Increased Surgical View. Used to retract the abdominal wall and viscera to provide good exposure during a laparotomy Suprapubic abdominal retractor with locking mechanism and compass (self-holding). Consists of three parts: 1 compass (frame) + 2 lateral blades + 1 central blade The compass is a support with hinged arms. Fenestrated lateral blades: approx. 35 x 70 mm. Solid center blade: approx. 45 x 80 mm. Maximum wound opening: approx. 180 mm. Material: Martensitic steel (quenched, magnetic steel). composition: From 0.42% to 0.55% carbon; From 12% to 15% chromium; From 0.45% to 0.9% molybdenum. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
451 | Medical & Surgical Instruments | RETRACTORCOLLIN | Retractor, abdominal, Collin, 3 blades | Retractor, abdominal, Collin, 3 blades | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used to retract the abdominal wall and viscera to provide good exposure of the operative field. Product description: Abdominal retractor with locking mechanism (self-retaining). 3 parts: 1 frame; 2 lateral blades and 1 central blade The two side(lateral) blades being mobile and the middle blade removable. Broad Lateral Blades For Clean Retraction Central Blade To Ensure Clean Surgical Plane Progressive setting and locking of the ratchet Material: Martensitic steel (quenched, magnetic steel). Composition: From 0.42% to 0.55% carbon; From 12% to 15% chromium; From 0.45% to 0.9% molybdenum. Solid lateral blades: approx. 38 x 60 mm. Solid center blade: approx. 38 x 55 mm. Maximum wound opening: approx. 100 mm. Solid center blade: approx. 45 x 80 mm. Maximum wound opening: approx. 180 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. CE mark Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Used to retract the abdominal wall and viscera to provide good exposure of the operative field. Product description: Abdominal retractor with locking mechanism (self-retaining). 3 parts: 1 frame; 2 lateral blades and 1 central blade The two side(lateral) blades being mobile and the middle blade removable. Broad Lateral Blades For Clean Retraction Central Blade To Ensure Clean Surgical Plane Progressive setting and locking of the ratchet Material: Martensitic steel (quenched, magnetic steel). Composition: From 0.42% to 0.55% carbon; From 12% to 15% chromium; From 0.45% to 0.9% molybdenum. Solid lateral blades: approx. 38 x 60 mm. Solid center blade: approx. 38 x 55 mm. Maximum wound opening: approx. 100 mm. Solid center blade: approx. 45 x 80 mm. Maximum wound opening: approx. 180 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. CE mark Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
452 | Medical & Surgical Instruments | RETRACTORDOYEN | Retractor, vaginal, Doyen, 8.5 x 4.5cm | Retractor, vaginal, Doyen, 8.5 x 4.5cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to expose the vaginal cavity and to retract viscera in abdominal surgery. Product description: Vaginal retractor (also called vaginal speculum), Doyen, for spreading and opening the posterior wall of the vagina to visualize the cervix of the uterus; used to display the cervix in operations of the uterine cavity Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Weight which fits on the shank of the blade. Weight should be soldered/ welded rather than screwed. Lateral edges must be blunt. Blade: medium, hollow shaped. Blade length: approx. 85mm. Blade width: approx. 45mm. Length: approx. 240mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To expose the vaginal cavity. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to expose the vaginal cavity and to retract viscera in abdominal surgery. Product description: Vaginal retractor (also called vaginal speculum), Doyen, for spreading and opening the posterior wall of the vagina to visualize the cervix of the uterus; used to display the cervix in operations of the uterine cavity Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Weight which fits on the shank of the blade. Weight should be soldered/ welded rather than screwed. Lateral edges must be blunt. Blade: medium, hollow shaped. Blade length: approx. 85mm. Blade width: approx. 45mm. Length: approx. 240mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To expose the vaginal cavity. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
453 | Medical & Surgical Instruments | RETRACTORGOSSET | Retractor, abdominal, Gosset, 3 blades | Retractor, abdominal, Gosset, 3 blades | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to separate the margins of an abdominal incision typically during deep or extensive surgery, to expose or access organs for examination or intervention. Product description: Abdominal self-retaining Gosset retractor with 1 fixed arm and 1 mobile arm. The instrument has wire-like blades that hold the abdomen open and an additional third blade that is tensioned using a series of eyeholes or notches 3 blades including two sliding blades and a central blade and central blade used to retract skin, fatty tissue, muscles, peritoneum and bladder area after the abdominal incision to expose the operative field. Material: Martensitic steel (quenched, magnetic steel) composition: 18 to 20% chromium; 8 to 10% nickel. Self retaining. Sliding system enabling a variable opening (maximum 110mm). Mobile, adjustable Central blade: 57 x 60mm Fenestrated lateral blades: approx. depth 58mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to separate the margins of an abdominal incision typically during deep or extensive surgery, to expose or access organs for examination or intervention. Product description: Abdominal self-retaining Gosset retractor with 1 fixed arm and 1 mobile arm. The instrument has wire-like blades that hold the abdomen open and an additional third blade that is tensioned using a series of eyeholes or notches 3 blades including two sliding blades and a central blade and central blade used to retract skin, fatty tissue, muscles, peritoneum and bladder area after the abdominal incision to expose the operative field. Material: Martensitic steel (quenched, magnetic steel) composition: 18 to 20% chromium; 8 to 10% nickel. Self retaining. Sliding system enabling a variable opening (maximum 110mm). Mobile, adjustable Central blade: 57 x 60mm Fenestrated lateral blades: approx. depth 58mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: To retract skin, fatty tissue, muscles or viscera after the incision to expose the operative field. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol), if applicable. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate or equivalent harmonised symbol. Information for handling, if applicable (or equivalent harmonised symbol). Symbols used according ISO 15223 Over packaging: Packaging unit. CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
454 | Medical & Surgical Instruments | SAFETYPINMAYO | Safety Pin, forceps holder, Mayo, 11.5cm | Safety Pin, forceps holder, Mayo, 11.5cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: s a specialized sterilization instrument that the operator uses for hanging up ring handle devices with comfort, in order to sterilize them in an autoclave or other apparatus Product description: Mayo Safety Pin for holding ring handle forceps. Measurements: approx. 114 mm long. Material: Stainless Steel. Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for holding ring handel forceps. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) safety pin in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: s a specialized sterilization instrument that the operator uses for hanging up ring handle devices with comfort, in order to sterilize them in an autoclave or other apparatus Product description: Mayo Safety Pin for holding ring handle forceps. Measurements: approx. 114 mm long. Material: Stainless Steel. Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used for holding ring handel forceps. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) safety pin in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
455 | Medical & Surgical Instruments | SAFETYPINSCHAEDEL | Safety Pin, forceps holder, Schaedel, 9cm | Safety Pin, forceps holder, Schaedel, 9cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is a commonly used surgical instrument that is used to grasp tissues, drapes, gauze, and towels during surgeries Product description: Schaedel Towel Clamp Shaedel Safety Pin forceps folder. Measurements: 9cm long. Material: Stainless Steel. Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) safety pin in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is a commonly used surgical instrument that is used to grasp tissues, drapes, gauze, and towels during surgeries Product description: Schaedel Towel Clamp Shaedel Safety Pin forceps folder. Measurements: 9cm long. Material: Stainless Steel. Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) safety pin in a plastic bag. Symbols used according ISO 15223 CE mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 7151 Surgical instruments -- Non-cutting, articulated instruments -- General requirements and test methods ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
456 | Medical & Surgical Instruments | SC_UMBIL_USA_STR | Scissors (Mod. USA) for umbilical cord cutting, sterile | Scissors (Mod. USA) for umbilical cord cutting, sterile | Intended purpose of use: is used for the precise cutting of the umbilical cord. Product Description: Sterile scissors (Mod. USA) Material: Martensitic stainless steel (AISI 440 or equivalent) or strong medical grade polymer. Length: 10 to 14 cm. Blades: Cutting, sharp. Single-edged preferred, double-edged acceptable (inner blade edges can be sharp, outer edges and blade tips must be blunt). Blade tips: Blunt. Scissors may come with a blunt protective sheath. Sterile. Supplied with: Manufacturer's instructions for use (written and/or pictorial) are printed onto the primary package. If scissors are supplied as 'initially sterile' but are inteded for reuse: Instructions for cleaning, disinfection and sterilization are required. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Scissors are designed for cutting the umbilical cord after child birth. Scissors are supplied as initially sterile. IMPORTANT: The scissors are part of IARH kit 2A - Individual Delivery. Prevention of the kit's individual plastic bag puncturing is to be taken into consideration in the product design or packaging, e.g. by introducing sterile steath for the blade. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol), if applicable. The words "check the package integrity before use", or similar warning. The words "dispose after use", if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety and Product Standards (current versions of the following standards): ISO 11137 or equivalent for the applied sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993 -1: Biological evaluation of medical devices - Part 1 Evaluation and testing within a risk management process ISO 7153-1: Surgical instruments - Metallic materials - Part 1 Stainless steel ISO 13402: Surgical and dental hand instruments - Determination of resistance against autoclaving corrosion and thermal exposure ISO 7741: Instruments for surgery - Scissors and shears - General requirements and test methods ISO 17664: Sterilization of Medical Devices - Information to be provided by the manufacturer for the processing of resterilizable medical devices Environmental Requirements: ISO 14001: Environmental management systems | Intended purpose of use: is used for the precise cutting of the umbilical cord. Product Description: Sterile scissors (Mod. USA) Material: Martensitic stainless steel (AISI 440 or equivalent) or strong medical grade polymer. Length: 10 to 14 cm. Blades: Cutting, sharp. Single-edged preferred, double-edged acceptable (inner blade edges can be sharp, outer edges and blade tips must be blunt). Blade tips: Blunt. Scissors may come with a blunt protective sheath. Sterile. Supplied with: Manufacturer's instructions for use (written and/or pictorial) are printed onto the primary package. If scissors are supplied as 'initially sterile' but are inteded for reuse: Instructions for cleaning, disinfection and sterilization are required. Accessories / Spare Parts / Consumables (available but not supplied): N/A Intended use: Scissors are designed for cutting the umbilical cord after child birth. Scissors are supplied as initially sterile. IMPORTANT: The scissors are part of IARH kit 2A - Individual Delivery. Prevention of the kit's individual plastic bag puncturing is to be taken into consideration in the product design or packaging, e.g. by introducing sterile steath for the blade. Packaging and Labelling: One (1) unit in a protective packaging with a label that includes: Device identity and intended purpose. Manufacturer's product code or reference number. Manufacturer identification. Address of the manufacturing site. EC Rep identification. How the device should be used, maintained and stored. Lot/Batch and MFD and EXP. Any residual device risks, warnings, limitations or contraindications. Symbols used according ISO 15223. CE mark. The word "sterile" (or equivalent harmonised symbol). Sterilisation method (or equivalent harmonised symbol). The words "for single use" (or equivalent harmonised symbol), if applicable. The words "check the package integrity before use", or similar warning. The words "dispose after use", if applicable. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety and Product Standards (current versions of the following standards): ISO 11137 or equivalent for the applied sterilization method. ISO 11607-1: Packaging for Terminally sterilized medical devices - Part 1 : Requirements for materials, sterile barrier systems and packaging systems. ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes. ISO 10993 -1: Biological evaluation of medical devices - Part 1 Evaluation and testing within a risk management process ISO 7153-1: Surgical instruments - Metallic materials - Part 1 Stainless steel ISO 13402: Surgical and dental hand instruments - Determination of resistance against autoclaving corrosion and thermal exposure ISO 7741: Instruments for surgery - Scissors and shears - General requirements and test methods ISO 17664: Sterilization of Medical Devices - Information to be provided by the manufacturer for the processing of resterilizable medical devices Environmental Requirements: ISO 14001: Environmental management systems
457 | Medical & Surgical Instruments | SCALPELHANDL7_15CM | Scalpel handle, no. 7, blade holder, 15 cm | Scalpel handle, no. 7, blade holder, 15 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 15cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 15cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
458 | Medical & Surgical Instruments | SCALPELHANDL7_16CM | Scalpel handle, no. 7, blade holder, 16 cm | Scalpel handle, no. 7, blade holder, 16 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 16cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 16cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
459 | Medical & Surgical Instruments | SCALPELHANDL7_18CM | Scalpel handle, no. 7, blade holder, 18 cm | Scalpel handle, no. 7, blade holder, 18 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 18cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 7, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blade no 10-15. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 18cm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
460 | Medical & Surgical Instruments | SCALPELHANDLE3 | Scalpel handle no. 3, blade holder, 13 cm | Scalpel handle no. 3, blade holder, 13 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 3, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: 10, 11, 12 and 15 Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 125mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 3, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: 10, 11, 12 and 15 Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 125mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.10. Scalpel blade, ster, disp, no.11. Scalpel blade, ster, disp, no.12. Scalpel blade, ster, disp, no.15 Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
461 | Medical & Surgical Instruments | SCALPELHANDLE4 | Scalpel handle no. 4, blade holder, 13 cm | Scalpel handle no. 4, blade holder, 13 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 4, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: . 20, 21, 22, 23. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 135mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.20. Scalpel blade, ster, disp, no.21. Scalpel blade, ster, disp, no.22. Scalpel blade, ster, disp, no.23. Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 4, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: . 20, 21, 22, 23. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 135mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.20. Scalpel blade, ster, disp, no.21. Scalpel blade, ster, disp, no.22. Scalpel blade, ster, disp, no.23. Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
462 | Medical & Surgical Instruments | SCALPELHANDLE4L | Scalpel handle no. 4, blade holder, long 21cm | Scalpel handle no. 4, blade holder, long 21cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 4, long, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: . 20, 21, 22, 23. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 210mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.20. Scalpel blade, ster, disp, no.21. Scalpel blade, ster, disp, no.22. Scalpel blade, ster, disp, no.23. Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: functions as a handle designed to mount a compatible surgical blade. Product description: Scalpel Handle no 4, long, a device that is an interchangeable component of a scalpel. This device when fitted with the blade become a single-bladed surgical instrument intended to be used by a physician for cutting or dissecting tissue. Suitable for blades: . 20, 21, 22, 23. Material: Austenitic steel (non-quenched, non-magnetic steel). composition: 17% to 19% chromium; 8 to 10% nickel The handle number indicates the characteristic of the distal end and therefore the choice of the blades The number is printed on the handle Length: 210mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Accessories/Spare parts/Consumables: Handle is compatible with following sized surgical blades (Note! Blades are sold separately): Scalpel blade, ster, disp, no.20. Scalpel blade, ster, disp, no.21. Scalpel blade, ster, disp, no.22. Scalpel blade, ster, disp, no.23. Instructions for use: To hold blade for surgical incisions. To be used with disposable blades. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used by surgeons, only. Packaging & Labelling: Primary packaging: Unit of use One (1) scalpel in a plastic bag. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, manufacturer's name and address, and CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Protective packaging: Box of 10 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. ISO 13402 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices Environmental requirements: ISO 14001
463 | Medical & Surgical Instruments | SCALPELKNIFE22_1 | Scalpel, knife, blade no. 22, with plastic handle, sterile | Scalpel, knife, blade no. 22, with plastic handle, sterile | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for use in minor surgical incisions, e.g. in cutting umbilical cord. Product description: Basic, sterile cutting instrument Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation Instructions for use: The surgical bladed scalpel is for single-use only. Supplied in sterile packaging, may only be used if packaging is undamaged. Use aseptic techniques. Wear gloves. Perform with extra caution: Before use, remove the protective plastic guard covering the sharp metal blade. Never remove the metal blade from the scalpel handle. Check visually the integrity of the device before use. Dispose used scalpels into a sealed container to prevent injuries. Avoid recapping of the blade with the guard before disposal. After use, collect and destroy the sealed containers by incineration in a controlled environment Packaging and labelling: Primary packaging: One (1) sterile, single-use unit individually wrapped in a reinforced laminated, sterile foil peel pack. (Unit description: whole scalpel including blade with integrated handle and protective hood for the blade). Symbols used according ISO 15223. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Storage Conditions: Avoid direct sunlight and humidity. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 1622-1:2012/ISO 19062-1:2015, Plastics – Polystyrene (PS) moulding and extrusion materials – Part 1: Designation system and basis for specification ISO 1622-2:1995/ISO 2580-2:2003, Plastics – moulding and extrusion materials – Part 2: Preparation of test specimens and determination of properties. ISO 11135:2014/ISO11137 - Sterilization of health-care products ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 7153-1: Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. Environmentally requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for use in minor surgical incisions, e.g. in cutting umbilical cord. Product description: Basic, sterile cutting instrument Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation Instructions for use: The surgical bladed scalpel is for single-use only. Supplied in sterile packaging, may only be used if packaging is undamaged. Use aseptic techniques. Wear gloves. Perform with extra caution: Before use, remove the protective plastic guard covering the sharp metal blade. Never remove the metal blade from the scalpel handle. Check visually the integrity of the device before use. Dispose used scalpels into a sealed container to prevent injuries. Avoid recapping of the blade with the guard before disposal. After use, collect and destroy the sealed containers by incineration in a controlled environment Packaging and labelling: Primary packaging: One (1) sterile, single-use unit individually wrapped in a reinforced laminated, sterile foil peel pack. (Unit description: whole scalpel including blade with integrated handle and protective hood for the blade). Symbols used according ISO 15223. Protective packaging: Symbols used according ISO 15223. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Storage Conditions: Avoid direct sunlight and humidity. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 1622-1:2012/ISO 19062-1:2015, Plastics – Polystyrene (PS) moulding and extrusion materials – Part 1: Designation system and basis for specification ISO 1622-2:1995/ISO 2580-2:2003, Plastics – moulding and extrusion materials – Part 2: Preparation of test specimens and determination of properties. ISO 11135:2014/ISO11137 - Sterilization of health-care products ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 7153-1: Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. Environmentally requirements: ISO14001
464 | Medical & Surgical Instruments | SCALPELKNIFE22_P10 | Scalpel, knife, blade no. 22, w/ plastic handle, sterile,P10 | Scalpel, knife, blade no. 22, w/ plastic handle, sterile,P10 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for use in minor surgical incisions, e.g. in cutting umbilical cord. Product description: Basic, sterile cutting instrument Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation Instructions for use: The surgical bladed scalpel is for single-use only. Supplied in sterile packaging, may only be used if packaging is undamaged. Use aseptic techniques. Wear gloves. Perform with extra caution: Before use, remove the protective plastic guard covering the sharp metal blade. Never remove the metal blade from the scalpel handle. Check visually the integrity of the device before use. Dispose used scalpels into a sealed container to prevent injuries. Avoid recapping of the blade with the guard before disposal. After use, collect and destroy the sealed containers by incineration in a controlled environment Packaging and labelling: Primary packaging: One (1) sterile, single-use unit individually wrapped in a reinforced laminated, sterile foil peel pack. (Unit description: whole scalpel including blade with integrated handle and protective hood for the blade). Symbols used according ISO 15223. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Storage Conditions: Avoid direct sunlight and humidity. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 1622-1:2012/ISO 19062-1:2015, Plastics – Polystyrene (PS) moulding and extrusion materials – Part 1: Designation system and basis for specification ISO 1622-2:1995/ISO 2580-2:2003, Plastics – moulding and extrusion materials – Part 2: Preparation of test specimens and determination of properties. ISO 11135:2014/ISO11137 - Sterilization of health-care products ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 7153-1: Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. Environmentally requirements: ISO14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for use in minor surgical incisions, e.g. in cutting umbilical cord. Product description: Basic, sterile cutting instrument Surgical, single-edged blade (no. 22) integrated with an easy-grip plastic handle (non-retractable). Blade covered with sterile, removable, protective plastic guard (hood or sheath). Non-glare, straight, robust, flat plastic handle for easy handling, molded laterally for safe non-splip grip and precise control (e.g. ridges in handle). Material, blade: Martensitic stainless steel (quenched, magnetic steel) Material, handle: Robust DEHP free plastic (ABS or high-impact-polystyrene). Latex-free. Material, guard/hood: DEHP free plastic (e.g. LDPE or similar). Latex-free. Hardness, blade: 50 HRC to 58 HRC. Sterile and single-use. Disposable. Length: approx. between 13-18 cm (blade+handle, excluding the blade guard) Sterilization method: Ethylene oxide gas/Gamma radiation Instructions for use: The surgical bladed scalpel is for single-use only. Supplied in sterile packaging, may only be used if packaging is undamaged. Use aseptic techniques. Wear gloves. Perform with extra caution: Before use, remove the protective plastic guard covering the sharp metal blade. Never remove the metal blade from the scalpel handle. Check visually the integrity of the device before use. Dispose used scalpels into a sealed container to prevent injuries. Avoid recapping of the blade with the guard before disposal. After use, collect and destroy the sealed containers by incineration in a controlled environment Packaging and labelling: Primary packaging: One (1) sterile, single-use unit individually wrapped in a reinforced laminated, sterile foil peel pack. (Unit description: whole scalpel including blade with integrated handle and protective hood for the blade). Symbols used according ISO 15223. Protective packaging: Box of 100 units, cardboard, labelling is the same as unit presentation with total quantity. Symbols used according ISO 15223. Labelling: Product name, size, reference number, expiry date, lot number, sterilization method, single-use, manufacturer's name and address, CE mark and reference number of notifying body. Must be multilingual: English, French and Spanish, others when available. Storage Conditions: Avoid direct sunlight and humidity. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with following standards (as per latest revisions of the standards) ISO 13485 Medical devices -- Quality management systems - Requirements for regulatory purposes ISO 1622-1:2012/ISO 19062-1:2015, Plastics – Polystyrene (PS) moulding and extrusion materials – Part 1: Designation system and basis for specification ISO 1622-2:1995/ISO 2580-2:2003, Plastics – moulding and extrusion materials – Part 2: Preparation of test specimens and determination of properties. ISO 11135:2014/ISO11137 - Sterilization of health-care products ISO 11607-1: Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems ISO 11607-2: Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes ISO 7153-1: Surgical instruments -- Metallic materials -- Part 1: Stainless steel ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. EN 27740/ ISO 7740: Scalpels with detachable blades Fitting dimensions. Environmentally requirements: ISO14001
465 | Medical & Surgical Instruments | SCISSORDEAVER14C | Scissors, Deaver, 14cm, curved, sharp/blunt | Scissors, Deaver, 14cm, curved, sharp/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: The tip configuration sharp/blunt allow these scissors to be used for multipurpose cutting and dissecting. Used for basic surgical operations (sutures or for treating abscesses). Product description: Surgical dissecting scissors, Deaver, have all the functionality of standard operating scissors, but also feature a delicate pattern. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Curved, with one pointed and one blunt end blade. Length: approx. 140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 Over packaging: Packaging unit. CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: The tip configuration sharp/blunt allow these scissors to be used for multipurpose cutting and dissecting. Used for basic surgical operations (sutures or for treating abscesses). Product description: Surgical dissecting scissors, Deaver, have all the functionality of standard operating scissors, but also feature a delicate pattern. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Curved, with one pointed and one blunt end blade. Length: approx. 140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 Over packaging: Packaging unit. CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
466 | Medical & Surgical Instruments | SCISSOREPISIOTOM20 | Scissors, fistula, 20cm curved | Scissors, fistula, 20cm curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: part of fistula kit. Product description: fistula scissor, sharp and strong Used for fistula repair Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Method of sterilization: it should be able to withstand autoclave or steam sterilization. Blades: curved Length: 20cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended use: part of fistula kit. Product description: fistula scissor, sharp and strong Used for fistula repair Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Method of sterilization: it should be able to withstand autoclave or steam sterilization. Blades: curved Length: 20cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark. Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
467 | Medical & Surgical Instruments | SCISSOREPISIOTOMY | Scissors, Episiotomy, Braun, 14.5, angular | Scissors, Episiotomy, Braun, 14.5, angular | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is used for cutting through the delicate perineal skin and tissue during an episiotomy. The scissor slices effortlessly through the soft tissues to perform a clean incision to enlarge the birth canal to facilitate the passage of the foetal head. Product description: Surgical scissors, Braun, non-ratcheted Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with one guarded blade. Length: 145 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors for episiotomy. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: is used for cutting through the delicate perineal skin and tissue during an episiotomy. The scissor slices effortlessly through the soft tissues to perform a clean incision to enlarge the birth canal to facilitate the passage of the foetal head. Product description: Surgical scissors, Braun, non-ratcheted Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with one guarded blade. Length: 145 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors for episiotomy. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
468 | Medical & Surgical Instruments | SCISSORSBANDAGE14 | Scissors, Bandage, Lister, 14 cm, angular, sharp/blunt | Scissors, Bandage, Lister, 14 cm, angular, sharp/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Product description: Bandage scissor Supplied with: Material: Martensitic steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with a small foot on one of the blade tips (which allows them to be placed directly against patient's skin when cutting bandages). Length: 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Product description: Bandage scissor Supplied with: Material: Martensitic steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with a small foot on one of the blade tips (which allows them to be placed directly against patient's skin when cutting bandages). Length: 140 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
469 | Medical & Surgical Instruments | SCISSORSBANDAGE18 | Scissors, Bandage, Lister, 18 cm, angular, sharp / blunt | Scissors, Bandage, Lister, 18 cm, angular, sharp / blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Product description: Bandage scissor Supplied with: Material: Martensitic steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with a small foot on one of the blade tips (which allows them to be placed directly against patient's skin when cutting bandages). Length: 180 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Product description: Bandage scissor Supplied with: Material: Martensitic steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Angular blades, with a small foot on one of the blade tips (which allows them to be placed directly against patient's skin when cutting bandages). Length: 180 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For general use in nursing care and first aid; used for cutting bandages, dressings and clothing. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer). Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
470 | Medical & Surgical Instruments | SCISSORSBOYD17 | Scissors, Boyd, 17cm, double curved | Scissors, Boyd, 17cm, double curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Surgical dissecting scissors used in the fine dissection of tissues. Product description: Dissecting scissors Curved 17 cm sharp tipped or semi-sharp tipped S-shaped (double curved) tissue scissors. Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Surgical dissecting scissors used in the fine dissection of tissues. Product description: Dissecting scissors Curved 17 cm sharp tipped or semi-sharp tipped S-shaped (double curved) tissue scissors. Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Method of sterilization: it should be able to withstand autoclave or steam sterilization. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
471 | Medical & Surgical Instruments | SCISSORSDEAVER14S | Scissors, Deaver, 14cm, straight, sharp/blunt | Scissors, Deaver, 14cm, straight, sharp/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Operating scissors that have all the functionality of standard operating scissors, but also feature a delicate pattern. The tip configuration sharp/blunt allow these scissors to be used for sharp and blunt dissection during surgical procedures Product description: Long surgical dissecting scissors, Deaver, for basic surgical operations (sutures or for treating abscesses) and for fistula repair. Material: Martensitic stainless steel (quenched, magnetic steel) Steel composition: 0.35 to 0.40 % carbon; 14% to 15 % chromium Straight, with one pointed and one blunt end blade Length: Approx.140mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Operating scissors that have all the functionality of standard operating scissors, but also feature a delicate pattern. The tip configuration sharp/blunt allow these scissors to be used for sharp and blunt dissection during surgical procedures Product description: Long surgical dissecting scissors, Deaver, for basic surgical operations (sutures or for treating abscesses) and for fistula repair. Material: Martensitic stainless steel (quenched, magnetic steel) Steel composition: 0.35 to 0.40 % carbon; 14% to 15 % chromium Straight, with one pointed and one blunt end blade Length: Approx.140mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: For basic surgical operations. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol)(if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
472 | Medical & Surgical Instruments | SCISSORSGYNAECOL | Scissors, gynaecological, 20cm, curved, blunt/blunt | Scissors, gynaecological, 20cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Surgical scissors used For cutting the perineal skin and tissue (episiotomy)to facilitate the passage of the foetal head. Can be used also for decapitation of dead foetus. Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Curved, with blunt end blades. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used For cutting the perineal skin and tissue (episiotomy)to facilitate the passage of the foetal head. Can be used also for decapitation of dead foetus This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable or equivalent harmonised symbol. Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Surgical scissors used For cutting the perineal skin and tissue (episiotomy)to facilitate the passage of the foetal head. Can be used also for decapitation of dead foetus. Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Curved, with blunt end blades. Length: approx. 200mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Used For cutting the perineal skin and tissue (episiotomy)to facilitate the passage of the foetal head. Can be used also for decapitation of dead foetus This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Primary packaging: Unit of use. One (1) scissors in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable or equivalent harmonised symbol. Secondary packaging: Protected unit. Ten (10) scissors in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Over packaging: Packaging unit. Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
473 | Medical & Surgical Instruments | SCISSORSMAYO14C | Scissors, Mayo, 14cm, curved, blunt/blunt | Scissors, Mayo, 14cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 140mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
474 | Medical & Surgical Instruments | SCISSORSMAYO17C | Scissors, Mayo, 17cm, curved, blunt/blunt | Scissors, Mayo, 17cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 170mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 170mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
475 | Medical & Surgical Instruments | SCISSORSMAYO17S | Scissors, Mayo, 17 cm, straight, blunt/blunt | Scissors, Mayo, 17 cm, straight, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical scissors, Mayo, used for cutting sutures and dressings. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, straight, with blunt end blades. Length: approx. 170mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical scissors, Mayo, used for cutting sutures and dressings. Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, straight, with blunt end blades. Length: approx. 170mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
476 | Medical & Surgical Instruments | SCISSORSMAYO23C | Scissors, Mayo, 23cm, curved, blunt/blunt | Scissors, Mayo, 23cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 230mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Curved scissor is most commonly utilized for cutting or dissecting dense or deep tissue and the straight version is typically used for cutting surface tissue or sutures. Both may be used interchangeably depending on procedure and physician preference Product description: Surgical dissecting scissors, Mayo Material: Martensitic stainless steel (quenched, magnetic steel) composition: 0.40 % carbon; 14% chromium. Instrument that cuts by shearing Thin, curved, with blunt end blades Length: approx. 230mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used for non-delicate dissections and to cut sutures and dressings. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE MArk Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
477 | Medical & Surgical Instruments | SCISSORSMETZEN14C | Scissors, Metzenbaum/Nelson, 14cm, curved, blunt/blunt | Scissors, Metzenbaum/Nelson, 14cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.140 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.140 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
478 | Medical & Surgical Instruments | SCISSORSMETZEN18C | Scissors, Metzenbaum/Nelson, 18cm, curved, blunt/blunt | Scissors, Metzenbaum/Nelson, 18cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.180 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.180 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
479 | Medical & Surgical Instruments | SCISSORSMETZEN23C | Scissors, Metzenbaum/Nelson, 23cm, curved, blunt/blunt | Scissors, Metzenbaum/Nelson, 23cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.230 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
480 | Medical & Surgical Instruments | SCISSORSMETZEN24C | Scissors, Metzenbaum/Nelson, 24 cm, curved, blunt/blunt | Scissors, Metzenbaum/Nelson, 24 cm, curved, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.260 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: typically used for blunt dissection of delicate tissues. They are not appropriate for cutting sutures, drains, or heavy tissue Product description: Surgical dissecting scissors, Metzenbaum/Nelson, ideal for dissecting delicate tissues in confined places Material: Martensitic stainless steel (quenched, magnetic steel) composition :From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Thin, curved, with blunt end blades. Length: Approx.260 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical dissecting scissors used in the fine dissection of tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
481 | Medical & Surgical Instruments | SCISSORSSTICH | Scissors, stitch, 15cm | Scissors, stitch, 15cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to isolate and dissect sutures during various procedures. Product description: Stitch Scissors, with a special hook tip for sliding beneath and grasping sutures Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Straight upper blade. Hooked or curved outward lower blade. Blunt bottom blade. Length: 150 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Scissors for cutting sutures. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to isolate and dissect sutures during various procedures. Product description: Stitch Scissors, with a special hook tip for sliding beneath and grasping sutures Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From 14% to 15 % chromium. Straight upper blade. Hooked or curved outward lower blade. Blunt bottom blade. Length: 150 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use Instructions for use: Scissors for cutting sutures. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
482 | Medical & Surgical Instruments | SCISSORSSTICH18S | Scissors, suture stich, 18cm, straight, sharp/blunt | Scissors, suture stich, 18cm, straight, sharp/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in basic surgical operations to cut sutures or non delicate tissues. Product description: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues. Straight edge with one sharp and one blunt tipped blade. Length: 18cm Material: Martensitic steel (quenched, magnetic steel). Material: Martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used in basic surgical operations to cut sutures or non delicate tissues. Product description: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues. Straight edge with one sharp and one blunt tipped blade. Length: 18cm Material: Martensitic steel (quenched, magnetic steel). Material: Martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol) (if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
483 | Medical & Surgical Instruments | SCISSORSTHOREK | Scissors, Thorek, 20cm, curved | Scissors, Thorek, 20cm, curved | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Material: Martensitic stainless steel (quenched, magnetic steel) Full curve scissors, blunt angled tip, 90 degrees tip. Length: approx..20cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Material: Martensitic stainless steel (quenched, magnetic steel) Full curve scissors, blunt angled tip, 90 degrees tip. Length: approx..20cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
484 | Medical & Surgical Instruments | SCISSORSTHOREK18C | Scissors, Thorek, 18cm, curved, sharp/blunt | Scissors, Thorek, 18cm, curved, sharp/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Curved edge with one sharp and one blunt tipped blade. Length: 18cm Material: Martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Curved edge with one sharp and one blunt tipped blade. Length: 18cm Material: Martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.4, chromium 14%, silicon 0.5 % and manganese 0.4% Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Surgical scissors used in basic surgical operations to cut sutures or non delicate tissues This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
485 | Medical & Surgical Instruments | SCISSORSUMBILICAL | Scissors, umbilical, USA-model, 10 cm | Scissors, umbilical, USA-model, 10 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to cut, clamp and hold the cord after dissection. Cuts the umbilical cord before tying it off after delivery of the baby Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From14% to 15 % chromium. Length: Approx.100 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Scissors for cutting the umbilical cord after child birth. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to cut, clamp and hold the cord after dissection. Cuts the umbilical cord before tying it off after delivery of the baby Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel) composition: From 0.35 to 0.40 % carbon; From14% to 15 % chromium. Length: Approx.100 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: Scissors for cutting the umbilical cord after child birth. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
486 | Medical & Surgical Instruments | SCISSORSUTERINE | Scissors, uterine, Sims, curved, 20 cm, blunt/blunt | Scissors, uterine, Sims, curved, 20 cm, blunt/blunt | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:s a multi-purpose surgical scissor used for blunt dissection of superficial and deep tissues during gynecological and general surgeries. Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.35 to 0.40 % carbon; From14% to 15 % chromium Curved, with blunt end blades. Length:approx.. 200 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Uterine scissors, Sims. must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use:s a multi-purpose surgical scissor used for blunt dissection of superficial and deep tissues during gynecological and general surgeries. Product description: Gynecological Scissors Material: Martensitic stainless steel (quenched, magnetic steel). composition: From 0.35 to 0.40 % carbon; From14% to 15 % chromium Curved, with blunt end blades. Length:approx.. 200 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Uterine scissors, Sims. must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
487 | Medical & Surgical Instruments | SCISSTHOREK18C90 | Scissors, Thorek, 18cm, curved 90 degrees | Scissors, Thorek, 18cm, curved 90 degrees | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Material: Martensitic stainless steel (quenched, magnetic steel) Full curve scissors, blunt angled tip, 90 degrees tip. Length: approx..18cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for blunt dissection of soft tissues. used for cutting the final attachments of the uterus to the vagina during abdominal hysterectomy and in fistula repair surgery. Product description: Dissecting Scissor, have long shafts with short fully curved blades. Material: Martensitic stainless steel (quenched, magnetic steel) Full curve scissors, blunt angled tip, 90 degrees tip. Length: approx..18cm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use. Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
488 | Medical & Surgical Instruments | SCOOPUTERINESIMON | Scoop, uterine, Simon, 6mm, sharp | Scoop, uterine, Simon, 6mm, sharp | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used during biopsy procedures of the uterus for diagnostic as well as therapeutic procedures. Also, used to access and scrape a particular area of the lining of the uterus for its removal during a gynaecological procedure. Product description: Uterine Curette. A heavy, hand-held, manual surgical instrument with a spoon-like, sharp-edged distal end on a long shaft. Material: Martensitic stainless steel (quenched, magnetic steel) Shank type: Rigid. Blade shape: Oval form of a spoon. Blade edge type: Sharp Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium. Length: Shaft 250 mm. Width: Spoon 6 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used during biopsy procedures of the uterus for diagnostic as well as therapeutic procedures. Also, used to access and scrape a particular area of the lining of the uterus for its removal during a gynaecological procedure. Product description: Uterine Curette. A heavy, hand-held, manual surgical instrument with a spoon-like, sharp-edged distal end on a long shaft. Material: Martensitic stainless steel (quenched, magnetic steel) Shank type: Rigid. Blade shape: Oval form of a spoon. Blade edge type: Sharp Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.20% carbon; 13% chromium. Length: Shaft 250 mm. Width: Spoon 6 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
489 | Medical & Surgical Instruments | SPECULUMAUVARD | Speculum, vaginal, Auvard, retractor, weighted,8 x 3.8cm | Speculum, vaginal, Auvard, retractor, weighted,8 x 3.8cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for access to the vaginal wall, cervix, uterus and uterine contents. It is useful for retracting the posterior vaginal wall for therapeutic gynecological procedures Product description: Vaginal retractor/speculum, Auvard Wide vaginal speculum with a weight which fits on the shank of the blade. The weight should be soldered/ welded rather than screwed. Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Blade length: approx. 80mm. Blade width: approx. 38mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for access to the vaginal wall, cervix, uterus and uterine contents. It is useful for retracting the posterior vaginal wall for therapeutic gynecological procedures Product description: Vaginal retractor/speculum, Auvard Wide vaginal speculum with a weight which fits on the shank of the blade. The weight should be soldered/ welded rather than screwed. Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. Blade length: approx. 80mm. Blade width: approx. 38mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
490 | Medical & Surgical Instruments | SPECULUMAUVARD1 | Speculum, vaginal, Auvard, retractor, weighted, 10 x 4 cm | Speculum, vaginal, Auvard, retractor, weighted, 10 x 4 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for access to the vaginal wall, cervix, uterus and uterine contents. It is useful for retracting the posterior vaginal wall for therapeutic gynecological procedures Product description: Vaginal retractor/speculum, Auvard Wide vaginal speculum with a weight which fits on the shank of the blade. The weight should be soldered/ welded rather than screwed. Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. - Blade Length: 100mm. - Blade Width: 40mm Weight 2.5 lb. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for access to the vaginal wall, cervix, uterus and uterine contents. It is useful for retracting the posterior vaginal wall for therapeutic gynecological procedures Product description: Vaginal retractor/speculum, Auvard Wide vaginal speculum with a weight which fits on the shank of the blade. The weight should be soldered/ welded rather than screwed. Material: Austenitic stainless steel (non-quenched, non-magnetic steel) composition: 18% to 20% chromium; 8 to 10% nickel. - Blade Length: 100mm. - Blade Width: 40mm Weight 2.5 lb. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument - Name and address of supplier (manufacturer) Symbols used according ISO 15223 CE Mark Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
491 | Medical & Surgical Instruments | SPECULUMSIMSL | Speculum, vaginal, Sims, retractor, double-ended, large | Speculum, vaginal, Sims, retractor, double-ended, large | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to visualize the female reproductive tract. It can be rotated 360 degrees, giving better visualization of the cervix and vaginal walls. Commonly used in diagnostic gynecological procedures Product description: Vaginal speculum, Sims, Single blade speculum with blunt lateral edges Hollow blade double-ended speculum Blade: one side 8x3.5cm, other 9x4cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 to 18% chromium; 2 to 3 % molybdenum; 10 to 14% nickel; 1% silicon; 2% manganese (AISI 316/316L). Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Speculum for examining the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark.Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to visualize the female reproductive tract. It can be rotated 360 degrees, giving better visualization of the cervix and vaginal walls. Commonly used in diagnostic gynecological procedures Product description: Vaginal speculum, Sims, Single blade speculum with blunt lateral edges Hollow blade double-ended speculum Blade: one side 8x3.5cm, other 9x4cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 to 18% chromium; 2 to 3 % molybdenum; 10 to 14% nickel; 1% silicon; 2% manganese (AISI 316/316L). Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Speculum for examining the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark.Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
492 | Medical & Surgical Instruments | SPECULUMSIMSM | Speculum, vaginal, Sims, retractor, double-ended, medium | Speculum, vaginal, Sims, retractor, double-ended, medium | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to visualize the female reproductive tract. It can be rotated 360 degrees, giving better visualization of the cervix and vaginal walls. Commonly used in diagnostic gynecological procedures Product description: Vaginal speculum, Sims, Single blade speculum with blunt lateral edges Hollow blade double-ended speculum Blade: one side 7.5x3cm, other 8x3.5cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 to 18% chromium; 2 to 3 % molybdenum; 10 to 14% nickel; 1% silicon; 2% manganese (AISI 316/316L). Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Speculum for examining the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark.Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: to visualize the female reproductive tract. It can be rotated 360 degrees, giving better visualization of the cervix and vaginal walls. Commonly used in diagnostic gynecological procedures Product description: Vaginal speculum, Sims, Single blade speculum with blunt lateral edges Hollow blade double-ended speculum Blade: one side 7.5x3cm, other 8x3.5cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 to 18% chromium; 2 to 3 % molybdenum; 10 to 14% nickel; 1% silicon; 2% manganese (AISI 316/316L). Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: Speculum for examining the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark.Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
493 | Medical & Surgical Instruments | SPECULVAGCUSCOL | Speculum, vaginal, Cusco, 11.5 x 3.5 cm | Speculum, vaginal, Cusco, 11.5 x 3.5 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenchedVnon-magnetic steel). Double beaked,vaginal speculum self-retaining. Blade length: approximately 115mm. Blade width: approximately 35 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenchedVnon-magnetic steel). Double beaked,vaginal speculum self-retaining. Blade length: approximately 115mm. Blade width: approximately 35 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
494 | Medical & Surgical Instruments | SPECULVAGCUSCOM | Speculum, vaginal, Cusco, 9.5 x 3.5 cm | Speculum, vaginal, Cusco, 9.5 x 3.5 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From 10.5% - 13.5% nickel; From 2 % to 2.5% molybdenum. Double beaked,vaginal speculum self-retaining. Blade length: approximately 95mm. Blade width: approximately 35 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From 10.5% - 13.5% nickel; From 2 % to 2.5% molybdenum. Double beaked,vaginal speculum self-retaining. Blade length: approximately 95mm. Blade width: approximately 35 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
495 | Medical & Surgical Instruments | SPECULVAGCUSCOS | Speculum, vaginal, Cusco, 7.5 x 2 cm | Speculum, vaginal, Cusco, 7.5 x 2 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum. Double beaked,vaginal speculum self-retaining. Blade length: approximately 75mm. Blade width: approximately 20 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used to visualize the vagina and cervix during DNC, DNE as well as IUCD insertion and regular pelvic examination. Product description: aginal speculum, Cusco Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum. Double beaked,vaginal speculum self-retaining. Blade length: approximately 75mm. Blade width: approximately 20 mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
496 | Medical & Surgical Instruments | SPECULVAGGRAVESL | Speculum, vaginal, Graves, 115 x 35mm | Speculum, vaginal, Graves, 115 x 35mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 115mm. Blade width: approximately 35mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 115mm. Blade width: approximately 35mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
497 | Medical & Surgical Instruments | SPECULVAGGRAVESM | Speculum, vaginal, Graves, 95 x 35mm | Speculum, vaginal, Graves, 95 x 35mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 95mm. Blade width: approximately 35mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 95mm. Blade width: approximately 35mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
498 | Medical & Surgical Instruments | SPECULVAGGRAVESS | Speculum, vaginal, Graves, 75 x 20mm | Speculum, vaginal, Graves, 75 x 20mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 75mm. Blade width: approximately 20mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: Opens the vagina for to visualize the cervix, uterus and uterine contents as well as access the vaginal walls Product description: Vaginal speculum, Graves Material: Austenitic steel (non-quenched, non-magnetic steel). composition: From 16.5% to18.5% chromium; From10.5% - 13.5% nickel; From2 % to 2.5% molybdenum Double beaked,vaginal speculum self-retaining. Blade length: approximately 75mm. Blade width: approximately 20mm. Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Supplied with: Manufacturer's instruction for use and care including cleaning, disinfection and sterilization instructions Instructions for use: To examine the walls of the vagina and cervix of the uterus. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) speculum in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
499 | Medical & Surgical Instruments | TOWELCLAMPBACKH | Clamp, towel, Backhaus, 120 mm | Clamp, towel, Backhaus, 120 mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for attaching towels around the surgical incision, either with or without piercing the skin. Product description: Towel clamp, Backhaus, Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.1-1%, chromium 12-14%, molybdenum 0.2-1%, silicon 0.5-1% and manganese 0.4-2%. Spring-type Hard ratchet, fine, single tooth in each jaw. Lockable. Length: approx. 120-125 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to attach towels around the surgical field, with or without piercing the skin This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) clamp in a plastic bag. Labelling on the primary packaging: Labelling as per regulations applied to the device. Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark Secondary packaging: Protected unit. Four (4) Backhaus towel clamps in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for attaching towels around the surgical incision, either with or without piercing the skin. Product description: Towel clamp, Backhaus, Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.1-1%, chromium 12-14%, molybdenum 0.2-1%, silicon 0.5-1% and manganese 0.4-2%. Spring-type Hard ratchet, fine, single tooth in each jaw. Lockable. Length: approx. 120-125 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to attach towels around the surgical field, with or without piercing the skin This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) clamp in a plastic bag. Labelling on the primary packaging: Labelling as per regulations applied to the device. Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark Secondary packaging: Protected unit. Four (4) Backhaus towel clamps in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
500 | Medical & Surgical Instruments | TOWELCLAMPBACKH9 | Clamp, towel, Backhaus, 90 mm | Clamp, towel, Backhaus, 90 mm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for attaching towels around the surgical incision, either with or without piercing the skin. Product description: Towel clamp, Backhaus, Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.1-1%, chromium 12-14%, molybdenum 0.2-1%, silicon 0.5-1% and manganese 0.4-2%. Spring-type Hard ratchet, fine, single tooth in each jaw. Lockable. Length: approx. 90 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to attach towels around the surgical field, with or without piercing the skin This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) clamp in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: for attaching towels around the surgical incision, either with or without piercing the skin. Product description: Towel clamp, Backhaus, Material: martensitic stainless steel (quenched, magnetic steel) composition of carbon 0.1-1%, chromium 12-14%, molybdenum 0.2-1%, silicon 0.5-1% and manganese 0.4-2%. Spring-type Hard ratchet, fine, single tooth in each jaw. Lockable. Length: approx. 90 mm Supplied non sterile, Hygenically packed for medical use after sterilization. Reusable: Can withstand repeated reprocessing including steam sterilization, including manual brushing and the use of disinfectant chemicals Instructions for use: Used to attach towels around the surgical field, with or without piercing the skin This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Supplied with: Manufacturer's instruction for use. Packaging and labelling: Primary packaging: Unit of use. One (1) clamp in a plastic bag. Labelling on the primary packaging: Name and/or trademark and address of the manufacturer.. Manufacturer's product reference. Type of product and main characteristics. Lot number prefixed by the word "LOT"(or equivalent harmonised symbol, if applicable). Information for particular storage conditions(temperature, pressure, light, humidity, etc., as appropriate (or equivalent harmonised symbol.) CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
501 | Medical & Surgical Instruments | UTERINESOUND30CM | Sound, uterine, Martin, malleable, 30 cm | Sound, uterine, Martin, malleable, 30 cm | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for determining the position of the uterus. Commonly used for IUCD insertion, dilatation and curettage as well as other gynecological procedures Product description: Uterine sound, malleable. Graduated in cm, shaft made of malleable metal, distal end bulbous, mounted on a handle. Length: approximately 30cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 - 18% chromium; 10 - 14% nickel; 2 - 3% molybdenum. Supplied with: Manufacturer's instruction for use and care. Instructions for use: risk of uterus perforation ! This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) item in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001 | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Intended purpose of use: used for determining the position of the uterus. Commonly used for IUCD insertion, dilatation and curettage as well as other gynecological procedures Product description: Uterine sound, malleable. Graduated in cm, shaft made of malleable metal, distal end bulbous, mounted on a handle. Length: approximately 30cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 16 - 18% chromium; 10 - 14% nickel; 2 - 3% molybdenum. Supplied with: Manufacturer's instruction for use and care. Instructions for use: risk of uterus perforation ! This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Packaging & Labelling: One (1) item in a plastic bag. Symbols used according ISO 15223 CE mark, Labelling as per regulations applied to the device. Shall include name of the device, name and address of the manufacturer, product code or manufacturers reference number etc. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards (as per latest revision of standards): ISO 14971:2007 Medical Devices- Application of risk management to medical devices ISO 7153-1:1991/Amd1:1999 Surgical instruments- Metallic material – Part 1 Stainless Steel ISO 13402:1995 Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure ISO 17664:2004 Sterilization of medical devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices ISO 7741 Instruments for surgery -- Scissors and shears -- General requirements and test methods Environmental requirements: ISO 14001
502 | Medical & Surgical Instruments | VACUUMEXTR_BD50&60 | Vacuum extractor, Bird, w/anterior & posterior cups, manual | Vacuum extractor, Bird, w/anterior & posterior cups, manual | Intended purpose of use: to facilitate the delivery of a foetus during vaginal childbirth by enabling traction to be applied to the foetal head by means of a suction cup that is attached to the scalp. It is typically used for dystocia, uterine inertia, maternal exhaustion, maternal/foetal distress. Product Description: Hand operated Vacuum extractor system, an assembly of devices, typically consists of a cup with a soft inner liner to prevent tissue damage, a hand-operated pump, and a handle used to apply traction. Comes with 50mm and 60mm Bird anterior cups made of stainless steel and silicone and one 50mm posterior bird cup. Hand-operated suction pump creates required vacuum. Vacuum range, adjustable: from 0 to -100kPa/-800mmHg. Gauge dial, minimal graduation: 5kPa/40mmHg. Accuracy vacuum gauge: ± 2.5kPa. All parts of the vacuum extractor can be disassembled, relevant parts can be autoclaved at 121 C. Requires no other than routine maintenance, manageable at the level of the delivery rooms. System is supplied as a complete set consisting of: 2X Bird type anterior suction/extraction cup (1x 50mm and 1x 60mm diameter) rounded edges. 1X Bird type posterior suction/extraction cup (1x50 mm diameter) rounded edges. 2x Extraction handles for Bird type anterior suction/extraction cup. Suction/extraction cups and handle: stainless steel, Chrome nickel plated. 2x Soft suction/extraction cup (1x50 and 1x60mm opening diameter), made of medical grade silicones 1 Hand operated aspiration device including: 1x Vacuum/collection bottle, plastic 500ml. 1x Cover plug for vacuum/collection bottle, with 3 holes (2 connector + 1 vacuum gauge). 1x Connector angled, chrome plated. 1x Connector angled with screw valve (vacuum release), chrome plated, with cap and gasket. 1x Vacuum gauge with straight connector, chrome plated. 1x Basket, epoxy coated, supporting vacuum/collection bottle and the suction pump, with large hook allowing to fit it to the delivery table. 2x Tube, silicone, transparent, inner/outer diameter approx. 6/11mm, length 50 and 150cm. Intended use: Manual vacuum extractor system to assist vaginal deliveries in a delivery room setting with stainless steel anterior and posterior cups. Hospital equipment for delivery room. The vacuum extractor must be cleaned and disinfected after each use. In view of the fact that the vacuum extractor is used in situations of acute need, it must always be readily available and in good working condition. The vacuum extractor must be operated by adequately trained staff only. It is recommended to follow the manufacturer's instructions at all times. Supplied with: 1x Spare vacuum/collection bottle, 500ml. 1x Spare tube, silicone, transparent, inner/outer diameter approx. 6/11mm, length 300cm. 1x Spare gasket for the screw valve (vacuum release). 20x Polypropylene bottom plate, for stainless steel suction/extraction cup, 10 for each size, 50 and 60mm. 1x List of accessories and parts, each identified with their product reference. 1x Instructions for use and maintenance, and contact details for repair service, in English, French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Vendor specific. Packaging and labelling: Packaging: One (1) manual Bird vacuum system with accessories in protective packaging. Labels include: Name and/ or trade name, and address of the manufacturer. Manufacturer's product code / reference. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 10079-2: Medical suction equipment — Part 2: Manually powered suction equipment. ISO 3529-1: Vacuum technology — Vocabulary — Part 1: General terms. ISO 3529-2: Vacuum technology — Vocabulary — Part 2: Vacuum pumps and related terms. ISO 21360-1: Vacuum technology - Standard methods for measuring vacuum-pump performance - Part 1: General description. ISO 17664: Sterilization devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO 14001: Environmental management systems. | Intended purpose of use: to facilitate the delivery of a foetus during vaginal childbirth by enabling traction to be applied to the foetal head by means of a suction cup that is attached to the scalp. It is typically used for dystocia, uterine inertia, maternal exhaustion, maternal/foetal distress. Product Description: Hand operated Vacuum extractor system, an assembly of devices, typically consists of a cup with a soft inner liner to prevent tissue damage, a hand-operated pump, and a handle used to apply traction. Comes with 50mm and 60mm Bird anterior cups made of stainless steel and silicone and one 50mm posterior bird cup. Hand-operated suction pump creates required vacuum. Vacuum range, adjustable: from 0 to -100kPa/-800mmHg. Gauge dial, minimal graduation: 5kPa/40mmHg. Accuracy vacuum gauge: ± 2.5kPa. All parts of the vacuum extractor can be disassembled, relevant parts can be autoclaved at 121 C. Requires no other than routine maintenance, manageable at the level of the delivery rooms. System is supplied as a complete set consisting of: 2X Bird type anterior suction/extraction cup (1x 50mm and 1x 60mm diameter) rounded edges. 1X Bird type posterior suction/extraction cup (1x50 mm diameter) rounded edges. 2x Extraction handles for Bird type anterior suction/extraction cup. Suction/extraction cups and handle: stainless steel, Chrome nickel plated. 2x Soft suction/extraction cup (1x50 and 1x60mm opening diameter), made of medical grade silicones 1 Hand operated aspiration device including: 1x Vacuum/collection bottle, plastic 500ml. 1x Cover plug for vacuum/collection bottle, with 3 holes (2 connector + 1 vacuum gauge). 1x Connector angled, chrome plated. 1x Connector angled with screw valve (vacuum release), chrome plated, with cap and gasket. 1x Vacuum gauge with straight connector, chrome plated. 1x Basket, epoxy coated, supporting vacuum/collection bottle and the suction pump, with large hook allowing to fit it to the delivery table. 2x Tube, silicone, transparent, inner/outer diameter approx. 6/11mm, length 50 and 150cm. Intended use: Manual vacuum extractor system to assist vaginal deliveries in a delivery room setting with stainless steel anterior and posterior cups. Hospital equipment for delivery room. The vacuum extractor must be cleaned and disinfected after each use. In view of the fact that the vacuum extractor is used in situations of acute need, it must always be readily available and in good working condition. The vacuum extractor must be operated by adequately trained staff only. It is recommended to follow the manufacturer's instructions at all times. Supplied with: 1x Spare vacuum/collection bottle, 500ml. 1x Spare tube, silicone, transparent, inner/outer diameter approx. 6/11mm, length 300cm. 1x Spare gasket for the screw valve (vacuum release). 20x Polypropylene bottom plate, for stainless steel suction/extraction cup, 10 for each size, 50 and 60mm. 1x List of accessories and parts, each identified with their product reference. 1x Instructions for use and maintenance, and contact details for repair service, in English, French and Spanish. Accessories / Spare Parts / Consumables (available but not supplied): Vendor specific. Packaging and labelling: Packaging: One (1) manual Bird vacuum system with accessories in protective packaging. Labels include: Name and/ or trade name, and address of the manufacturer. Manufacturer's product code / reference. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol). Expiry date: month and year. CE marking and Notified Body number. Symbols used according ISO 15223 and EN 980. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards (current version of the following standards): ISO 14971: Medical Devices- Application of risk management to medical devices. ISO 10079-2: Medical suction equipment — Part 2: Manually powered suction equipment. ISO 3529-1: Vacuum technology — Vocabulary — Part 1: General terms. ISO 3529-2: Vacuum technology — Vocabulary — Part 2: Vacuum pumps and related terms. ISO 21360-1: Vacuum technology - Standard methods for measuring vacuum-pump performance - Part 1: General description. ISO 17664: Sterilization devices -- Information to be provided by the manufacturer for the processing of re-sterilizable medical devices. Environmental requirements: ISO 14001: Environmental management systems.
503 | Medical & Surgical Instruments | VACUUMEXTRCTRBIRD | Vacuum extractor, Bird, manual, complete set | Vacuum extractor, Bird, manual, complete set | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Manually operated by 1 person. Handle = hand-operated suction device. Manual aspiration pump with finger activated vacuum release valve. Vacuum indicator (mmHg, in Hg, kPa, bar). Screw thread to attach the fluid trap, non-autoclavable. Anterior suction cup autoclavable at 121°C. Diameter: 5cm. Height: 2cm. Plastic fluid trap cup with lid. Containing a disk attached with wire. O-ring on the edge autoclavable at 121°C. Aspiration tube. With inner tubing connecting the fluid trap to the cup. Autoclavable at 121°C. Length: 19cm. Supplied with: 2 pumps/handles. 1 cup. 2 belt disks. 1 trap assembly (lid and cup). 2 tubes. 2 belts. 2 O-rings for Trap 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish) with DVD Accessories/Spare parts/Consumables: Spare parts: 1 pumps/handles. 1 cup. 1 belt disks. 1 trap assembly (lid and cup). 1 tubes. 1 belts. 1 O-rings for Trap Instructions for use: Place the cup safely and correctly over the flexion point of the fetal head. Exert the traction in line with the pelvic axis to correct malpositions of the foetal head so the optimal birth canal diameter and autorotation will occur naturally. The traction is delivered by pressing the thumb- or finger-activated vacuum release valve. This detachable Fluid Trap and vacuum stem must be cleaned, disinfected after each use, and sterilised in a steam steriliser. (PalmPump must not be included in this sterilization.)(source : http://www.vaccareusableomnicup.com/) Packaging & Labelling: Primary packaging Unit presentation: 5 (five) OmniCup w/traction force indicator + box with Manufacturer's instruction for use, spare parts and accessories. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) - Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485:2016 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: Not use PVC polymers | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Vacuum extractor system to assist vaginal deliveries. Hand-operated suction pump supports vacuum extraction. Vacuum range, adjustable: from 0 to -85kPa/-638mmHg Gauge dial, minimal graduation: 5kPa/40mmHg Accuracy vacuum gauge: ± 2.5kPa All parts of the vacuum extractor can be disassembled. All parts can be disassembled and are autoclavable at 121°C Requires no other than routine maintenance, manageable at the level of delivery rooms. Supplied as a completed set consisting of: 1 x Vacuum pump 2 x Bird type anterior suction cup, stainless steel: diameter 50 and 60 mm rounded edges 2 x Soft vacuum extractor cups "silc-cup": diameter 50mm and 60mm 2 x Extraction handle, stainless steel 1 x Cover plug for vacuum collection bottle, with 3 holes (2 connectors +1 vacuum gauge) 1 x Connector, angled, chrome plated 1 x Connector angled with screw valve, chrome plated screw, cap and gasket 1 x Vacuum gauge (0 to -100 kPa/- 800mmHg) with straight connector, chrome plated 1 x Vacuum collection bottle, plastic 500ml 1 x Basket, epoxy coated, supporting vacuum collection bottle and the vacuum pump, with a large hook allowing it to fit to the delivery table. 2 x Tube, silicone transparent, inner/outer diameter approx.: 6/11mm, respective lengths: 50 and 150cm Accessories/Spare parts/Consumables: 1 x Spare vacuum bottle, plastic, 500ml. 1 x Spare tube, transparent silicone, inner/outer diameter: 6/11mm, length: 300cm. 1 x Spare gasket for the screw valve (vacuum release). 20 x Polypropylene bottom plate, for stainless steel suction/extraction cup, 10 for each size, 50 and 60mm. Instructions for use: Vacuum extraction should only be performed or supervised by trained and experience physicians. The manufacturer does not provide clinical training on vacuum extraction. The instructions for use included with the devices provides the necessary details for physicians trained on carrying out vacuum extractions not familiar with this specific device. Packaging & Labelling: Primary packaging The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) - Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note: For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device Safety & product Standards: Must comply with the following standards: ISO 13485:2016 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process - EN ISO 10079-2:2009 Medical suction equipment - Part 2: Manually powered suction equipment. - ISO 10993-1:2018 Biological evaluation of medical devices — Part 1: Evaluation and testing within a risk management process. - ISO 10993-5:2009 Biological evaluation of medical devices. Tests for in vitro cytotoxicity Environmental requirements: Not use PVC polymers
504 | Medical & Surgical Instruments | VACUUMEXTRCTRVAC | Vacuum extractor, Vacca OmniCup, hand-held, 1 person, reusable | Vacuum extractor, Vacca OmniCup, hand-held, 1 person, reusable | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Manually operated by 1 person. Handle = hand-operated suction device. Manual aspiration pump with finger activated vacuum release valve. Vacuum indicator (mmHg, in Hg, kPa, bar). Screw thread to attach the fluid trap, non-autoclavable. Anterior suction cup autoclavable at 121°C. Diameter: 5cm. Height: 2cm. Plastic fluid trap cup with lid. Containing a disk attached with wire. O-ring on the edge autoclavable at 121°C. Aspiration tube. With inner tubing connecting the fluid trap to the cup. Autoclavable at 121°C. Length: 19cm. Supplied with: 2 pumps/handles. 1 cup. 2 belt disks. 1 trap assembly (lid and cup). 2 tubes. 2 belts. 2 O-rings for Trap 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish) with DVD Accessories/Spare parts/Consumables: Spare parts: 1 pumps/handles. 1 cup. 1 belt disks. 1 trap assembly (lid and cup). 1 tubes. 1 belts. 1 O-rings for Trap Instructions for use: Place the cup safely and correctly over the flexion point of the fetal head. Exert the traction in line with the pelvic axis to correct malpositions of the foetal head so the optimal birth canal diameter and autorotation will occur naturally. The traction is delivered by pressing the thumb- or finger-activated vacuum release valve. This detachable Fluid Trap and vacuum stem must be cleaned, disinfected after each use, and sterilised in a steam steriliser. (PalmPump must not be included in this sterilization.)(source : http://www.vaccareusableomnicup.com/) Packaging & Labelling: Primary packaging Unit presentation: 5 (five) OmniCup w/traction force indicator + box with Manufacturer's instruction for use, spare parts and accessories. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) - Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: Not use PVC polymers | IMPORTANT NOTE: These technical specifications are generic and may vary depending on the selected Vendor. Product description: Manually operated by 1 person. Handle = hand-operated suction device. Manual aspiration pump with finger activated vacuum release valve. Vacuum indicator (mmHg, in Hg, kPa, bar). Screw thread to attach the fluid trap, non-autoclavable. Anterior suction cup autoclavable at 121°C. Diameter: 5cm. Height: 2cm. Plastic fluid trap cup with lid. Containing a disk attached with wire. O-ring on the edge autoclavable at 121°C. Aspiration tube. With inner tubing connecting the fluid trap to the cup. Autoclavable at 121°C. Length: 19cm. Supplied with: 2 pumps/handles. 1 cup. 2 belt disks. 1 trap assembly (lid and cup). 2 tubes. 2 belts. 2 O-rings for Trap 1 x Instructions for assembly, use and maintenance, and list of accessories and spare parts with product reference code, in 3 languages (English, French and Spanish) with DVD Accessories/Spare parts/Consumables: Spare parts: 1 pumps/handles. 1 cup. 1 belt disks. 1 trap assembly (lid and cup). 1 tubes. 1 belts. 1 O-rings for Trap Instructions for use: Place the cup safely and correctly over the flexion point of the fetal head. Exert the traction in line with the pelvic axis to correct malpositions of the foetal head so the optimal birth canal diameter and autorotation will occur naturally. The traction is delivered by pressing the thumb- or finger-activated vacuum release valve. This detachable Fluid Trap and vacuum stem must be cleaned, disinfected after each use, and sterilised in a steam steriliser. (PalmPump must not be included in this sterilization.)(source : http://www.vaccareusableomnicup.com/) Packaging & Labelling: Primary packaging Unit presentation: 5 (five) OmniCup w/traction force indicator + box with Manufacturer's instruction for use, spare parts and accessories. The following should appear on the packaging: - Designation of the instrument. - Name and address of supplier (manufacturer) - Extra information required: Number of units per secondary packaging. Symbols used according ISO 15223 CE Mark Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with the following standards: ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 10993-1:2009 Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process Environmental requirements: Not use PVC polymers
505 | Medical & Surgical Instruments | New | Forceps, Dressing, Standard Pattern, 12.7 cm | Forceps, Dressing, Standard Pattern, 12.7 cm | GENERAL DESCRIPTION Forceps, dressing, standard, 127 mm, straight. INTENDED USE Dressing forceps for gripping, dissecting tissue and coagulation of vessels. Used in surgery and nursing activities. Forceps without teeth are used for dissecting delicate tissues. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: dissecting forceps, springy See Martin Catalogue no. 24 Ref: 12 140 25. Straight without teeth. Flexible arms. Serrated jaws Length:12.7cm. Material: Martensitic steel (quenched, magne tic steel). Martensitic steel composition: 0.20% carbon, 13% chromium. Reusable. Packaging and labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark" Instructions for use: Used for dressing in the abdominal area / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel Environmental conditions: Storage and transport conditions: 5°C to 40°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forceps, dressing, standard, 127 mm, straight. INTENDED USE Dressing forceps for gripping, dissecting tissue and coagulation of vessels. Used in surgery and nursing activities. Forceps without teeth are used for dissecting delicate tissues. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: dissecting forceps, springy See Martin Catalogue no. 24 Ref: 12 140 25. Straight without teeth. Flexible arms. Serrated jaws Length:12.7cm. Material: Martensitic steel (quenched, magne tic steel). Martensitic steel composition: 0.20% carbon, 13% chromium. Reusable. Packaging and labelling: One (1) forceps in a plastic bag. Symbols used according ISO 15223 CE mark" Instructions for use: Used for dressing in the abdominal area / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. ISO 7153-1:1991 Surgical instruments --Metallic materials --Part 1: Stainless steel Environmental conditions: Storage and transport conditions: 5°C to 40°C. Humidity: 30% RH to 70% RH.
506 | Medical & Surgical Instruments | New | Forceps, Intestinal, Allis, Delicate, 15 cm (5x6 teeth) | Forceps, Intestinal, Allis, Delicate, 15 cm (5x6 teeth) | GENERAL DESCRIPTION Forceps, tissue seizing, Allis, 150 mm. INTENDED USE For grasping organs, soft or heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: tissue grasping forceps, springy. See Martin Catalogue no. 24 Ref: 30 132 15. Precise adjustment of the teeth. Hard ratchet, lockable. Length: 15 cm. Tip configuration :5x6 teeth Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Primary packaging: Unit of use. One (1) forceps in a plastic bag. Secondary packaging: Protected unit. Ten (10) forceps in a box. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Information for particular storage conditions (temperature, pressure , light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Any symbol used on the labelling should be in accordance of the ISO 15223-1: 2016: If the item has a CE marking (or equivalent), the CE logo (or equivalent) must be visible on the packaging. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forceps, tissue seizing, Allis, 150 mm. INTENDED USE For grasping organs, soft or heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: tissue grasping forceps, springy. See Martin Catalogue no. 24 Ref: 30 132 15. Precise adjustment of the teeth. Hard ratchet, lockable. Length: 15 cm. Tip configuration :5x6 teeth Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Primary packaging: Unit of use. One (1) forceps in a plastic bag. Secondary packaging: Protected unit. Ten (10) forceps in a box. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Information for particular storage conditions (temperature, pressure , light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable(or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Any symbol used on the labelling should be in accordance of the ISO 15223-1: 2016: If the item has a CE marking (or equivalent), the CE logo (or equivalent) must be visible on the packaging. Instructions for use: Used to grip soft tissue This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
507 | Medical & Surgical Instruments | New | Richardson-Eastman retractors, small size (for suprapubic procedure) | Retractor, double-ended, Farabeuf, 120 mm, pair | GENERAL DESCRIPTION Retractor, double-ended, Farabeuf, 120 mm, pair. INTENDED USE To retract skin, fatty tissue, muscles or viscera during superficial surgery. Suprapubic means that the surgery is done through an incision in your lower abdomen, above your pubic bone. An incision is made in your bladder, and the center of your prostate gland is removed. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function instrument: retractor with handle. See Martin Catalogue no. 24 Ref: 15 663 00. Double-ended retractor. Slightly curved blade at each end. Smooth blunt blades with slight upward lip at the end. Small: approx. 12 cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 18 to 20% chromium, 8 to 10% nickel. Supplied in pairs. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Retractor, double-ended, Farabeuf, 120 mm, pair. INTENDED USE To retract skin, fatty tissue, muscles or viscera during superficial surgery. Suprapubic means that the surgery is done through an incision in your lower abdomen, above your pubic bone. An incision is made in your bladder, and the center of your prostate gland is removed. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function instrument: retractor with handle. See Martin Catalogue no. 24 Ref: 15 663 00. Double-ended retractor. Slightly curved blade at each end. Smooth blunt blades with slight upward lip at the end. Small: approx. 12 cm. Material: Austenitic steel (non-quenched, non-magnetic steel). Austenitic steel composition: 18 to 20% chromium, 8 to 10% nickel. Supplied in pairs. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
508 | Medical & Surgical Instruments | New | Army-Navy retractors, double-ended (for subumbilical procedure) | Army-Navy retractors, double-ended (for subumbilical procedure) | GENERAL DESCRIPTION Army Navy Retractor is used in plastic surgery to expose surface layers of skin. Army Navy is a surgical instrument with a double-ended fenestrated handle with a 90-degree angled set of two blades. INTENDED USE A retractor is a surgical instrument that is used to hold back underlying organs and tissues, so that body parts under the incision may be accessed without any fear of damage to the organs lying at the above levels. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function instrument: retractor with handle. Finish: Satin Franchise: Surgical Instruments Grade: OR Item Type: Retractors Length - Overall (mm): 215 Length 1 - Tip/Jaw (mm): 45.5 Physician/Developer: Army-Navy Shape: Right-Angled Therapeutic Specialty: Rhinology Usage: Reusable Width 1 - Tip/Jaw (mm): 15 Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Army Navy Retractor is used in plastic surgery to expose surface layers of skin. Army Navy is a surgical instrument with a double-ended fenestrated handle with a 90-degree angled set of two blades. INTENDED USE A retractor is a surgical instrument that is used to hold back underlying organs and tissues, so that body parts under the incision may be accessed without any fear of damage to the organs lying at the above levels. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function instrument: retractor with handle. Finish: Satin Franchise: Surgical Instruments Grade: OR Item Type: Retractors Length - Overall (mm): 215 Length 1 - Tip/Jaw (mm): 45.5 Physician/Developer: Army-Navy Shape: Right-Angled Therapeutic Specialty: Rhinology Usage: Reusable Width 1 - Tip/Jaw (mm): 15 Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
509 | Medical & Surgical Instruments | New | Forceps, Artery, Kelly, Straight, 14 cm | Forceps, Artery, Kelly, Straight, 14 cm | GENERAL DESCRIPTION Forceps, artery, Kelly, 140 mm, Straight INTENDED USE For haemostasis, clamping of large blood vessels or manipulating heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function Pressure force instrument: haemostatic forceps, springy. Flexible arms. Variable settings of the ratchet, lockable. Adjustment of the jaws. Length: 14 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forceps, artery, Kelly, 140 mm, Straight INTENDED USE For haemostasis, clamping of large blood vessels or manipulating heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function Pressure force instrument: haemostatic forceps, springy. Flexible arms. Variable settings of the ratchet, lockable. Adjustment of the jaws. Length: 14 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
510 | Medical & Surgical Instruments | New | Forceps mosquito, delicate, curved, 14 cm | Forceps mosquito, delicate, curved, 14 cm | GENERAL DESCRIPTION Forceps, artery, Halsted-Mosquito, 140 mm, curved. INTENDED USE For haemostasis, clamping of large blood vessels or manipulating heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function Pressure force instrument: haemostatic forceps, springy. Flexible arms. Variable settings of the ratchet, lockable. Adjustment of the jaws. Length: 14 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. S9910000 - Surg.inst., abdominal /SET S9910001 - Surg.inst., basic surgery /SET Product attributes Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forceps, artery, Halsted-Mosquito, 140 mm, curved. INTENDED USE For haemostasis, clamping of large blood vessels or manipulating heavy tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function Pressure force instrument: haemostatic forceps, springy. Flexible arms. Variable settings of the ratchet, lockable. Adjustment of the jaws. Length: 14 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.20% carbon, 12-13% chromium. Reusable. S9910000 - Surg.inst., abdominal /SET S9910001 - Surg.inst., basic surgery /SET Product attributes Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
511 | Medical & Surgical Instruments | New | Scissors, Tonsil, Metzenbaum, Curved, 18 cm | Scissors, Tonsil, Metzenbaum, Curved, 18 cm | GENERAL DESCRIPTION Scissors, Metzenbaum, 180 mm, curved, blunt / blunt. INTENDED USE Surgical scissors used for cutting or dissecting delicate tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function See Martin Catalogue no. 24 Ref: 11 255 14. Instrument that cuts by shearing: surgical scissors. Thin, curved, with blunt end blades. Length: 18 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.40 % carbon; 13-14 % chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Scissors, Metzenbaum, 180 mm, curved, blunt / blunt. INTENDED USE Surgical scissors used for cutting or dissecting delicate tissue. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Static function See Martin Catalogue no. 24 Ref: 11 255 14. Instrument that cuts by shearing: surgical scissors. Thin, curved, with blunt end blades. Length: 18 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0.15-0.40 % carbon; 13-14 % chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
512 | Medical & Surgical Instruments | New | Ramathobodi tubal hook | Ramathobodi tubal hook | GENERAL DESCRIPTION Ramathibodi Tubal Hook comes with blunt tips to prevent piercing injuries to the tissues. INTENDED USE The streamlined built of the instrument makes it especially useful for manipulating the Fallopian tubes during sterilization surgeries.Surgical . PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Manipulates the Fallopian tubes and pulls them through the narrow space for surgical manipulation. Especially includes a long shank for deep access to the tissue. The tip has a ergonomic hook that slips across the tube and pulls it for access. In addition, the instrument has a straight and slender profile for better access to the tubes.Moreover, the smooth edges of the instrument allow comfortable insertion with minimal damage to surrounding tissue. The hook has an overall length of 25 cm which allows it to reach the organ comfortably. The light weight and streamlined design of the instrument decreases user fatigue. Moreover, the broad handle of the Ramathibodi Tubal Hook allows precise control. The streamlined built of the instrument makes it especially useful for manipulating the Fallopian tubes during sterilization surgeries. Made in premium grade stainless steel with a glare free satin finish. The instrument can be sterilized and reused. Handle Type Solid Handle Overall Length 25 cm – 10" Finish Satin Material Stainless Steel CE Marking Yes Reusable Yes Packaging and labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Ramathibodi Tubal Hook comes with blunt tips to prevent piercing injuries to the tissues. INTENDED USE The streamlined built of the instrument makes it especially useful for manipulating the Fallopian tubes during sterilization surgeries.Surgical . PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Manipulates the Fallopian tubes and pulls them through the narrow space for surgical manipulation. Especially includes a long shank for deep access to the tissue. The tip has a ergonomic hook that slips across the tube and pulls it for access. In addition, the instrument has a straight and slender profile for better access to the tubes.Moreover, the smooth edges of the instrument allow comfortable insertion with minimal damage to surrounding tissue. The hook has an overall length of 25 cm which allows it to reach the organ comfortably. The light weight and streamlined design of the instrument decreases user fatigue. Moreover, the broad handle of the Ramathibodi Tubal Hook allows precise control. The streamlined built of the instrument makes it especially useful for manipulating the Fallopian tubes during sterilization surgeries. Made in premium grade stainless steel with a glare free satin finish. The instrument can be sterilized and reused. Handle Type Solid Handle Overall Length 25 cm – 10" Finish Satin Material Stainless Steel CE Marking Yes Reusable Yes Packaging and labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
513 | Medical & Surgical Instruments | New | Ramathobodi uterine elevator, 28 cm | Ramathobodi uterine elevator, 28 cm | GENERAL DESCRIPTION Ramathibodi Uterine Manipulator comes with blunt tips to prevent piercing injuries to the tissues. INTENDED USE The streamlined built of the instrument makes it especially useful for manipulating the uterine cavity during surgeries. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS manipulates uterine cavity, directing tissue as needed. useful during laparoscopic pelvic surgeries, especially laparoscopic sterilization. Especially includes long shanks for deep access to the tissue. In addition, the instrument has a curved profile for better access to the uterus. Moreover, the smooth edges of the instrument allow comfortable insertion with minimal damage to surrounding tissue. The manipulator has an overall length of 27.5 cm which allows it to reach the organ comfortably. The light weight and streamlined design of the instrument decreases user fatigue. Moreover, the finger ring handles of the Ramathibodi Uterine Manipulator allows precise control. The streamlined built of the instrument makes it especially useful for manipulating the uterine cavity during surgeries. Made in premium grade stainless steel with a glare free satin finish. The instrument can be sterilized and reused. Instrument Profile Curved Handle Type Ring Handle Overall Length 27.5 cm – 11" Finish Satin Material Stainless Steel CE Marking Yes Reusable Yes Packaging and labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Ramathibodi Uterine Manipulator comes with blunt tips to prevent piercing injuries to the tissues. INTENDED USE The streamlined built of the instrument makes it especially useful for manipulating the uterine cavity during surgeries. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS manipulates uterine cavity, directing tissue as needed. useful during laparoscopic pelvic surgeries, especially laparoscopic sterilization. Especially includes long shanks for deep access to the tissue. In addition, the instrument has a curved profile for better access to the uterus. Moreover, the smooth edges of the instrument allow comfortable insertion with minimal damage to surrounding tissue. The manipulator has an overall length of 27.5 cm which allows it to reach the organ comfortably. The light weight and streamlined design of the instrument decreases user fatigue. Moreover, the finger ring handles of the Ramathibodi Uterine Manipulator allows precise control. The streamlined built of the instrument makes it especially useful for manipulating the uterine cavity during surgeries. Made in premium grade stainless steel with a glare free satin finish. The instrument can be sterilized and reused. Instrument Profile Curved Handle Type Ring Handle Overall Length 27.5 cm – 11" Finish Satin Material Stainless Steel CE Marking Yes Reusable Yes Packaging and labelling: One (1) Hook in a plastic bag. Symbols used according ISO 15223 Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
514 | Medical & Surgical Instruments | New | Forceps Schroeder-Braun uterine tenaculum, 24.8 cm | Forceps, uterine, tenaculum, Duplay, 248 mm, curved | GENERAL DESCRIPTION Forceps, uterine, tenaculum, Duplay, 248 mm, curved. INTENDED USE Schroeder Uterine Tenaculum Forceps is a specialized OB/GYN instrument that surgeons commonly use for grasping, holding and mobilizing soft tissues and blood vessels during gynaecological procedures. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: gripping forceps, springy. See Martin Catalogue no. 24 Ref: 32 645 28. Uterine traction forceps, vulsellum / tenaculum, double curved. Hard ratchet, lockable. Length: 28 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0-15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forceps, uterine, tenaculum, Duplay, 248 mm, curved. INTENDED USE Forceps, uterine, tenaculum, Duplay, 248 mm, curved is a specialized OB/GYN instrument that surgeons commonly use for grasping, holding and mobilizing soft tissues and blood vessels during gynaecological procedures. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Pressure force instrument: gripping forceps, springy. See Martin Catalogue no. 24 Ref: 32 645 28. Uterine traction forceps, vulsellum / tenaculum, double curved. Hard ratchet, lockable. Length: 28 cm. Material: Martensitic steel (quenched, magnetic steel). Martensitic steel composition: 0-15-0.20% carbon, 12-13% chromium. Reusable. Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
515 | Medical & Surgical Instruments | New | Forceps, sponge, Foerster, straight, 24 cm | Forceps, sponge, Foerster, straight, 24 cm | GENERAL DESCRIPTION Forester sponge forceps are another type of forceps that are smooth, so during any medical procedures, the tip of the forceps don’t cause any irritation to the wound of the patient. INTENDED USE Sponge Forceps are the forceps that are used to hold sponges and swabs during medical procedures so, there should be no direct physical contact with the wound. These are standard gynecological and thoracic instruments to minimize damage and trauma during surgery. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS These forceps are also available in sterile stainless-steel material. Length: 24 cm. Floor Grade quality Packaged in case quantities Instruments kept in individual sleeves Single-Use Sterile Durable design Stainless steel Constructed of stainless steel OR grade quality Ideal for hospitals and specialty clinics Autoclavable Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH. | GENERAL DESCRIPTION Forester sponge forceps are another type of forceps that are smooth, so during any medical procedures, the tip of the forceps don’t cause any irritation to the wound of the patient. INTENDED USE Sponge Forceps are the forceps that are used to hold sponges and swabs during medical procedures so, there should be no direct physical contact with the wound. These are standard gynecological and thoracic instruments to minimize damage and trauma during surgery. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS These forceps are also available in sterile stainless-steel material. Length: 24 cm. Floor Grade quality Packaged in case quantities Instruments kept in individual sleeves Sterilized prior to use Durable design Stainless steel Constructed of stainless steel OR grade quality Ideal for hospitals and specialty clinics Autoclavable Packaging and labelling: Unit presentation: individually, in protective packaging. The following should appear on the packaging: - Designation of the instrument. - Name and address and supplier (manufacturer) Symbols used according ISO 15223 CE Mark Instructions for use: Used for dressing / swabbing of vagina in preparation for surgical intervention. Also used as serving forceps (used with jar for forceps) This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 1°C to 30°C. Humidity: 30% RH to 70% RH.
516 | Medical & Surgical Instruments | New | 1 Jackson vaginal retractor, 38mm X 76mm (deep blade) (for alternative method for inserting the uterine elevator) | 1 Jackson vaginal retractor, 38mm X 76mm (deep blade) (for alternative method for inserting the uterine elevator) | GENERAL DESCRIPTION Jackson Vaginal Retractor is also known as a flat-blade retractor with a curved handle to separate the vaginal walls, allowing better visualization and easy access to the cervix. INTENDED USE To expose the vaginal cavity. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Length :17.8cm, Blade 3 .8cm X 7.6cm MATERIAL: Stainless Steel INSTRUMENT TYPE: Retractors APPLICATION: Vaginal Retractor BLADES: Flat Blades HANDLE: Curve Handle SURFACE: Satin REUSABLE: Yes AUTOCLAVED: Yes RUST PREVENTION PROCEDURE: Passivated ULTRASONIC CLEANED: Yes TESTS PERFORMED: Boil Test, Performance Test, Shape Test Packaging and labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Any symbol used on the labelling should be in accordance of the ISO 15223- 1: 2016: If the item has a CE marking (or equivalent), the CE logo (or equivalent) must be visible on the packaging. Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 5°C to 40°C. Humidity: 40% RH to 70% RH. | GENERAL DESCRIPTION Jackson Vaginal Retractor is also known as a flat-blade retractor with a curved handle to separate the vaginal walls, allowing better visualization and easy access to the cervix. INTENDED USE To expose the vaginal cavity. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Length :17.8cm, Blade 3 .8cm X 7.6cm MATERIAL: Stainless Steel INSTRUMENT TYPE: Retractors APPLICATION: Vaginal Retractor BLADES: Flat Blades HANDLE: Curve Handle SURFACE: Satin REUSABLE: Yes AUTOCLAVED: Yes RUST PREVENTION PROCEDURE: Passivated ULTRASONIC CLEANED: Yes TESTS PERFORMED: Boil Test, Performance Test, Shape Test Packaging and labelling: Primary packaging: Unit of use One (1) retractor in a plastic bag. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol). Information for handling, if applicable (or equivalent harmonised symbol). Over packaging: Packaging unit Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: Number of units per box. Any symbol used on the labelling should be in accordance of the ISO 15223- 1: 2016: If the item has a CE marking (or equivalent), the CE logo (or equivalent) must be visible on the packaging. Instructions for use: To retract skin, fatty tissue, muscles or viscera during superficial surgery. This item must be cleaned, disinfected after each use, and sterilised in a steam steriliser. Used to grip soft tissue Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: ISO 7153-1:2016: Surgical instruments - Materials - Part 1: Metals. ISO 7151:1988: Surgical instruments - Non-cutting, articulated instruments - General requirements and test methods. ISO 13402:1995: Surgical and dental hand instruments - Determination of resistance against autoclaving, corrosion and thermal exposure. ISO 17664:2017: Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices. ISO 14971:2007 Medical devices - Application of risk management to medical devices. Environmental conditions: Storage and transport conditions: 5°C to 40°C. Humidity: 40% RH to 70% RH.
517 | Medical Electrical Equipment | New | ECG Recorder,portable,w/access | ECG Recorder,portable,w/access | GENERAL DESCRIPTION ECG recorder, portable, w/access Electrocardiograph (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, portable, AC and battery powered , with printer and accessories. Multichannel, interpretive electrocardiograph, including printer. INTENDED USE Electrocardiographs detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time. They are used to diagnose and assist in treating some types of heart disease and arrhythmias. Electrocardiographs differ from ECG monitors, which provide a continuous real-time display of the ECG waveform and other physiologic measurements on a screen and warn of potentially life-threatening or precursor rhythms. Electrocardiographs are used to diagnose cardiac abnormalities, determine a patient's response to drug therapy, and reveal trends or changes in heart function. Multichannel electrocardiographs record and display signals from two or more leads simultaneously." "PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS ECG analysis and full interpretation (rhythm and events), real-time and manual interpretation. Three (3) or more recording channels. SIGNAL ACQUISITION Simultaneous 12 lead acquisition ( aVR, aVL, aVF, I, II, III and V1 - V6) derived from 10 electrodes (RA, LA, RL, LL, V1 - V6). Arrythmia and ST elevation detection. Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV. Lead switching, either manual, automatic or both (manual and auto options). Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50 mm/s. Minimum HR range 30 – 300 bpm with rhythm analysis. Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz. Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better. Calibration signal of 1 mV, manual and/or automatic. Selectable/adjustable filters for baseline drift, muscle artefacts, mains power. Pacemaker detection. Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater. Input impedance > 50 MΩ. Internal noise level < 12.5 µV peak-to-peak. Automatic baseline centring. Defibrillation fluctuation/overload protection. Baseline recovery < 5 s after defibrillation. AC fluctuation protection. RECORDER AND PRINTER Minimum of 3 recording channels. Recorder display includes date/time, patient data and heart rate and basic settings. Capable of displaying one group of at least three channels simultaneously. Recorder waveform display includes lead marker and timing marker. Integrated/built-in printer. Capable of printing user selected number of channels. Capable of printing one group of at least three channels simultaneously. Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s. Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility). DATA INPUT/OUTPUT, STORAGE AND ALARMS Integrated alpha-numeric keyboard. Patient data input fields include name, age, height and weight, gender. Backlit Liquid Crystal Display (LCD) display screen, minimum of 12.7 cm (5 inches) diagonally. Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission. Expandable storage for additional waveforms if required, via USB. Capable of exporting waveform data and reports, via USB or LAN. Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors. Automatic self-test at start up. POWER PROVISIONS Built-in rechargeable lithium-ion battery. Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording. Automatic switch to battery in case of power failure, automatic recharge on connection to mains. Maximum battery charging time to full charge is 8 hours. Power requirements: 100 - 240 Volts - 50/60 Hz. Designed for frequent and easy dismount and disinfection with hospital-grade products. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x Patient cable. 2 x full sets of chest electrodes. 2 x full sets of limb electrodes. 1 x supply of 960 Z-folded sheets suitable for the unit. (if accepts Z-folded sheets), or 5 x supply of rolls suitable for the unit (if the unit accepts rolls). 3 x bottle of 100 ml of ECG conductive gel. 1 x spare rechargeable battery pack. REGULATORY & CONFORMITY REQUIREMENTS : CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION : Classified under EU MDR 2017/745 as Class IIa device " "SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-25:2011 Medical Electrical Equipment – Part 2-25: Particular requirements for the safety of electrocardiographs." "INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. COMPONENT OF A KIT Not part of a kit. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C – 40°C / 25% - 85% RH. - Storage conditions: -20°C - 40°C / 25% - 95% RH. - Atmospheric pressure: 860 ~ 1060 hPa. ESTIMATED LIFE SPAN Five years. WARRANTY Two years from shipping date. | GENERAL DESCRIPTION ECG recorder, portable, w/access Electrocardiograph (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, portable, AC and battery powered , with printer and accessories. Multichannel, interpretive electrocardiograph, including printer. INTENDED USE Electrocardiographs detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time. They are used to diagnose and assist in treating some types of heart disease and arrhythmias. Electrocardiographs differ from ECG monitors, which provide a continuous real-time display of the ECG waveform and other physiologic measurements on a screen and warn of potentially life-threatening or precursor rhythms. Electrocardiographs are used to diagnose cardiac abnormalities, determine a patient's response to drug therapy, and reveal trends or changes in heart function. Multichannel electrocardiographs record and display signals from two or more leads simultaneously. PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS ECG analysis and full interpretation (rhythm and events), real-time and manual interpretation. For recording and printing, it is required to have Minimum of 3 recording channels simultaneously. SIGNAL ACQUISITION Simultaneous 12 lead acquisition ( aVR, aVL, aVF, I, II, III and V1 - V6) derived from 10 electrodes (RA, LA, RL, LL, V1 - V6). Arrythmia and ST elevation detection. Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV. Lead switching, either manual, automatic or both (manual and auto options). Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50 mm/s. Minimum HR range 30 – 300 bpm with rhythm analysis. Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz. Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better. Calibration signal of 1 mV, manual and/or automatic. Selectable/adjustable filters for baseline drift, muscle artefacts, mains power. Pacemaker detection. Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater. Input impedance > 50 MΩ. Internal noise level < 12.5 µV peak-to-peak. Automatic baseline centring. Defibrillation fluctuation/overload protection. Baseline recovery < 5 s after defibrillation. AC fluctuation protection. RECORDER AND PRINTER Minimum of 3 recording channels. Recorder display includes date/time, patient data and heart rate and basic settings. Capable of displaying one group of at least three channels simultaneously. Recorder waveform display includes lead marker and timing marker. Integrated/built-in printer. Capable of printing user selected number of channels. Capable of printing one group of at least three channels simultaneously. Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s. Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility). DATA INPUT/OUTPUT, STORAGE AND ALARMS Integrated alpha-numeric keyboard. Patient data input fields include name, age, height and weight, gender. Backlit Liquid Crystal Display (LCD) display screen, minimum of 12.7 cm (5 inches) diagonally. Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission. Expandable storage for additional waveforms if required, via USB. Capable of exporting waveform data and reports, via USB or LAN. Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors. Automatic self-test at start up. POWER PROVISIONS Built-in rechargeable lithium-ion battery. Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording. Automatic switch to battery in case of power failure, automatic recharge on connection to mains. Maximum battery charging time to full charge is 8 hours. Power requirements: 100 - 240 Volts - 50/60 Hz. Designed for frequent and easy dismount and disinfection with hospital-grade products. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x Patient cable. 2 x full sets of chest electrodes. 2 x full sets of limb electrodes. 1 x supply of 960 Z-folded sheets suitable for the unit. (if accepts Z-folded sheets), or 5 x supply of rolls suitable for the unit (if the unit accepts rolls). 3 x bottle of 100 ml of ECG conductive gel. 1 x spare rechargeable battery pack. REGULATORY & CONFORMITY REQUIREMENTS : CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION : Classified under EU MDR 2017/745 as Class IIa device " "SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-25:2011 Medical Electrical Equipment – Part 2-25: Particular requirements for the safety of electrocardiographs." "INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. COMPONENT OF A KIT Not part of a kit. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C – 40°C / 25% - 85% RH. - Storage conditions: -20°C - 40°C / 25% - 95% RH. - Atmospheric pressure: 860 ~ 1060 hPa. ESTIMATED LIFE SPAN Five years. WARRANTY Two years from shipping date.
518 | Med. Diagnostics Equipment & Supplies | New | Sphygmomanaometer, Aneroid, Adult | Sphygmomanaometer, Aneroid, Adult | GENERAL DESCRIPTION Sphygmomanometer, aneroid, adult INTENDED USE A device designed to measure blood pressure consisting of an inflatable cuff that fits around a limb (arm or thigh), an inflation bulb for controlling the air pressure within the cuff, an aneroid manometer, and tubing. The aneroid manometer consists of a metal bellows, which expands as the pressure in the cuff increases, and a mechanical amplifier that transmits this expansion through a lever to an indicator needle, which rotates around a circular, calibrated scale. The manometer is hand held (portable); blood pressure measurement is taken in conjunction with a stethoscope." PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Device used for the indirect (non-invasive) measurement of arterial blood pressure The sphygmomanometer is composed of cloth cuff containing an inflatable bag. Connected via a tube to a flexible bulb with valve and integrated manometer needle gauge. Material for cuff: durable nylon, non-deformable, washable at 30ºC. Material for tube: rubber or other suitable material, e.g. silicone rubber, crack resistant. Very strong cuff with double velcro fastening, enabling it to be adjusted to fit tightly around the arm. Cuff reinforced at both ends. Dimensions of cuff: approx. 570 x 145mm. Gauge graduated 0 - 300mmHg (min) in 2 (max) mmHg increments, with pressure release valve. Accuracy as per ISO 81060-1: +/- 3mm Hg. Latex and mercury free design. The bag is inflated by means of a flexible bulb connected via a tube Length of tube between: 50 to 70 cm and reliable quick connector). The quick connector can easily be connected to all types of inflation bulbs. Rubber inflation bulb with integrated manometer needle gauge. Glass and metal aneroid pressure gauge with needle. Dial graduation: 0 to 300mmHG. With pressure release valve. Easy to grip bulb." "SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective case or pouch. INSTRUCTION FOR USE For the measurement of arterial blood pressure. PACKAGING & LABELLING : One (1) sphygmomanometer (adult) in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark with Notified Number Body" REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class Im device SAFETY AND PRODUCT STANDARDS EN 1060-3:1997+A2:2009 Non-invasive sphygmomanometers - Part 3: Supplementary requirements for electro-mechanical blood pressure measuring systems EN 1060-4:2004 Non-invasive sphygmomanometers - Part 4: Test procedures to determine the overall system accuracy of automated non-invasive sphygmomanometers ISO 81060-1:2007 Non-invasive sphygmomanometers - Part 1: Requirements and test methods for non-automated measurement type ISO 10993-1:2009" ESTIMATED LIFE SPAN 5 - 10 years WARRANTY Two years Environmental requirements: To avoid contaminant paint. | GENERAL DESCRIPTION Sphygmomanometer, aneroid, adult INTENDED USE A device designed to measure blood pressure consisting of an inflatable cuff that fits around a limb (arm or thigh), an inflation bulb for controlling the air pressure within the cuff, an aneroid manometer, and tubing. The aneroid manometer consists of a metal bellows, which expands as the pressure in the cuff increases, and a mechanical amplifier that transmits this expansion through a lever to an indicator needle, which rotates around a circular, calibrated scale. The manometer is hand held (portable); blood pressure measurement is taken in conjunction with a stethoscope." PRODUCT DESCRIPTION / TECHNICAL SPECIFICATIONS Device used for the indirect (non-invasive) measurement of arterial blood pressure The sphygmomanometer is composed of cloth cuff containing an inflatable bag. Connected via a tube to a flexible bulb with valve and integrated manometer needle gauge. Material for cuff: durable nylon, non-deformable, washable at 30ºC. Material for tube: rubber or other suitable material, e.g. silicone rubber, crack resistant. Very strong cuff with double velcro fastening, enabling it to be adjusted to fit tightly around the arm. Cuff reinforced at both ends. Dimensions of cuff: approx. 570 x 145mm. Gauge graduated 0 - 300mmHg (min) in 2 (max) mmHg increments, with pressure release valve. Accuracy as per ISO 81060-1: +/- 3mm Hg. Latex and mercury free design. The bag is inflated by means of a flexible bulb connected via a tube Length of tube between: 50 to 70 cm and reliable quick connector). The quick connector can easily be connected to all types of inflation bulbs. Rubber inflation bulb with integrated manometer needle gauge. Glass and metal aneroid pressure gauge with needle. Dial graduation: 0 to 300mmHG. With pressure release valve. Easy to grip bulb." "SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective case or pouch. INSTRUCTION FOR USE For the measurement of arterial blood pressure. PACKAGING & LABELLING : One (1) sphygmomanometer (adult) in a box or case or bag with manufacturer's instruction for use in English, French and Spanish, spare parts and accessories (when applicable). Symbols used according ISO 15223 CE mark with Notified Number Body" REGULATORY & CONFORMITY REQUIREMENTS CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class Im device SAFETY AND PRODUCT STANDARDS EN 1060-3:1997+A2:2009 Non-invasive sphygmomanometers - Part 3: Supplementary requirements for electro-mechanical blood pressure measuring systems EN 1060-4:2004 Non-invasive sphygmomanometers - Part 4: Test procedures to determine the overall system accuracy of automated non-invasive sphygmomanometers ISO 81060-1:2007 Non-invasive sphygmomanometers - Part 1: Requirements and test methods for non-automated measurement type ISO 10993-1:2009" ESTIMATED LIFE SPAN 5 - 10 years WARRANTY Two years Environmental requirements: To avoid contaminant paint.
519 | Anaesthesia & Resus. Equip | New | Anaesthesia machine Portable | Anaesthesia machine Portable | GENERAL DESCRIPTION: A portable, stand-alone, mechanical device allowing anesthesia without pressurized gaz for spontaneous or assisted ventilation. The DPA tank is a thermostabilized and calibrated vaporiser. Its low internal resistance allows ambient air to be used as the carrier gas for volatile agents (draw-over anaesthesia). The DPA tank can also be used with medical gas continuous flow systems, or on anaesthetic respirators. The DPA02 is used to provide draw-over anesthesia technique. It allows the use of halogene agents (halothane, isoflurane). It can be used with low pressure oxygen source such as oxygen concentrators. The new DPA02 valve and patient circuit removes the valve from the vicinity of the patient airway and allows the use of standard Y circuit. This valve is more suitable to the spontaneous ventilation than the previous ones. SPECIFICATION : The configuration chosen my MSF includes 2 adult circuits: EANEANAA411 2 paediatric circuits (children >10 kg): EANEANAA412 2 Jackson Rees modified - T circuit Mapleson F - 0.5 l balloon open tail (T piece system): EANEANAA305 TECHNICAL SPECIFICATIONS : DPA tank Height: 330 mm - estimated Width (inlet to outlet): 139 mm - estimated Diameter of the tank: 50 mm - estimated Weight: 2.6 kg - estimated Capacity: 150 ml - estimated Arrow engraved on the tank, indicating the gas flow direction Connections: 22 mm Calibration: 0.5% - 5% halothane Internal resistance: less than 1 cm H2O at 40 litres/min Constant concentration for a flow rate between 4 and 8 litres/min Compatible for isoflurane Width: 47 cm, height: 33 cm, depth: 17 cm - estimate Weight: 9.8 kg - estimate PACKAGING : Contained in a shockproof and waterproof Peli case. SUPPLIED WITH THE ARTICLE: 1 DPA02 tank 1 oxygen reservoir balloon 2 l (neoprene) 1 conservation bloc plus 7.5 cm H2O valve 1 connection tube for oxygen supply 1 flowmeter to regulate oxygen supply 1 DPA02 respiratory valve 2 adult 2-branch silicone circuits 2 paediatric 2-branch silicone circuits 2 Y-piece adult 2 Y-piece paediatric 2 Jackson Rees modified - Mapleson F circuit 0.5 l balloon open tail (T piece system) 1 adult self-inflating silicone balloon 1 paediatric self inflating silicone balloon 1 extension tube for self inflating balloon 1 extension tube for scavenger 1 PEEP valve 1 connector for PEEP valve and scavenger tube 1 transport suit-case 1 scavenging system for anesthesia gas To be Ordered Separately isoflurane bottles (regulated transport) INSTRUCTION FOR USE For draw-over inhaling anaesthesia without N2O. The vaporiser is calibrated either for isoflurane, either for halothane (for the new circuits sent by MSF only Isoflurane is used). The gaseous mixture must be enriched with oxygen (bottled or from an oxygen concentrator)." PRECAUTION FOR USE Halothane and isoflurane must never be mixed. REGULATORY AND COMPLIANCE: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device MAINTENANCE The use of an antibacterial-antiviral filter is mandatory between each patient. The patient circuit should be cleaned once a week with water and soap. The system should be wiped with a detergent-disinfectant solution for surfaces after each procedure. In case of shortage of filter, the DPA2 valve and the patient circuits should be autoclaved between each patient. At that time, T piece circuit should be considered as single use. Basic maintenance can be performed by the anaesthetist. MSF REQUIREMENTS : The DPA is the vaporiser of choice for draw-over anaesthesia in MSF. | GENERAL DESCRIPTION: A portable, stand-alone, mechanical device allowing anesthesia without pressurized gaz for spontaneous or assisted ventilation. The DPA tank is a thermostabilized and calibrated vaporiser. Its low internal resistance allows ambient air to be used as the carrier gas for volatile agents (draw-over anaesthesia). The DPA tank can also be used with medical gas continuous flow systems, or on anaesthetic respirators. The DPA02 is used to provide draw-over anesthesia technique. It allows the use of halogene agents (halothane, isoflurane). It can be used with low pressure oxygen source such as oxygen concentrators. The new DPA02 valve and patient circuit removes the valve from the vicinity of the patient airway and allows the use of standard Y circuit. This valve is more suitable to the spontaneous ventilation than the previous ones. SPECIFICATION : The configuration chosen my MSF includes 2 adult circuits: EANEANAA411 2 paediatric circuits (children >10 kg): EANEANAA412 2 Jackson Rees modified - T circuit Mapleson F - 0.5 l balloon open tail (T piece system): EANEANAA305 TECHNICAL SPECIFICATIONS : DPA tank Height: 330 mm - estimated Width (inlet to outlet): 139 mm - estimated Diameter of the tank: 50 mm - estimated Weight: 2.6 kg - estimated Capacity: 150 ml - estimated Arrow engraved on the tank, indicating the gas flow direction Connections: 22 mm Calibration: 0.5% - 5% halothane Internal resistance: less than 1 cm H2O at 40 litres/min Constant concentration for a flow rate between 4 and 8 litres/min Compatible for isoflurane Width: 47 cm, height: 33 cm, depth: 17 cm - estimate Weight: 9.8 kg - estimate PACKAGING : Contained in a shockproof and waterproof Peli case. SUPPLIED WITH THE ARTICLE: 1 DPA02 tank 1 oxygen reservoir balloon 2 l (neoprene) 1 conservation bloc plus 7.5 cm H2O valve 1 connection tube for oxygen supply 1 flowmeter to regulate oxygen supply 1 DPA02 respiratory valve 2 adult 2-branch silicone circuits 2 paediatric 2-branch silicone circuits 2 Y-piece adult 2 Y-piece paediatric 2 Jackson Rees modified - Mapleson F circuit 0.5 l balloon open tail (T piece system) 1 adult self-inflating silicone balloon 1 paediatric self inflating silicone balloon 1 extension tube for self inflating balloon 1 extension tube for scavenger 1 PEEP valve 1 connector for PEEP valve and scavenger tube 1 transport suit-case 1 scavenging system for anesthesia gas To be Ordered Separately isoflurane bottles (regulated transport) INSTRUCTION FOR USE For draw-over inhaling anaesthesia without N2O. The vaporiser is calibrated either for isoflurane, either for halothane (for the new circuits sent by MSF only Isoflurane is used). The gaseous mixture must be enriched with oxygen (bottled or from an oxygen concentrator)." PRECAUTION FOR USE Halothane and isoflurane must never be mixed. REGULATORY AND COMPLIANCE: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device MAINTENANCE The use of an antibacterial-antiviral filter is mandatory between each patient. The patient circuit should be cleaned once a week with water and soap. The system should be wiped with a detergent-disinfectant solution for surfaces after each procedure. In case of shortage of filter, the DPA2 valve and the patient circuits should be autoclaved between each patient. At that time, T piece circuit should be considered as single use. Basic maintenance can be performed by the anaesthetist. MSF REQUIREMENTS : The DPA is the vaporiser of choice for draw-over anaesthesia in MSF.
520 | Medical Electrical Equipment | New | Electrocardiograph with 12 channels | Electrocardiograph with 12 channels | GENERAL DESCRIPTION Multichannel, interpretive electrocardiograph, including printer. Product Description Electrocardiograph (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, portable, AC and battery powered, with printer and accessories. TECHNICAL SPECIFICATIONS ECG analysis and full interpretation (rhythm and events), real-time and manual interpretation. Three (3) or more recording channels. SIGNAL ACQUISITION Simultaneous 12 lead acquisition ( aVR, aVL, aVF, I, II, III and V1 - V6) derived from 10 electrodes (RA, LA, RL, LL, V1 - V6). Arrythmia and ST elevation detection. Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV. Lead switching, either manual, automatic or both (manual and auto options). Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50 mm/s. Minimum HR range 30 – 300 bpm with rhythm analysis. Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz. Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better. Calibration signal of 1 mV, manual and/or automatic. Selectable/adjustable filters for baseline drift, muscle artefacts, mains power. Pacemaker detection. Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater. Input impedance > 50 MΩ. Internal noise level < 12.5 µV peak-to-peak. Automatic baseline centring. Defibrillation fluctuation/overload protection. Baseline recovery < 5 s after defibrillation. AC fluctuation protection." INTENDED USE Electrocardiographs detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time. They are used to diagnose and assist in treating some types of heart disease and arrhythmias. Electrocardiographs differ from ECG monitors, which provide a continuous real-time display of the ECG waveform and other physiologic measurements on a screen and warn of potentially life-threatening or precursor rhythms. Electrocardiographs are used to diagnose cardiac abnormalities, determine a patient's response to drug therapy, and reveal trends or changes in heart function. Multichannel electrocardiographs record and display signals from two or more leads simultaneously." RECORDER AND PRINTER Minimum of 3 recording channels. Recorder display includes date/time, patient data and heart rate and basic settings. Capable of displaying one group of at least three channels simultaneously. Recorder waveform display includes lead marker and timing marker. Integrated/built-in printer. Capable of printing user selected number of channels. Capable of printing one group of at least three channels simultaneously. Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s. Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility)." DATA INPUT/OUTPUT, STORAGE AND ALARMS Integrated alpha-numeric keyboard. Patient data input fields include name, age, height and weight, gender. Backlit Liquid Crystal Display (LCD) display screen, minimum of 12.7 cm (5 inches) diagonally. Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission. Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors. Automatic self-test at start up. POWER PROVISIONS Built-in rechargeable lithium-ion battery. Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording. Automatic switch to battery in case of power failure, automatic recharge on connection to mains.Maximum battery charging time to full charge is 8 hours. Power requirements: 100 - 240 Volts - 50/60 Hz.Designed for frequent and easy dismount and disinfection with hospital-grade products." SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x Patient cable. 2 x full sets of chest electrodes. 2 x full sets of limb electrodes. 1 x supply of 960 Z-folded sheets suitable for the unit. (if accepts Z-folded sheets), or 5 x supply of rolls suitable for the unit (if the unit accepts rolls). 3 x bottle of 100 ml of ECG conductive gel. 1 x spare rechargeable battery pack." ENVIRONMENTAL CONDITIONS- Operating conditions: 5°C – 40°C / 25% - 85% RH. - Storage conditions: -20°C - 40°C / 25% - 95% RH. - Atmospheric pressure: 860 ~ 1060 hPa." INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-25:2011 Medical Electrical Equipment – Part 2-25: Particular requirements for the safety of electrocardiographs." | GENERAL DESCRIPTION Multichannel, interpretive electrocardiograph, including printer. Product Description Electrocardiograph (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, portable, AC and battery powered, with printer and accessories. TECHNICAL SPECIFICATIONS ECG analysis and full interpretation (rhythm and events), real-time and manual interpretation. For recording and printing it is required to have at least: Three (3) or more recording channels simultaneously. SIGNAL ACQUISITION Simultaneous 12 lead acquisition ( aVR, aVL, aVF, I, II, III and V1 - V6) derived from 10 electrodes (RA, LA, RL, LL, V1 - V6). Arrythmia and ST elevation detection. Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV. Lead switching, either manual, automatic or both (manual and auto options). Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50 mm/s. Minimum HR range 30 – 300 bpm with rhythm analysis. Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz. Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better. Calibration signal of 1 mV, manual and/or automatic. Selectable/adjustable filters for baseline drift, muscle artefacts, mains power. Pacemaker detection. Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater. Input impedance > 50 MΩ. Internal noise level < 12.5 µV peak-to-peak. Automatic baseline centring. Defibrillation fluctuation/overload protection. Baseline recovery < 5 s after defibrillation. AC fluctuation protection." INTENDED USE Electrocardiographs detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time. They are used to diagnose and assist in treating some types of heart disease and arrhythmias. Electrocardiographs differ from ECG monitors, which provide a continuous real-time display of the ECG waveform and other physiologic measurements on a screen and warn of potentially life-threatening or precursor rhythms. Electrocardiographs are used to diagnose cardiac abnormalities, determine a patient's response to drug therapy, and reveal trends or changes in heart function. Multichannel electrocardiographs record and display signals from two or more leads simultaneously." RECORDER AND PRINTER Minimum of 3 recording channels. Recorder display includes date/time, patient data and heart rate and basic settings. Capable of displaying one group of at least three channels simultaneously. Recorder waveform display includes lead marker and timing marker. Integrated/built-in printer. Capable of printing user selected number of channels. Capable of printing one group of at least three channels simultaneously. Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s. Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility)." DATA INPUT/OUTPUT, STORAGE AND ALARMS Integrated alpha-numeric keyboard. Patient data input fields include name, age, height and weight, gender. Backlit Liquid Crystal Display (LCD) display screen, minimum of 12.7 cm (5 inches) diagonally. Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission. Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors. Automatic self-test at start up. POWER PROVISIONS Built-in rechargeable lithium-ion battery. Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording. Automatic switch to battery in case of power failure, automatic recharge on connection to mains.Maximum battery charging time to full charge is 8 hours. Power requirements: 100 - 240 Volts - 50/60 Hz.Designed for frequent and easy dismount and disinfection with hospital-grade products." SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x Patient cable. 2 x full sets of chest electrodes. 2 x full sets of limb electrodes. 1 x supply of 960 Z-folded sheets suitable for the unit. (if accepts Z-folded sheets), or 5 x supply of rolls suitable for the unit (if the unit accepts rolls). 3 x bottle of 100 ml of ECG conductive gel. 1 x spare rechargeable battery pack." ENVIRONMENTAL CONDITIONS- Operating conditions: 5°C – 40°C / 25% - 85% RH. - Storage conditions: -20°C - 40°C / 25% - 95% RH. - Atmospheric pressure: 860 ~ 1060 hPa." INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS User training prior to utilization is recommended. MAINTENANCE/USER REQUIREMENTS As per user and service manuals. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-25:2011 Medical Electrical Equipment – Part 2-25: Particular requirements for the safety of electrocardiographs.
521 | Medical Electrical Equipment | New | CTG (Cardiotocography) instrument | CTG (Cardiotocography) instrument | GENERAL DESCRIPTION Cardiotocograph (CTG) system for antepartum foetal and maternal monitoring. Product Description Capable of external monitoring of foetal heart rate (FHR) and maternal uterine activity (UA). Alphanumeric display shows FHR1, FHR2 and UCs and alarms. Automatically detects transducers when they are plugged in. Includes remote switch for event marking. Automatic self-test on power up. System reports with status and alarms. Unit includes a battery allowing for > 2 hours continuous operation. Power requirements: 100 - 240 Volts - 50/60 Hz. FETAL HEART RATE MONITORING Foetal heart rate detected by ultrasonic transducer/probe. Capable of measuring FHR in the 50-240 beats per minute (bpm) range with 1 bpm resolution and 2 bpm accuracy. Can monitor twins (i.e., has two transducers). Includes high/low audible and visual alarms. Includes signal quality/loss indicators and alarm(s). Provides audible feedback on signal quality. MATERNAL UTERINE ACTIVITY MONITORING Uterine contractions measured though a pressure sensitive transducer. Capable of measuring relative uterine contractions (UCs) in the range of 0-100 units with at least 1 unit resolution and 1 unit accuracy. Includes signal loss due to unplugged transducer. Intended Purpose Cardiotocography (CTG) is a continuous recording of the fetal heart rate obtained via an ultrasound transducer placed on the mother’s abdomen. CTG is widely used in pregnancy as a method of assessing fetal well‐being, predominantly in pregnancies with increased risk of complications." PRINTER/RECORDER Integrated thermal printer with automatic and manual print-out modes. Prints FHR1, FHR2, UCs, marked events and parameters and relevant alarms. Print speeds include 1, 2 and 3 cm/min. Compatible with z-folded thermal paper. Two different model are supplied: One model allows for adjusting paper scale between 20 and 30 bmp/cm, the other model has a fixed paper scale of 30 bmp/cm Includes alarms or another indicator for end of paper." SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x UC transducer. 2 x FHR ultrasound transducers. 1 x Remote switch event marker with cable. 3 x Adjustable transducer belts for ultrasound (2 FHRs) and toco (UC). 2 x Box of thermal recording paper, total 100 z-folded sheets. 2 x Bottle of ultrasound gel, approximately 250ml. Instruction for Use The QVCAD System is indicated for use as an aid to the reader during screening procedures in searching images of female breasts produced by the somo.v Automated Breast Ultrasound System (screening mammography BI-RADS® Assessment Category 1 or 2, and BI-RADS Composition/Density c or d) to detect mammography-occult lesions in regions not known to have suspicious findings. The indicators produced by the QVCAD System are not intended to be used for diagnostic characterization of suspicious findings. Accessories /Spares/ Consumables At extra costs additional optional features can be requested: - An optional cart can be included. - An optional back up battery can be included." Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety and Product standards - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-37:2007+AMD1:2015 CSV Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. Environmental requirements Context-dependent requirements Handling environment temperature: –20 to 60 °C" | GENERAL DESCRIPTION Cardiotocograph (CTG) system for antepartum foetal and maternal monitoring. Product Description Capable of external monitoring of foetal heart rate (FHR) and maternal uterine activity (UA). Alphanumeric display shows FHR1, FHR2 and UCs and alarms. Automatically detects transducers when they are plugged in. Includes remote switch for event marking. Automatic self-test on power up. System reports with status and alarms. Unit includes a battery allowing for > 2 hours continuous operation. Power requirements: 100 - 240 Volts - 50/60 Hz. FETAL HEART RATE MONITORING Foetal heart rate detected by ultrasonic transducer/probe. Capable of measuring FHR in the 50-240 beats per minute (bpm) range with 1 bpm resolution and 2 bpm accuracy. Can monitor twins (i.e., has two transducers). Includes high/low audible and visual alarms. Includes signal quality/loss indicators and alarm(s). Provides audible feedback on signal quality. MATERNAL UTERINE ACTIVITY MONITORING Uterine contractions measured though a pressure sensitive transducer. Capable of measuring relative uterine contractions (UCs) in the range of 0-100 units with at least 1 unit resolution and 1 unit accuracy. Includes signal loss due to unplugged transducer. Intended Purpose Cardiotocography (CTG) is a continuous recording of the fetal heart rate obtained via an ultrasound transducer placed on the mother’s abdomen. CTG is widely used in pregnancy as a method of assessing fetal well‐being, predominantly in pregnancies with increased risk of complications." PRINTER/RECORDER Integrated thermal printer with automatic and manual print-out modes. Prints FHR1, FHR2, UCs, marked events and parameters and relevant alarms. Print speeds include 1, 2 and 3 cm/min. Compatible with z-folded thermal paper. Two different model are supplied: One model allows for adjusting paper scale between 20 and 30 bmp/cm, the other model has a fixed paper scale of 30 bmp/cm Includes alarms or another indicator for end of paper." SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x UC transducer. 2 x FHR ultrasound transducers. 1 x Remote switch event marker with cable. 3 x Adjustable transducer belts for ultrasound (2 FHRs) and toco (UC). 2 x Box of thermal recording paper, total 100 z-folded sheets. 2 x Bottle of ultrasound gel, approximately 250ml. Instruction for Use The QVCAD System is indicated for use as an aid to the reader during screening procedures in searching images of female breasts produced by the somo.v Automated Breast Ultrasound System (screening mammography BI-RADS® Assessment Category 1 or 2, and BI-RADS Composition/Density c or d) to detect mammography-occult lesions in regions not known to have suspicious findings. The indicators produced by the QVCAD System are not intended to be used for diagnostic characterization of suspicious findings. Accessories /Spares/ Consumables At extra costs additional optional features can be requested: - An optional cart can be included. - An optional back up battery can be included." Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device Safety and Product standards - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-37:2007+AMD1:2015 CSV Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. Environmental requirements Context-dependent requirements Handling environment temperature: –20 to 60 °C"
522 | Medical Electrical Equipment | New | Automated external defibrillator system | Automated external defibrillator system | GENERAL DESCRIPTION Compact fully automatic external (AED) defibrillator including Cardiopulmonary Resuscitation (CPR) support feature. Product Description Adult and paediatric settings. Biphasic output waveform Automatic ECG (VT/VF) detection and analysis. Automatic impedance compensation based on patient. Paediatric dose attenuation. Automatic switch between AED and CPR modes based on analysis. Analysis time less than 10 seconds after having been switched on. Maximum energy output for adults between 150 and 200 Joules. Maximum energy output for paediatrics 50 Joules, or adjustable between 30 and 70 (depending on the model supplied). Charge time to maximum energy output is 5 seconds. Includes step-by-step device and CPR user guide, either in durable plastic-coated manual and/or on machine. Audible metronome for CPR procedure. Built-in discharge feature for safety. Integrated control panel with all parameters and controls. Step-by-step pictograms on the control panel for ready and easy operation. Audio and/or visual indications of operational status and step-by-step operation. Audio and/or visual alarms for operational status, electrodes, battery status and system errors. Automatic self-test and continuous check of pads and electrodes connection. Conductive surface area of adult electrodes is at least 80 cm² Conductive surface area of paediatric electrodes is at least 80 cm² Shelf life of electrodes is at least 2 years. Replaceable internal battery, non-rechargeable. Battery type LiSO2 or LiM no (depending on the model supplied). The internal battery, when full supports at least 140 full discharges at 200J, or two hours continuous ECG monitoring. Weighs less than 3 kg estimated Intended Use Fully automated external defibrillators (AEDs) deliver a high-voltage electrical impulse to the heart in order to restore normal rhythm and contractile function in patients who are experiencing ventricular fibrillation (VF) or ventricular tachycardia (VT) that is not accompanied by a palpable pulse. AEDs differ from conventional manual defibrillators in that AEDs can analyse the ECG rhythm to determine whether defibrillation is necessary; this eliminates the need for the user to interpret the cardiac rhythm before delivering a shock. AEDs are designed for use primarily by first responders to cardiac emergencies, who may not be fully trained in advanced cardiac life support (ACLS). In the prehospital setting, these emergency personnel can include EMTs, firefighters, and law enforcement officers. Flight attendants, security guards, and others −sometimes called targeted responders may be expected to use PAD units. AEDs can also be used in areas of the hospital where advanced life support personnel are not readily available. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x carry case, with storage pockets for leads and accessories. 3 x set of adult adhesive external pads, colour-coded and with pictograms. 3 x set of paediatric adhesive external pads, colour-coded and with pictograms. 1 x plastic enclosed Quick Reference Guide (step-by-step AED and CPR). 1 x CD-ROM/DVD/USB with training material. 2 x sets of spare batteries, packed separately. Instruction for Use Step 1: Turn the defibrillator on by pressing the green button and follow its instructions. Step 2: Peel off the sticky pads and attach them to the patient’s skin, one on each side of the chest, as shown in the picture on the defibrillator. Step 3: Once the pads have been attached, stop CPR and don’t touch the patient. The defibrillator will then check the patient’s heart rhythm. Step 4: The defibrillator will decide whether a shock is needed and if so, it will tell you to press the shock button. An automatic defibrillator will shock the patient without prompt. Don’t touch the patient while they are being shocked. Step 5: The defibrillator will tell you when the shock has been delivered and whether you need to continue CPR. Step 6: Continue with chest compressions until the patient shows signs of life or the defibrillator tells you to stop so it can analyse the heartbeat again. Anyone can use a defibrillator and you don’t need training. Once you turn it on, it will give clear instructions on how to attach the defibrillator pads. The device checks the heart rhythm and will only tell you to shock if it’s needed. Y ou can't shock someone accidentally. Accessories/Spares/Consumables Defibrillator with AED and External Pacemaker – 01 Adult with Built in Paediatric External Paddles - 01 Patient cables - 01 ECG Rolls – 50 Adult SpO2 reusable Sensor – 01 Adult NIBP Cuff and Hose – 01 88 etCO2 Tubing (box of 20) – 01 box AED Multifunction Pads for Adults - 10 pairs with Each unit Packaging and Labelling Disposable, self-adhesive, single-use electrodes; ready-to-use, sealed and packed with connector outside pouch General Device Labeling - 21 CFR Part 801 Use of Symbols - 21 CFR Part 801.15 In Vitro Diagnostic Products - 21 CFR Part 809 Investigational Device Exemptions - 21 CFR Part 812 Unique Device Identification - 21CFR Part 830 Good Manufacturing Practices - 21 CFR Part 820 General Electronic Products - 21 CFR Part 1010 The Federal Food, Drug and Cosmetic Act (FFDCA) is the law under which the FDA takes action against regulated products. Specifically:Section 201(k) defines 'label' as a: 'display of written, printed, or graphic matter upon the immediate container of any article...' The term 'immediate container' does not include package liners. Any word, statement, or other information appearing on the immediate container must also appear 'on the outside container or wrapper, if any there be, or the retail package of such article, or is easily legible through the outside container of wrapper.' Section 201(m) defines 'labeling' as: 'all labels and other written, printed, or graphic matter (1) upon any article or any of its containers or wrappers, or (2) accompanying such article' at any time while a device is held for sale after shipment or delivery for shipment in interstate commerce. The term 'accompanying' is interpreted liberally to mean more than physical association with the product. It extends to posters, tags, pamphlets, circulars, booklets, brochures, instruction books, direction sheets, fillers, etc. ' Accompanying' also includes labeling that is brought together with the device after shipment or delivery for shipment in interstate commerce. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device" Safety and Product Standards - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-1-6:2010 Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - IEC 60601-2-4:2010+AMD1:2018 CSV Consolidated version Medical electrical equipment - Part 2-4: Particular requirements for the basic safety and essential performance of cardiac defibrillators. Environmental Complainces - Operating conditions: 0°C - 50°C / 10 % – 95% RH. - Storage conditions: -10°C - 50°C / 10% – 95% RH. - Atmospheric pressure up to 700 hPa (altitude of 3,000m). - Ingress protection rating: IP44. | GENERAL DESCRIPTION Compact fully automatic external (AED) defibrillator including Cardiopulmonary Resuscitation (CPR) support feature. Product Description Adult and paediatric settings. Biphasic output waveform Automatic ECG (VT/VF) detection and analysis. Automatic impedance compensation based on patient. Paediatric dose attenuation. Automatic switch between AED and CPR modes based on analysis. Analysis time less than 10 seconds after having been switched on. Maximum energy output for adults between 150 and 200 Joules. Maximum energy output for paediatrics 50 Joules, or adjustable between 30 and 70 (depending on the model supplied). Charge time to maximum energy output is 5 seconds. Includes step-by-step device and CPR user guide, either in durable plastic-coated manual and/or on machine. Audible metronome for CPR procedure. Built-in discharge feature for safety. Integrated control panel with all parameters and controls. Step-by-step pictograms on the control panel for ready and easy operation. Audio and/or visual indications of operational status and step-by-step operation. Audio and/or visual alarms for operational status, electrodes, battery status and system errors. Automatic self-test and continuous check of pads and electrodes connection. Conductive surface area of adult electrodes is at least 80 cm² Conductive surface area of paediatric electrodes is at least 80 cm² Shelf life of electrodes is at least 2 years. Replaceable internal battery, non-rechargeable. Battery type LiSO2 or LiM no (depending on the model supplied). The internal battery, when full supports at least 140 full discharges at 200J, or two hours continuous ECG monitoring. Weighs less than 3 kg estimated Intended Use Fully automated external defibrillators (AEDs) deliver a high-voltage electrical impulse to the heart in order to restore normal rhythm and contractile function in patients who are experiencing ventricular fibrillation (VF) or ventricular tachycardia (VT) that is not accompanied by a palpable pulse. AEDs differ from conventional manual defibrillators in that AEDs can analyse the ECG rhythm to determine whether defibrillation is necessary; this eliminates the need for the user to interpret the cardiac rhythm before delivering a shock. AEDs are designed for use primarily by first responders to cardiac emergencies, who may not be fully trained in advanced cardiac life support (ACLS). In the prehospital setting, these emergency personnel can include EMTs, firefighters, and law enforcement officers. Flight attendants, security guards, and others −sometimes called targeted responders may be expected to use PAD units. AEDs can also be used in areas of the hospital where advanced life support personnel are not readily available. SUPPLIED WITH Instructions for assembly, use and maintenance in English, French and Spanish. 1 x carry case, with storage pockets for leads and accessories. 3 x set of adult adhesive external pads, colour-coded and with pictograms. 3 x set of paediatric adhesive external pads, colour-coded and with pictograms. 1 x plastic enclosed Quick Reference Guide (step-by-step AED and CPR). 1 x CD-ROM/DVD/USB with training material. 2 x sets of spare batteries, packed separately. 1x External Paddles - 01 10 pairs x AED Multifunction Pads for Adults Instruction for Use Step 1: Turn the defibrillator on by pressing the green button and follow its instructions. Step 2: Peel off the sticky pads and attach them to the patient’s skin, one on each side of the chest, as shown in the picture on the defibrillator. Step 3: Once the pads have been attached, stop CPR and don’t touch the patient. The defibrillator will then check the patient’s heart rhythm. Step 4: The defibrillator will decide whether a shock is needed and if so, it will tell you to press the shock button. An automatic defibrillator will shock the patient without prompt. Don’t touch the patient while they are being shocked. Step 5: The defibrillator will tell you when the shock has been delivered and whether you need to continue CPR. Step 6: Continue with chest compressions until the patient shows signs of life or the defibrillator tells you to stop so it can analyse the heartbeat again. Anyone can use a defibrillator and you don’t need training. Once you turn it on, it will give clear instructions on how to attach the defibrillator pads. The device checks the heart rhythm and will only tell you to shock if it’s needed. Y ou can't shock someone accidentally. Packaging and Labelling Disposable, self-adhesive, single-use electrodes; ready-to-use, sealed and packed with connector outside pouch General Device Labeling - 21 CFR Part 801 Use of Symbols - 21 CFR Part 801.15 In Vitro Diagnostic Products - 21 CFR Part 809 Investigational Device Exemptions - 21 CFR Part 812 Unique Device Identification - 21CFR Part 830 Good Manufacturing Practices - 21 CFR Part 820 General Electronic Products - 21 CFR Part 1010 The Federal Food, Drug and Cosmetic Act (FFDCA) is the law under which the FDA takes action against regulated products. Specifically:Section 201(k) defines 'label' as a: 'display of written, printed, or graphic matter upon the immediate container of any article...' The term 'immediate container' does not include package liners. Any word, statement, or other information appearing on the immediate container must also appear 'on the outside container or wrapper, if any there be, or the retail package of such article, or is easily legible through the outside container of wrapper.' Section 201(m) defines 'labeling' as: 'all labels and other written, printed, or graphic matter (1) upon any article or any of its containers or wrappers, or (2) accompanying such article' at any time while a device is held for sale after shipment or delivery for shipment in interstate commerce. The term 'accompanying' is interpreted liberally to mean more than physical association with the product. It extends to posters, tags, pamphlets, circulars, booklets, brochures, instruction books, direction sheets, fillers, etc. ' Accompanying' also includes labeling that is brought together with the device after shipment or delivery for shipment in interstate commerce. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class III device" Safety and Product Standards - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-1-6:2010 Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability. - IEC 60601-2-4:2010+AMD1:2018 CSV Consolidated version Medical electrical equipment - Part 2-4: Particular requirements for the basic safety and essential performance of cardiac defibrillators. Environmental Complainces - Operating conditions: 0°C - 50°C / 10 % – 95% RH. - Storage conditions: -10°C - 50°C / 10% – 95% RH. - Atmospheric pressure up to 700 hPa (altitude of 3,000m). - Ingress protection rating: IP44.
523 | Med. Diagnostics & Supplies | New | Handheld breast ultrasound | Handheld breast ultrasound | General Description: Ultrasound Breast ultrasound is an imaging technique that uses sound waves to create images of breast tissue." "Product Description Linear Ultrasound Scanner Enjoy superior imaging to 11cm in our most versatile linear scanner. ✓ High-Definition wireless ultrasound ✓ Easy-to-use app for iOS and Android ✓ Steadfast wireless connection Applications: Breast, Lung, MSK, Ocular, Small Parts, Vascular [Included] ✓ B-Mode ✓ Color Doppler ✓ M-Mode ✓ 1 hour personal training [Options] - DICOM - Needle Enhance - Pulsed-Wave Doppler - Elastography" Intended purpose Breast ultrasound is an imaging test that uses sound waves to look at the inside of your breasts. It can help your healthcare provider find breast problems. It also lets your healthcare provider see how well blood is flowing to areas in your breasts. Supplied with Battery life for hand-held ultrasound system. several hours . They define the type of image that will be displayed. The most common ultrasound modes and their applications are as follows: A-mode: this is the simplest type of ultrasound. No image will be displayed on the screen, only a one-dimensional representation of reflected sound waves. This mode is used for mesuring distances. B-mode (or 2B mode): linear array transducers simultaneously scan a plane through the body that can be seen as a 2D image on the screen. This is the most commonly used mode in various displines. M-mode:this consists in a fast B-mode scanning sequence with successive images on the screen allowing doctors to see and measure the amplitude of movements ( the perimeters of the organs producing the reflections move in relation to the transducer) Doppler mode: this mode allows you to visualize and analyze blood flow, vascularization, or even the direction of flow if the Doppler is in color. This mode is used in cardiology to analyze the condition of heart valves or the vascularization of certain tissue. Elastography mode: with this mode it is possible to measure the hardness of tissue and to diagnose, for example, a potential hepatic fibrosis. 3D and real-time 3D (4D) mode: this mode is used during ultrasound examinations of pregnant women to obtain a three-dimensional image of the fetus and observe its movements in utero. But the risks to the fetus are not negligible (see question 2). Wifi type : 802.11n/5G/450Mbps. Working system : iOS/Android /Windows. Charging mode : USB charging or wireless charging. Android phone brand : SUMSUNG,HUAWEI, HTC, SONY, Google Nexus,MI. Instruction for use The ultrasound system should be used far away from the strong electromagnetic field (e.g. transformer). Otherwise, the system will be disturbed The ultrasound system should be used far away from high-frequency radioactive device (e.g. mobile phone). Otherwise, the system will get damaged or be affected. : " Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device" Safety and Product standards - FDA - CE - ISO13485 Federal Food, Drug and Cosmetic Act (FFDCA), Chapter V, Subchapter C - Electronic Product Radiation Control. Manufacturers of ultrasound imaging products are responsible for compliance with all applicable requirements of Title 21 Code of Federal Regulations (Subchapter J, Radiological Health) Parts 1000 through 1005: 1000 - General 1002 - Records and Reports 1003 - Notification of defects or failure to comply 1004 - Repurchase, repairs, or replacement of electronic products 1005 - Importation of electronic products There are no federal radiation safety performance standards for diagnostic ultrasound. Industry Guidance - Documents of Interest Guidance for Industry and FDA Staff - Information for Manufacturers Seeking Marketing Clearance of Diagnostic Ultrasound Systems and Transducers (PDF Version) (PDF - 474KB) Declaration for Imported Electronic Products Subject to Radiation Control Standards (PDF - 399KB) Ultrasound Letter to Industry (PDF 41KB) Variance Application Process" Environmental Compliances Working environment  The ultrasound system should be operated, preserved and transported under the following conditions: Conditions Parameters Temperature Operation 5 ℃～40 ℃ Preservation Transportation5 ℃ Humidity 30%～80%， ～40 ℃ Less than 80%30 ℃～55 ℃，no condensation no condensation Less than 95%，no condensation Atmospheric pressure 70 kPa～106kPa 70 kPa ～106kPa 50kPa～106kPa C" | General Description: Ultrasound Breast ultrasound is an imaging technique that uses sound waves to create images of breast tissue." "Product Description Linear Ultrasound Scanner Enjoy superior imaging to 11cm in our most versatile linear scanner. ✓ High-Definition wireless ultrasound ✓ Easy-to-use app for iOS and Android ✓ Steadfast wireless connection Applications: Breast, Lung, MSK, Ocular, Small Parts, Vascular [Included] ✓ B-Mode ✓ Color Doppler ✓ M-Mode ✓ 1 hour personal training [Options] - DICOM - Needle Enhance - Pulsed-Wave Doppler - Elastography" Intended purpose Breast ultrasound is an imaging test that uses sound waves to look at the inside of your breasts. It can help your healthcare provider find breast problems. It also lets your healthcare provider see how well blood is flowing to areas in your breasts. Supplied with Battery life for hand-held ultrasound system. several hours . They define the type of image that will be displayed. The most common ultrasound modes and their applications are as follows: A-mode: this is the simplest type of ultrasound. No image will be displayed on the screen, only a one-dimensional representation of reflected sound waves. This mode is used for mesuring distances. B-mode (or 2B mode): linear array transducers simultaneously scan a plane through the body that can be seen as a 2D image on the screen. This is the most commonly used mode in various displines. M-mode:this consists in a fast B-mode scanning sequence with successive images on the screen allowing doctors to see and measure the amplitude of movements ( the perimeters of the organs producing the reflections move in relation to the transducer) Doppler mode: this mode allows you to visualize and analyze blood flow, vascularization, or even the direction of flow if the Doppler is in color. This mode is used in cardiology to analyze the condition of heart valves or the vascularization of certain tissue. Elastography mode: with this mode it is possible to measure the hardness of tissue and to diagnose, for example, a potential hepatic fibrosis. 3D and real-time 3D (4D) mode: this mode is used during ultrasound examinations of pregnant women to obtain a three-dimensional image of the fetus and observe its movements in utero. But the risks to the fetus are not negligible (see question 2). Wifi type : 802.11n/5G/450Mbps. Working system : iOS/Android /Windows. Charging mode : USB charging or wireless charging. Android phone brand : SUMSUNG,HUAWEI, HTC, SONY, Google Nexus,MI. Instruction for use The ultrasound system should be used far away from the strong electromagnetic field (e.g. transformer). Otherwise, the system will be disturbed The ultrasound system should be used far away from high-frequency radioactive device (e.g. mobile phone). Otherwise, the system will get damaged or be affected. : " Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device" Safety and Product standards - FDA - CE - ISO13485 Federal Food, Drug and Cosmetic Act (FFDCA), Chapter V, Subchapter C - Electronic Product Radiation Control. Manufacturers of ultrasound imaging products are responsible for compliance with all applicable requirements of Title 21 Code of Federal Regulations (Subchapter J, Radiological Health) Parts 1000 through 1005: 1000 - General 1002 - Records and Reports 1003 - Notification of defects or failure to comply 1004 - Repurchase, repairs, or replacement of electronic products 1005 - Importation of electronic products There are no federal radiation safety performance standards for diagnostic ultrasound. Industry Guidance - Documents of Interest Guidance for Industry and FDA Staff - Information for Manufacturers Seeking Marketing Clearance of Diagnostic Ultrasound Systems and Transducers (PDF Version) (PDF - 474KB) Declaration for Imported Electronic Products Subject to Radiation Control Standards (PDF - 399KB) Ultrasound Letter to Industry (PDF 41KB) Variance Application Process" Environmental Compliances Working environment  The ultrasound system should be operated, preserved and transported under the following conditions: Conditions Parameters Temperature Operation 5 ℃～40 ℃ Preservation Transportation5 ℃ Humidity 30%～80%， ～40 ℃ Less than 80%30 ℃～55 ℃，no condensation no condensation Less than 95%，no condensation Atmospheric pressure 70 kPa～106kPa 70 kPa ～106kPa 50kPa～106kPa C"
524 | Medical Electrical Equipment | New | Ultrasound, Scanner, Digital w/access | Ultrasound, Scanner, Digital w/access | General Description Ultrasound, scanner, digital, portable, with accessories Product Description: General purpose ultrasound scanner, with obstetrics package and probes. System includes scanner and printer mounted on a trolley. The system supports B mode and M mode. The unit includes: one (1) convex array probe, 2.5-5 MHz and one (1) endovaginal probe, 5.5-7.5 MHz. System includes scanner and printer mounted on a trolley. An obstetrics analysis package is included which provides for foetal weight and anatomy, gestational age and foetal growth calculations. The unit comes standard with packages for: abdomen, gynaecology/obstetrics, cardiology and small parts. Depending on the model supplied Pulsed Wave Doppler (back/white), is included in the unit. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit)." IMAGE DISPLAY Digital callipers. Cine loop. Cine loop. A minimum of 4 x zoom 256 levels of greyscale. A freeze function. Capable of lateral and vertical inversion in B mode. Capable of displaying B+M modes simultaneously. Full screen annotation, for real-time and frozen images." Cart /Trolley Includes at least two side probe holders. Side handle(s) for transport and positioning. Surface space for scanner and keyboard. Storage space (shelf) for printer. Four anti-static braking swivel castors, at least 10 cm high. Data storage and Communication Internal hard drive for limited data storage. Supported communication protocols: USB, S-video, VGA, and depending on the model: RS 232 or ethernet). Supported formats for image export: JPEG. Intended use Portable ultrasonic scanning systems are used in physician offices, radiology departments, cardiology departments, emergency departments, critical care units and by mobile imaging service providers. They can also have important applications in disaster triage and mass casualty incidents. The main benefit of point of care ultrasounds units is that it allows physicians to quickly determine whether an abnormality is present so they can make appropriate patient management decisions." Supplied with Instructions for assembly, use and maintenance in English, French and Spanish: 1 x Plastic protective dustcover. 3 x bottle of 100 ml (or 2 X 150 ml) of ultrasound gel. 1 x Set of spare fuses. OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - A linear array probe 5.5- 8.5 MHz can be added to the unit. - A micro-convex probe, 5.0 -7.5 MHz can be included in the unit. - A rechargeable battery (standard in one model, optional in other model). - DICOM 3.0 compatibility (Included standard in one model, optional in other model). - A thermal printer, supporting a 256 greyscale, 325 dpi resolution at a minimum, including printer head cleaning kit. INSTALLATION REQUIREMENTS This product does not have complex assembly and/or commissioning needs requiring a technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-37:2007+AMD1:2015 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C – 40°C / 25% - 80% RH. - Storage conditions: -20°C - 55°C / 25% - 93% RH. - Atmospheric pressure: 860 ~ 1060 hPa." "This product does not have complex assembly and/or commissioning needs requiring a technician. Thermofisher Reference :https://www.thermofisher.com 3730 x DNA analyser" | General Description Ultrasound, scanner, digital, portable, with accessories Product Description: General purpose ultrasound scanner, with obstetrics package and probes. System includes scanner and printer mounted on a trolley. The system supports B mode and M mode. The unit includes: one (1) convex array probe, 2.5-5 MHz and one (1) endovaginal probe, 5.5-7.5 MHz. System includes scanner and printer mounted on a trolley. An obstetrics analysis package is included which provides for foetal weight and anatomy, gestational age and foetal growth calculations. The unit comes standard with packages for: abdomen, gynaecology/obstetrics, cardiology and small parts. Depending on the model supplied Pulsed Wave Doppler (back/white), is included in the unit. Power requirements: 100 - 240 Volts - 50/60 Hz (not necessarily in a single unit)." IMAGE DISPLAY Digital callipers. Cine loop. Cine loop. A minimum of 4 x zoom 256 levels of greyscale. A freeze function. Capable of lateral and vertical inversion in B mode. Capable of displaying B+M modes simultaneously. Full screen annotation, for real-time and frozen images." Cart /Trolley Includes at least two side probe holders. Side handle(s) for transport and positioning. Surface space for scanner and keyboard. Storage space (shelf) for printer. Four anti-static braking swivel castors, at least 10 cm high. Data storage and Communication Internal hard drive for limited data storage. Supported communication protocols: USB, S-video, VGA, and depending on the model: RS 232 or ethernet). Supported formats for image export: JPEG. Intended use Portable ultrasonic scanning systems are used in physician offices, radiology departments, cardiology departments, emergency departments, critical care units and by mobile imaging service providers. They can also have important applications in disaster triage and mass casualty incidents. The main benefit of point of care ultrasounds units is that it allows physicians to quickly determine whether an abnormality is present so they can make appropriate patient management decisions." Supplied with Instructions for assembly, use and maintenance in English, French and Spanish: 1 x Plastic protective dustcover. 3 x bottle of 100 ml (or 2 X 150 ml) of ultrasound gel. 1 x Set of spare fuses. OPTIONAL ACCESSORIES AVAILABLE At extra costs additional optional features can be requested: - A linear array probe 5.5- 8.5 MHz can be added to the unit. - A micro-convex probe, 5.0 -7.5 MHz can be included in the unit. - A rechargeable battery (standard in one model, optional in other model). - DICOM 3.0 compatibility (Included standard in one model, optional in other model). - A thermal printer, supporting a 256 greyscale, 325 dpi resolution at a minimum, including printer head cleaning kit. INSTALLATION REQUIREMENTS This product does not have complex assembly and/or commissioning needs requiring a technician. TRAINING REQUIREMENTS User training prior to utilization is recommended. Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS - IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. - IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. - IEC 60601-2-37:2007+AMD1:2015 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. ENVIRONMENTAL CONDITIONS - Operating conditions: 5°C – 40°C / 25% - 80% RH. - Storage conditions: -20°C - 55°C / 25% - 93% RH. - Atmospheric pressure: 860 ~ 1060 hPa." "This product does not have complex assembly and/or commissioning needs requiring a technician. Thermofisher Reference :https://www.thermofisher.com 3730 x DNA analyser"
525 | Med. Diagnostics & Supplies | New | Digital blood pressure monitor | Digital blood pressure monitor | General Description Automatic electronic sphygmomanometers non-invasive. Digital automatic non-invasive BP monitor Product Description An electrically powered device designed to non-invasivel meaysure BP, with a self-contained software program to regulate automatic arm-cuff inflation and measurement cycles. It typically displays current heart rate and mean arterial pressure in addition to systolic and diastolic BPbedside. Measurement interval, min: User selectable: ≥ 5 Measurement time (s) ≤ 60, Automatic 0 required. Display :LCD digital display Measurement range Pressure: 0 to 299 mmHg Pulse: 40 to 180 beats / min Inflation :Fuzzy-logic controlled by electric pump Deflation:Automatic pressure release valve Power source : 4 “AA” batteries 1.5V Battery life :Approximately 1000 measurements (using new alkaline batteries Operating temperature / humidity / air pressure 50°F to 104°F (10°C to 40°C) / 15 to 90% RH / 700 to 1060 hPa Storage temperature / humidity / air pressure-4°F to 140°F (-20°C to 60°C) / 10 to 95% RH / 700 to 1060 hPa Single-use cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm) adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes of the cuffs depend on the manufacturer but should not deviate by ± 5 cm from the stated ""Rubber tube (length > 30 cm), reusable cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The cuff is made of durable material (e.g. nylon), which is non-deformable, and washable at 30°C. The cuff is fitted with double Velcro fastening, enabling a tight and secure fit around arms.The cuff is reinforced at both sides. The bag is inflated by means of the flexible bulb connected via a tube. Material tube: rubber or other suitable material, e.g. silicone rubber, crack resistant. Length tube between: 50 to 70 cm." Intended purpose A device designed to measure blood pressure consisting of an cuff inflate itself with the push of a button .It measures blood pressure and pulse simply and quickly and is easy to use, reliable and portable." Supplied with 1BP Machine Device, 4 AA Alkaline Battery, 1 Water Proof Carry pouch, 1 Adjustable Arm Cuff (22-36 CM), Power Source Both Battery-Powered & USB Powered (USB wire not Included in Package)1" "Instruction for use Do not subject the monitor to strong shocks, such as dropping the unit on the floor. Do not submerge the device or any of the components in water. Store the device and the components in a clean, safe location. Changes or modification not approved by the manufacturer will void the user warranty. Do not disassemble or attempt to repair the unit or components. Accessories (if relevant)/Spares /Consumables Sterilization process for accessories (if relevant) Not applicable Consumables and reagents (if relevant) Single-use cuffs in the following sizes : Neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). Reusable cuffs in the following sizes: Neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes off the cuffs may vary by manufacturer but should not deviate by ± 5 cm from the stated sizes. 28Spare parts (if relevant)Rubber tube (length > 30 cm), reusable cuffs of various sizes 29Other components (if relevant) Protective container" Packaging and Labelling Sterility status on delivery (if relevant) Single-use cuffs must be delivered sterile. Shelf life (if relevant) Minimum shelf life for single-use cuffs must be 1 year from the date of reception. Transportation and storage (if relevant) Storage environment humidity: 10–95% relative humidity. Storage environment temperature: –20 to 60 °C""""""" Labelling (if relevant) With the proper certification and validation requested, plus those required in each country" Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im" Safety and standards Depends on the country. Examples: Class A (Global Harmonization Task Force Rule 4); Class II (USA); Class I (Australia, Canada and Japan); Class IIa (European Union)"" ""Proof of regulatory compliance (e.g. registration, clearance, approval) must be provided as appropriate per the product’s risk classification, by regulatory agency (e.g. by a founding member of IMDRF-EU, USA, Canada, Australia, Japan). Else approved by local national regulatory agency."" ""Standards applicable to the product and to the manufacturing process are listed below. Compliance to the last available version of the international standard or to its local equivalent standard is recommended and proof of compliance must be provided. Non-exhaustive list of standards applicable to general quality systems for medical devices and specific for BPMD: • ISO 13485:2016, Medical devices – Quality management systems – Requirements for regulatory purposes • EN ISO 14971:2012, Medical devices – Application of risk management to medical devices • ISO 14155:2011, Clinical investigation of medical devices for human subjects – Good clinical practice • ISO 14971:2007, Medical devices – Application of risk management to medical devices • IEC 80601-2-30:2018 Medical electrical equipment – Part 2-30: Particular requirements for basic safety and essential performance of automated non-invasive sphygmomanometers • ISO 16142-1:2016, Medical devices – Recognized essential principles of safety and performance of medical devices – Part 1: General essential principles and additional specific essential principles for all non-IVD medical devices and guidance on the selection of standards Non-exhaustive list of standards applicable to electronic BP devices: • AAMI/ESH/ISO 81060 Universal Standard for the Validation of Blood Pressure Measuring Devices Noninvasive phygmomanometers – Part 2: Clinical investigation of automated measurement type • ISO 81060-2:2018(E) Non-invasive sphygmomanometer standard Part 2: Clinical investigation of intermittent automated measurement type • ISO/IEEE 11073-10407:2010 (Part 10407: Device specialization – Blood pressure monitor) • IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) • DS/EN 1060-3 Non-invasive sphygmomanometers – Part 3: Electro-mechanical blood pressure measuring system"" ""ANSI/AAMI SP10:2002 & ANSI/AAMI SP10:2002/A1:2003 (Manual, electronic or automated sphygmomanometers) DS/EN 1060-3 Non-invasive sphygmomanometers - Part 3: Electro-mechanical blood pressure measuring system GOST R 50267.30 Medical electrical equipment. Part 2. Particular requirements for safety of automatic cycling indirect blood pressure monitoring equipment JIS T 1115:2005 Non-invasive automated sphygmomanometers"" ""Compliance with (where applicable, but not limited to, and last available version): USA: CFR - Code of Federal Regulations, Title 21, Part 820 CFR - Code of Federal Regulations, Title 21, Part 870, Section 1130 / Non-invasive blood pressure measurement system Japan: MHLW Ordinance No. 16916156000 Aneroid sphygmomanometer European Commission: Council Directive 93/42/EEC of 14 June 1993 on Medical Devices Regulation (EU) 2017/745 of the European Parliament and the Council on Medical Devices Environmental requirements Context-dependent requirements Handling environment temperature: –20 to 60 °C | General Description Automatic electronic sphygmomanometers non-invasive. Digital automatic non-invasive BP monitor Product Description An electrically powered device designed to non-invasivel meaysure BP, with a self-contained software program to regulate automatic arm-cuff inflation and measurement cycles. It typically displays current heart rate and mean arterial pressure in addition to systolic and diastolic BPbedside. Measurement interval, min: User selectable: ≥ 5 Measurement time (s) ≤ 60, Automatic 0 required. Display :LCD digital display Measurement range Pressure: 0 to 299 mmHg Pulse: 40 to 180 beats / min Inflation :Fuzzy-logic controlled by electric pump Deflation:Automatic pressure release valve Power source : 4 “AA” batteries 1.5V Battery life :Approximately 1000 measurements (using new alkaline batteries Operating temperature / humidity / air pressure 50°F to 104°F (10°C to 40°C) / 15 to 90% RH / 700 to 1060 hPa Storage temperature / humidity / air pressure-4°F to 140°F (-20°C to 60°C) / 10 to 95% RH / 700 to 1060 hPa Single-use cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm) adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes of the cuffs depend on the manufacturer but should not deviate by ± 5 cm from the stated ""Rubber tube (length > 30 cm), reusable cuffs in the following sizes: neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The cuff is made of durable material (e.g. nylon), which is non-deformable, and washable at 30°C. The cuff is fitted with double Velcro fastening, enabling a tight and secure fit around arms.The cuff is reinforced at both sides. The bag is inflated by means of the flexible bulb connected via a tube. Material tube: rubber or other suitable material, e.g. silicone rubber, crack resistant. Length tube between: 50 to 70 cm." Intended purpose A device designed to measure blood pressure consisting of an cuff inflate itself with the push of a button .It measures blood pressure and pulse simply and quickly and is easy to use, reliable and portable." Supplied with 1BP Machine Device, 4 AA Alkaline Battery, 1 Water Proof Carry pouch, 1 Adjustable Arm Cuff (22-36 CM), Power Source Both Battery-Powered & USB Powered (USB wire not Included in Package)1" "Instruction for use Do not subject the monitor to strong shocks, such as dropping the unit on the floor. Do not submerge the device or any of the components in water. Store the device and the components in a clean, safe location. Changes or modification not approved by the manufacturer will void the user warranty. Do not disassemble or attempt to repair the unit or components. Accessories (if relevant)/Spares /Consumables Sterilization process for accessories (if relevant) Not applicable Consumables and reagents (if relevant) Single-use cuffs in the following sizes : Neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). Reusable cuffs in the following sizes: Neonatal (10–15 cm), paediatric (14–22 cm), adult (25–36 cm), large adult (34–43 cm), thigh (40–55 cm). The sizes off the cuffs may vary by manufacturer but should not deviate by ± 5 cm from the stated sizes. 28Spare parts (if relevant)Rubber tube (length > 30 cm), reusable cuffs of various sizes 29Other components (if relevant) Protective container" Packaging and Labelling Sterility status on delivery (if relevant) Single-use cuffs must be delivered sterile. Shelf life (if relevant) Minimum shelf life for single-use cuffs must be 1 year from the date of reception. Transportation and storage (if relevant) Storage environment humidity: 10–95% relative humidity. Storage environment temperature: –20 to 60 °C""""""" Labelling (if relevant) With the proper certification and validation requested, plus those required in each country" Regulatory & Conformity Requirements: CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class Im" Safety and standards Depends on the country. Examples: Class A (Global Harmonization Task Force Rule 4); Class II (USA); Class I (Australia, Canada and Japan); Class IIa (European Union)"" ""Proof of regulatory compliance (e.g. registration, clearance, approval) must be provided as appropriate per the product’s risk classification, by regulatory agency (e.g. by a founding member of IMDRF-EU, USA, Canada, Australia, Japan). Else approved by local national regulatory agency."" ""Standards applicable to the product and to the manufacturing process are listed below. Compliance to the last available version of the international standard or to its local equivalent standard is recommended and proof of compliance must be provided. Non-exhaustive list of standards applicable to general quality systems for medical devices and specific for BPMD: • ISO 13485:2016, Medical devices – Quality management systems – Requirements for regulatory purposes • EN ISO 14971:2012, Medical devices – Application of risk management to medical devices • ISO 14155:2011, Clinical investigation of medical devices for human subjects – Good clinical practice • ISO 14971:2007, Medical devices – Application of risk management to medical devices • IEC 80601-2-30:2018 Medical electrical equipment – Part 2-30: Particular requirements for basic safety and essential performance of automated non-invasive sphygmomanometers • ISO 16142-1:2016, Medical devices – Recognized essential principles of safety and performance of medical devices – Part 1: General essential principles and additional specific essential principles for all non-IVD medical devices and guidance on the selection of standards Non-exhaustive list of standards applicable to electronic BP devices: • AAMI/ESH/ISO 81060 Universal Standard for the Validation of Blood Pressure Measuring Devices Noninvasive phygmomanometers – Part 2: Clinical investigation of automated measurement type • ISO 81060-2:2018(E) Non-invasive sphygmomanometer standard Part 2: Clinical investigation of intermittent automated measurement type • ISO/IEEE 11073-10407:2010 (Part 10407: Device specialization – Blood pressure monitor) • IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) • DS/EN 1060-3 Non-invasive sphygmomanometers – Part 3: Electro-mechanical blood pressure measuring system"" ""ANSI/AAMI SP10:2002 & ANSI/AAMI SP10:2002/A1:2003 (Manual, electronic or automated sphygmomanometers) DS/EN 1060-3 Non-invasive sphygmomanometers - Part 3: Electro-mechanical blood pressure measuring system GOST R 50267.30 Medical electrical equipment. Part 2. Particular requirements for safety of automatic cycling indirect blood pressure monitoring equipment JIS T 1115:2005 Non-invasive automated sphygmomanometers"" ""Compliance with (where applicable, but not limited to, and last available version): USA: CFR - Code of Federal Regulations, Title 21, Part 820 CFR - Code of Federal Regulations, Title 21, Part 870, Section 1130 / Non-invasive blood pressure measurement system Japan: MHLW Ordinance No. 16916156000 Aneroid sphygmomanometer European Commission: Council Directive 93/42/EEC of 14 June 1993 on Medical Devices Regulation (EU) 2017/745 of the European Parliament and the Council on Medical Devices Environmental requirements Context-dependent requirements Handling environment temperature: –20 to 60 °C
526 | Medical Sterilization Equipmnt | New | Autoclave | Autoclave | General Description - Sterilizer steam autoclave Product Description Stand-alone table top steam sterilizer with drying cycle Program control unit for fully automated process control, with temperature up to 134 C Liquid temperature-control unit and pressure device (for sterilization of liquids) Internal chamber volume Autoclave recooling system Single door, self-sealing with high-quality silicone gasket Epoxy coated metal housing, interior chamber of stainless steel Fit with 5L water reservoir, auto-fill and autonomy for approximately 10 cycles Water circuit with high-efficiency bacteriological filter Soft-touch control panel allow easy cleaning Panel reports operating temperature, pressure and time, low-water level , as well as system errors (e.g. door) Safety feature protect against over-pressure and over-temperature Audio visual alarm at cycle end and in case of failure or potential danger Power requirements 110- 220V / 50Hz-60Hz, with voltage surge protection Capacity 28 L, 75 -200 L,110-880L available Intended purpose : Autoclaves are also known as steam sterilizers, and are typically used for healthcare or industrial applications. An autoclave is a machine that uses steam under pressure to kill harmful bacteria, viruses, fungi, and spores on items that are placed inside a pressure vessel." Supplied with UPS to ensure power supply for 30 minutes 2 x Bacteriological air filters (spare) 2 x Gaskets (spare) Start-up supplies Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish Instructions for use Bench-top equipment used to sterilize small items of equipment and reagents Used mainly at secondary, tertiary and reference facilities (laboratory network levels II-IV)" Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device" Safety and EMC standards EN61010-1 :2001 EN/61010-2 040:2005 EN61326-1:2006 EMC Directive 89/336 EN60529:1991 Degree of enclosure Protection (IPX4) EN60529:1991 Degree of enclosure Protection (IPX4)"Environmental Compliances Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Operating altitude: Atmospheric pressure: 1 – 70 kPa | General Description - Sterilizer steam autoclave Product Description Stand-alone table top steam sterilizer with drying cycle Program control unit for fully automated process control, with temperature up to 134 C Liquid temperature-control unit and pressure device (for sterilization of liquids) Internal chamber volume Autoclave recooling system Single door, self-sealing with high-quality silicone gasket Epoxy coated metal housing, interior chamber of stainless steel Fit with 5L water reservoir, auto-fill and autonomy for approximately 10 cycles Water circuit with high-efficiency bacteriological filter Soft-touch control panel allow easy cleaning Panel reports operating temperature, pressure and time, low-water level , as well as system errors (e.g. door) Safety feature protect against over-pressure and over-temperature Audio visual alarm at cycle end and in case of failure or potential danger Power requirements 110- 220V / 50Hz-60Hz, with voltage surge protection Capacity 28 L, 75 -200 L,110-880L available Intended purpose : Autoclaves are also known as steam sterilizers, and are typically used for healthcare or industrial applications. An autoclave is a machine that uses steam under pressure to kill harmful bacteria, viruses, fungi, and spores on items that are placed inside a pressure vessel." Supplied with UPS to ensure power supply for 30 minutes 2 x Bacteriological air filters (spare) 2 x Gaskets (spare) Start-up supplies Instructions for use, preventive maintenance and troubleshooting in English, French and Spanish Instructions for use Bench-top equipment used to sterilize small items of equipment and reagents Used mainly at secondary, tertiary and reference facilities (laboratory network levels II-IV)" Regulatory & Conformity Requirements CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Classification: Classified under EU MDR 2017/745 as Class IIb device" Safety and EMC standards EN61010-1 :2001 EN/61010-2 040:2005 EN61326-1:2006 EMC Directive 89/336 EN60529:1991 Degree of enclosure Protection (IPX4) EN60529:1991 Degree of enclosure Protection (IPX4)"Environmental Compliances Storage conditions: 0 - 50°C / 90% RH. Operating conditions: 10 - 40°C / 90% RH. Operating altitude: Atmospheric pressure: 1 – 70 kPa
527 | Medical Electrical Diagnostic Equipment | New | Colposcope | Colposcope | GENERAL DESCRIPTION A device used to examine the lower genital tract epithelium, including tissues of the cervix, vulva, vagina, and anogenital areas. INTENDED USE A colposcope is a low magnification, light-illuminated visualization device for examining the cervix, across an area measuring approximately 20 to 30 mm in diameter, with enough distance between the colposcope lens and the cervix to accommodate the surgical instruments needed for the examination and/or treatment. It allows the examiner to view the epithelial tissues of the cervix and other anogenital areas. PRODUCT DESCRIPTION A colposcope is a low magnification, light-illuminated visualization instrument primarily used alongside screening tools for triaging, diagnosing and managing precancerous cervical lesions in women. It allows the examiner to view the epithelial tissues of the cervix and other anogenital areas. For purposes of cervical precancer assessment, it helps determine the transformation zone type and the grade of suspected epithelial abnormality. In addition, colposcopy facilitates and optimises biopsy and excisional treatment. FUNCTIONAL REQUIREMENTS 1) the colposcope head housing the optics capable of magnification and focus 2) the light source and green light filters for improved visualization of vasculature 3) the body or stand to facilitate and optimises biopsy and excisional treatment TECHNICAL SPECIFICATIONS : DETAILED REQUIREMENTS - Magnification: A range of optical magnification between 3x to 15x (either stepped or continuously variable) - Illumination: Light sources shall be either halogen or LED to guarantee full-spectrum visible light (white light) - Halogen light: 15 V/150 W - LED: 20,000-35,000 LUX (at 300 mm working distance) - An illumination adjustment knob to change the intensity of light - A fan to cool the bulbs (if halogen bulbs are used) - Green light filters; however, blue filters are also acceptable but not the preferred option - Ingress protection rating: IPX2 (minimum) USER ADJUSTABLE SETTINGS - Coarse and fine optical magnification - Illumination UTILITY REQUIREMENTS The unit is suggested to be connected to a reliable power source. • Electrical source requirements (based on country/setting of use): o Plug type: Based on the country settings ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS • Stand or mount to allow for hands-free operation • LED TV or medical grade monitor if not integrated (optional) • Single-use sheath (if using invasive portable model) • Lamp bulbs and fuses TRANSPORTAION AND STORAGE Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% (non-condensing), and be protected from dripping water (IPX2)." PACKAGING AND LABELLING Labelling on the primary packaging should include the name and/ or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements. DOCUMENTATION REQUIREMENTS • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers." REGULATORY & CONFORMITY REQUIREMENTS Approval by regulatory body of country of manufacturer (if applicable) CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan. Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements" "• IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests For vaginally-inserted colposcopes: • Biocompatibility: o ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process o ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity o ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization • Endoscopy (pertaining to inserted scope): o ISO 8600-1: Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements o ISO 8600-3: Optics and optical instruments – Medical endoscopes and endoscopic accessories – Part 3: Determination of field of view and direction of view of endoscopes with optics o ISO 8600-4: Endoscopes – Medical endoscopes and endotherapy devices – Part 4: Determination of maximum width of insertion portion o ISO 8600-5: Optics and photonics – Medical endoscopes and endotherapy devices – Part 5: Determination of optical resolution of rigid endoscopes with optics o ISO 8600-6: Optics and photonics – Medical endoscopes and endotherapy devices – Part 6: vocabulary INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided. MAINTENANCE/USER REQUIREMENTS Limited maintenance requirements. Standard colposcopes may require replacement of worn parts including lamps, eyepiece rings, light guides and fuses. Follow device-specific service manual, as instructions are specific to each colposcope model. Spare parts availability post-warranty 8 years minimum, starting from installation/commissioning ENVIRONMENTAL CONDITIONS Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure." ESTIMATED LIFE SPAN 10 years WARRANTY Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided." | GENERAL DESCRIPTION A device used to examine the lower genital tract epithelium, including tissues of the cervix, vulva, vagina, and anogenital areas. INTENDED USE A colposcope is a low magnification, light-illuminated visualization device for examining the cervix, across an area measuring approximately 20 to 30 mm in diameter, with enough distance between the colposcope lens and the cervix to accommodate the surgical instruments needed for the examination and/or treatment. It allows the examiner to view the epithelial tissues of the cervix and other anogenital areas. PRODUCT DESCRIPTION A colposcope is a low magnification, light-illuminated visualization instrument primarily used alongside screening tools for triaging, diagnosing and managing precancerous cervical lesions in women. It allows the examiner to view the epithelial tissues of the cervix and other anogenital areas. For purposes of cervical precancer assessment, it helps determine the transformation zone type and the grade of suspected epithelial abnormality. In addition, colposcopy facilitates and optimises biopsy and excisional treatment. FUNCTIONAL REQUIREMENTS 1) the colposcope head housing the optics capable of magnification and focus 2) the light source and green light filters for improved visualization of vasculature 3) the body or stand to facilitate and optimises biopsy and excisional treatment TECHNICAL SPECIFICATIONS : DETAILED REQUIREMENTS - Magnification: A range of optical magnification between 3x to 15x (either stepped or continuously variable) - Illumination: Light sources shall be either halogen or LED to guarantee full-spectrum visible light (white light) - Halogen light: 15 V/150 W - LED: 20,000-35,000 LUX (at 300 mm working distance) - An illumination adjustment knob to change the intensity of light - A fan to cool the bulbs (if halogen bulbs are used) - Green light filters; however, blue filters are also acceptable but not the preferred option - Ingress protection rating: IPX2 (minimum) USER ADJUSTABLE SETTINGS - Coarse and fine optical magnification - Illumination UTILITY REQUIREMENTS The unit is suggested to be connected to a reliable power source. • Electrical source requirements (based on country/setting of use): o Plug type: Based on the country settings ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS • Stand or mount to allow for hands-free operation • LED TV or medical grade monitor if not integrated (optional) • Single-use sheath (if using invasive portable model) • Lamp bulbs and fuses TRANSPORTAION AND STORAGE Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% (non-condensing), and be protected from dripping water (IPX2)." PACKAGING AND LABELLING Labelling on the primary packaging should include the name and/ or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements. DOCUMENTATION REQUIREMENTS • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers." REGULATORY & CONFORMITY REQUIREMENTS Approval by regulatory body of country of manufacturer (if applicable) CE mark Comply to EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate and Declaration of Conformity for the CE mark is applicable. Note : For manufacturers which are supplying CE certificates and relevant documentation under the MDD, a declaration letter from the manufacturer that it is working towards compliance to MDR is required. Additionally, the manufacturer shall submit objective evidence to demonstrate compliance to MDR within one (1) year after supply awards have been given and signed. UNFPA reserves the right to change the probationary period of one year to an appropriate length of time which could be shorter than one year. OR US FDA 510(K) clearance CE certificate available for Class IIa and above devices Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan. Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 CLASSIFICATION Classified under EU MDR 2017/745 as Class IIa device SAFETY & PRODUCT STANDARDS Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements" "• IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests For vaginally-inserted colposcopes: • Biocompatibility: o ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process o ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity o ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization • Endoscopy (pertaining to inserted scope): o ISO 8600-1: Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements o ISO 8600-3: Optics and optical instruments – Medical endoscopes and endoscopic accessories – Part 3: Determination of field of view and direction of view of endoscopes with optics o ISO 8600-4: Endoscopes – Medical endoscopes and endotherapy devices – Part 4: Determination of maximum width of insertion portion o ISO 8600-5: Optics and photonics – Medical endoscopes and endotherapy devices – Part 5: Determination of optical resolution of rigid endoscopes with optics o ISO 8600-6: Optics and photonics – Medical endoscopes and endotherapy devices – Part 6: vocabulary INSTALLATION REQUIREMENTS This product does not require assembly or installation. TRAINING REQUIREMENTS Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided. MAINTENANCE/USER REQUIREMENTS Limited maintenance requirements. Standard colposcopes may require replacement of worn parts including lamps, eyepiece rings, light guides and fuses. Follow device-specific service manual, as instructions are specific to each colposcope model. Spare parts availability post-warranty 8 years minimum, starting from installation/commissioning ENVIRONMENTAL CONDITIONS Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure." ESTIMATED LIFE SPAN 10 years WARRANTY Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided."
528 | In Vitro Diagnostics |  | HIV 1+2 rapid test (Colloidal Test) | HIV 1+2 rapid test (Colloidal Test) | Product name: HIV test - 1+2 assay (Colloidal Test) Product code Type of Immunoassay/ Analyte: Inmuno-Chromatographic lateral flow device in a cassette format. Colloidal gold conjugated recombinant antigens corresponding to HIV-1 (gp120, gp41) and HIV-2 (gp-36) Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ capillary blood Accessories: Safety lancet, alcohol swab, disposable pipette, clock or timer, specimen collection container, centrifuge, biohazard waste container. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s Cassettes Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 4 - 30 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV test - 1+2 assay (Colloidal Test) Product code Type of Immunoassay/ Analyte: Inmuno-Chromatographic lateral flow device in a cassette format. Colloidal gold conjugated recombinant antigens corresponding to HIV-1 (gp120, gp41) and HIV-2 (gp-36) Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ capillary blood Accessories: Safety lancet, alcohol swab, disposable pipette, clock or timer, specimen collection container, centrifuge, biohazard waste container. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s Cassettes Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 4 - 30 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
529 | In Vitro Diagnostics |  | HIV 1/2 Rapid Antibody Test | HIV 1/2 Rapid Antibody Test | Product name: HIV test - 1/2 assay Product code Type of Immunoassay/ Analyte: Qualitative immunochromatographic test for the detection of IgG antibodies to HIV-1 and HIV-2 Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ capillary blood Accessories: Pouch, Test stands, Specimen collection loops Items required but not provided: Alcohol swabs, Sterile lancets Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA needs and products availability Shelf life: 24 /36 months - 30 months Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 2- 30 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV test - 1/2 assay Product code Type of Immunoassay/ Analyte: Qualitative immunochromatographic test for the detection of IgG antibodies to HIV-1 and HIV-2 Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ capillary blood Accessories: Pouch, Test stands, Specimen collection loops Items required but not provided: Alcohol swabs, Sterile lancets Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA needs and products availability Shelf life: 24 /36 months - 30 months Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 2- 30 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
530 | In Vitro Diagnostics |  | Anti-HIV 1+2 ELISA | Anti-HIV 1+2 ELISA | Product name: Anti-HIV 1+2 ELISA Product code Type of Immunoassay/ Analyte: Polystyrene microwell strips pre-coated with recombinant HIV antigens expressed in E.coli (recombinant HIV-1gp41, gp120, and recombinant HIV-2 gp36) Name of reagent Product code Specimen or sample type to be used: serum or plasma Accessories: 5 x 96wells, Microwell Plate 3 x 1ml vial, Negative Control 3 x 1ml vial, Positive Control-1 3 x 1ml vial, Postive Control-2 5 x 12ml vial, Hrp-Conjugate 2 x 125ml vial, Wash Buffer 1 x 60ml vial, Chromogen Solution A 1 x 60ml vial, Chromogen Solution B 1 x 60ml vial, Stop Solution 5, Plastic Sealable Bag 1, IFU 15, Cardboard Plate Cover Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s, 500s Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months 18 months. Do not use beyond its expiry date. Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 2- 8 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Anti-HIV 1+2 ELISA Product code Type of Immunoassay/ Analyte: Polystyrene microwell strips pre-coated with recombinant HIV antigens expressed in E.coli (recombinant HIV-1gp41, gp120, and recombinant HIV-2 gp36) Name of reagent Product code Specimen or sample type to be used: serum or plasma Accessories: 5 x 96wells, Microwell Plate 3 x 1ml vial, Negative Control 3 x 1ml vial, Positive Control-1 3 x 1ml vial, Postive Control-2 5 x 12ml vial, Hrp-Conjugate 2 x 125ml vial, Wash Buffer 1 x 60ml vial, Chromogen Solution A 1 x 60ml vial, Chromogen Solution B 1 x 60ml vial, Stop Solution 5, Plastic Sealable Bag 1, IFU 15, Cardboard Plate Cover Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s, 500s Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months 18 months. Do not use beyond its expiry date. Instructions for use Transport Storage and operating Temperature: Storage conditions: approximately 2- 8 C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
531 | In Vitro Diagnostics |  | Anti - HCV Rapid Test | Anti - HCV Rapid Test | Product name: Anti - HCV Rapid Test Product code Type of Immunoassay/ Analyte: Colloidal gold enhanced, rapid immunochromatographic assay for qualitative detection of antibodies to hepatitis C Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Test cassette, Dropper, Sample diluent, Sterile Safety lancet, Alcohol swab, package insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-30°C Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Anti - HCV Rapid Test Product code Type of Immunoassay/ Analyte: Colloidal gold enhanced, rapid immunochromatographic assay for qualitative detection of antibodies to hepatitis C Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Test cassette, Dropper, Sample diluent, Sterile Safety lancet, Alcohol swab, package insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-30°C Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
532 | In Vitro Diagnostics |  | HCV Antibody Test | HCV Antibody Test | Product name: HCV Antibody Test Product code Type of Immunoassay/ Analyte: Synthetic peptides and recombinant proteins from the core, NS3, and NS4 regions of the HCV genome (test line) and a goat anti-human IgG (control line) immobilized onto a nitrocellulose membrane. Name of reagent Product code Specimen or sample type to be used: Oral fluid, fingerstick whole blood, venipuncture whole blood, plasma specimens and serum Accessories: Test cassette, Dropper, Sample diluent, Sterile Safety lancet, Alcohol swab, package insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelflife: 24 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test kit should be stored at 2-30°C Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HCV Antibody Test Product code Type of Immunoassay/ Analyte: Synthetic peptides and recombinant proteins from the core, NS3, and NS4 regions of the HCV genome (test line) and a goat anti-human IgG (control line) immobilized onto a nitrocellulose membrane. Name of reagent Product code Specimen or sample type to be used: Oral fluid, fingerstick whole blood, venipuncture whole blood, plasma specimens and serum Accessories: Test cassette, Dropper, Sample diluent, Sterile Safety lancet, Alcohol swab, package insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 25, 50s, 100s 500 Tests Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelflife: 24 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test kit should be stored at 2-30°C Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
533 | In Vitro Diagnostics |  | Test Hepatitis B surface antigen (HBsAg) | Test Hepatitis B surface antigen (HBsAg) | Product name: Test Hepatitis B surface antigen (HBsAg) Product code Type of Immunoassay/ Analyte: In vitro immunochromatographic, rapid assay designed for the qualitative detection of Hepatitis B surface antigen Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ Serum Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test kit should be stored at 1 to 40 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Test Hepatitis B surface antigen (HBsAg) Product code Type of Immunoassay/ Analyte: In vitro immunochromatographic, rapid assay designed for the qualitative detection of Hepatitis B surface antigen Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ Serum Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test kit should be stored at 1 to 40 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
534 | In Vitro Diagnostics |  | Test Hepatitis C virus (HCV) | Test Hepatitis C virus (HCV) | Product name: Test Hepatitis C virus (HCV) Product code Type of Immunoassay/ Analyte: In vitro immunochromatographic, rapid assay designed for the qualitative detection of antibodies specific to HCV (core, NS3, NS4 and NS5) Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ Serum Accessories: Specimen procedure(s), Test devices with desiccant, in individual foil pouch, Assay diluent, Capillary pipette(s), Sterile lancet(s), Alcohol swap(s), Insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temeprature Storage conditions: The test kit should be stored at 1 °C to 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temeprature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Test Hepatitis C virus (HCV) Product code Type of Immunoassay/ Analyte: In vitro immunochromatographic, rapid assay designed for the qualitative detection of antibodies specific to HCV (core, NS3, NS4 and NS5) Name of reagent Product code Specimen or sample type to be used: Venous whole blood / plasma o/ Serum Accessories: Specimen procedure(s), Test devices with desiccant, in individual foil pouch, Assay diluent, Capillary pipette(s), Sterile lancet(s), Alcohol swap(s), Insert Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temeprature Storage conditions: The test kit should be stored at 1 °C to 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temeprature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
535 | In Vitro Diagnostics |  | HPV Test (Professional Use Only) | HPV Test (Professional Use Only) | Product name: HPV Test (Professional Use Only) Product code Type of Immunoassay/ Analyte: In vitro nucleic acid hybridization assay with signal amplification using microplate chemiluminescence for the qualitative detection of 14 highrisk types of HPV DNA. Name of reagent Product code Specimen or sample type to be used: Cervical specimens Accessories: Assay microplate, Negative Calibrator, Positive Calibrator, Reagent 1 (purple label), Indicator Dye, Stabilized biologics, Reconstitution diluents Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: in women 30 years and older ONLY Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored between 4°C and 25°C (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HPV Test (Professional Use Only) Product code Type of Immunoassay/ Analyte: In vitro nucleic acid hybridization assay with signal amplification using microplate chemiluminescence for the qualitative detection of 14 highrisk types of HPV DNA. Name of reagent Product code Specimen or sample type to be used: Cervical specimens Accessories: Assay microplate, Negative Calibrator, Positive Calibrator, Reagent 1 (purple label), Indicator Dye, Stabilized biologics, Reconstitution diluents Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: in women 30 years and older ONLY Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored between 4°C and 25°C (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
536 | In Vitro Diagnostics |  | HPV Test | HPV Test | Product name: HPV Test Product code Type of Immunoassay/ Analyte: In vitro test for the detection of the E6/E7 region of the viral DNA genome from high risk Human Papillomavirus (HPV). Name of reagent Product code Specimen or sample type to be used: Cervical specimens Accessories: Xpert HPV assay cartridges with integrated reaction tubes, Disposable 1 mL transfer pipettes, CD with Assay Definition Files (ADF) and instructions for use Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population (Women) Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-28 °C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HPV Test Product code Type of Immunoassay/ Analyte: In vitro test for the detection of the E6/E7 region of the viral DNA genome from high risk Human Papillomavirus (HPV). Name of reagent Product code Specimen or sample type to be used: Cervical specimens Accessories: Xpert HPV assay cartridges with integrated reaction tubes, Disposable 1 mL transfer pipettes, CD with Assay Definition Files (ADF) and instructions for use Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population (Women) Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-28 °C (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
537 | In Vitro Diagnostics |  | HIV Blot Elisa Test | HIV Blot Elisa Test | Product name: HIV Blot Test 2.2 Product code Type of Immunoassay/ Analyte: Qualitative enzyme immunoassay for the in vitro detection of antibodies to HIV-1 and HIV-2 in human serum or plasma Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Nitrocellulose Strips, Non-reactive control, Strong reactive control, Weak reactive control, Stock Buffer concentrate, Wash buffer concentrate, Conjugate, Substrate, Blotting powder, Incubation Tray, Insert, Forceps Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV Blot Test 2.2 Product code Type of Immunoassay/ Analyte: Qualitative enzyme immunoassay for the in vitro detection of antibodies to HIV-1 and HIV-2 in human serum or plasma Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Nitrocellulose Strips, Non-reactive control, Strong reactive control, Weak reactive control, Stock Buffer concentrate, Wash buffer concentrate, Conjugate, Substrate, Blotting powder, Incubation Tray, Insert, Forceps Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
538 | In Vitro Diagnostics |  | Anti-HIV 1+2 ELISA Blot | Anti-HIV 1+2 ELISA Blot | Product name: Anti-HIV 1+2 ELISA Product code Type of Immunoassay/ Analyte: A two-step incubation antigen “sandwich” enzyme immunoassay, which uses polystyrene microwell strips pre-coated with recombinant HIV antigens expressed in E.coli (recombinant HIV-1gp41, gp120, and recombinant HIV-2 gp36). Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Microwell Plate, Negative Control, Positive Control-1, Positive Control-2, Hrp-Conjugate, Wash Buffer, Chromogen Solution A, Chromogen Solution B, Stop Solution, Plastic Sealable Bag, Cardboard Plate Cover, IFU. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Anti-HIV 1+2 ELISA Product code Type of Immunoassay/ Analyte: A two-step incubation antigen “sandwich” enzyme immunoassay, which uses polystyrene microwell strips pre-coated with recombinant HIV antigens expressed in E.coli (recombinant HIV-1gp41, gp120, and recombinant HIV-2 gp36). Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Microwell Plate, Negative Control, Positive Control-1, Positive Control-2, Hrp-Conjugate, Wash Buffer, Chromogen Solution A, Chromogen Solution B, Stop Solution, Plastic Sealable Bag, Cardboard Plate Cover, IFU. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
539 | In Vitro Diagnostics |  | HIV Ag/Ab Combination | HIV Ag/Ab Combination | Product name: HIV Ag/Ab Combination Product code Type of Immunoassay/ Analyte: An enzyme immunoassay for the simultaneous qualitative detection of Human Immunodeficiency Virus (HIV) p24 antigen and antibodies to HIV type 1 (HIV-1 group O) and HIV type 2 (HIV-2) in human serum or plasma. Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Coated well, Sample diluent, Conjugate, Conjugate Diluent, Anti-HIV-1 Positive Control, Anti-HIV-2 Positive Control, HIV-1 p24 Positive Control, Negative Control, Substrate Diluent, Substrate Concentrate, Wash Fluid Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV Ag/Ab Combination Product code Type of Immunoassay/ Analyte: An enzyme immunoassay for the simultaneous qualitative detection of Human Immunodeficiency Virus (HIV) p24 antigen and antibodies to HIV type 1 (HIV-1 group O) and HIV type 2 (HIV-2) in human serum or plasma. Name of reagent Product code Specimen or sample type to be used: Serum or plasma Accessories: Coated well, Sample diluent, Conjugate, Conjugate Diluent, Anti-HIV-1 Positive Control, Anti-HIV-2 Positive Control, HIV-1 p24 Positive Control, Negative Control, Substrate Diluent, Substrate Concentrate, Wash Fluid Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable: Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2-8 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
540 | In Vitro Diagnostics |  | HIV (1+2) antibody (colloidal gold) | HIV (1+2) antibody (colloidal gold) | Product name: HIV (1+2) antibody (colloidal gold) Diagnostic Kit Product code Type of Immunoassay/ Analyte: An in vitro, visually read test for the qualitative determination of antibodies against HIV-1 and HIV-2 in human Name of reagent Product code Specimen or sample type to be used: serum, plasma, venous and capillary whole blood. Accessories: Coated well, Sample diluent, Conjugate, Conjugate Diluent, Anti-HIV-1 Positive Control, Anti-HIV-2 Positive Control, HIV-1 p24 Positive Control, Negative Control, Substrate Diluent, Substrate Concentrate, Wash Fluid Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 4-30°C (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV (1+2) antibody (colloidal gold) Diagnostic Kit Product code Type of Immunoassay/ Analyte: An in vitro, visually read test for the qualitative determination of antibodies against HIV-1 and HIV-2 in human Name of reagent Product code Specimen or sample type to be used: serum, plasma, venous and capillary whole blood. Accessories: Safety lancet, alcohol swab, disposable transfer pipette Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 4-30°C (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
541 | In Vitro Diagnostics |  | HIV1+2/Syphilis | HIV1+2/Syphilis | Product name: HIV 1+2/Syphilis Product code Type of Immunoassay/ Analyte: It is a rapid, qualitative test for the detection of antibodies to all isotypes (IgG, IgM, IgA) specific to HIV1 including subtype-O, HIV-2 and Syphilis (Treponema pallidum) in human serum, plasma or whole blood. Name of reagent Product code Specimen or sample type to be used: human serum, plasma or whole blood. Accessories: Test devices, Assay diluent, Specimen transfer devices (capillary pipettes), Lancets, Alcohol swabs Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 1 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV 1+2/Syphilis Product code Type of Immunoassay/ Analyte: It is a rapid, qualitative test for the detection of antibodies to all isotypes (IgG, IgM, IgA) specific to HIV1 including subtype-O, HIV-2 and Syphilis (Treponema pallidum) in human serum, plasma or whole blood. Name of reagent Product code Specimen or sample type to be used: human serum, plasma or whole blood. Accessories: Test devices, Assay diluent, Specimen transfer devices (capillary pipettes), Lancets, Alcohol swabs Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 1 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
542 | In Vitro Diagnostics |  | Syphilis Rapid Test | Syphilis Rapid Test | Product name: Syphilis Rapid Test Product code Type of Immunoassay/ Analyte: It is a rapid, qualitative, screening, in vitro diagnostic test for the detection of antibodies of all classes specific to Treponema pallidum in human serum, plasma or venous or capillary whole blood Name of reagent Product code Specimen or sample type to be used: human serum, plasma or whole blood. Accessories: Test devices pouch, Assay buffer bottle, Specimen transfer devices, Lancets, Alcohol swabs Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 4 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Syphilis Rapid Test Product code Type of Immunoassay/ Analyte: It is a rapid, qualitative, screening, in vitro diagnostic test for the detection of antibodies of all classes specific to Treponema pallidum in human serum, plasma or venous or capillary whole blood Name of reagent Product code Specimen or sample type to be used: human serum, plasma or whole blood. Accessories: Test devices pouch, Assay buffer bottle, Specimen transfer devices, Lancets, Alcohol swabs Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 4 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use Environmental requirements: Rohs compliance
543 | In Vitro Diagnostics |  | HIV Self Test | HIV Self Test | Product name: HIV Self Test Product code Type of Immunoassay/ Analyte: A visually read, qualitative immunochromatographic test for the detection of IgG antibodies to HIV-1 and HIV-2 in oral fluid Name of reagent Product code Specimen or sample type to be used: Oral fluid Accessories: Pouch, single use test device, desiccant, developer solution vial, test stand, disposal bag Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV Self Test Product code Type of Immunoassay/ Analyte: A visually read, qualitative immunochromatographic test for the detection of IgG antibodies to HIV-1 and HIV-2 in oral fluid Name of reagent Product code Specimen or sample type to be used: Oral fluid Accessories: Pouch, single use test device, desiccant, developer solution vial, test stand, disposal bag Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
544 | In Vitro Diagnostics |  | HIV AntibodyTest | HIV AntibodyTest | Product name: HIV Antibody Test Product code Type of Immunoassay/ Analyte: A qualitative, in vitro immunoassay. It detects antibodies to the human immunodeficiency virus types 1 and 2 (HIV-1/2) in human oral fluid, whole blood, serum or plasma (EDTA). Name of reagent Product code Specimen or sample type to be used: oral fluid, whole blood, serum or plasma (EDTA). Accessories: Pouch, Specimen collection loops, test stand. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temeprature Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: HIV Antibody Test Product code Type of Immunoassay/ Analyte: A qualitative, in vitro immunoassay. It detects antibodies to the human immunodeficiency virus types 1 and 2 (HIV-1/2) in human oral fluid, whole blood, serum or plasma (EDTA). Name of reagent Product code Specimen or sample type to be used: oral fluid, whole blood, serum or plasma (EDTA). Accessories: Pouch, Specimen collection loops, test stand. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 30, 50s, 100s test pouch Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temeprature Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
545 | In Vitro Diagnostics |  | Urine Test – strips | Urine Test – strips | Product name: Urine Test – strips Product code Description: Test strips for urine analysis. Parameters: Glucose neg, 100, 250, 500, 1000, 2000 mg/dl, or more Protein neg, trace, 30, 100, 300, 2000 mg/dl, or more pH 5.0, 6.0, 6.5, 7.0, 7.5, 8.0, 8.5 Leukocytes neg, trace, +, ++, +++ Nitrites neg, pos Ketone neg, 5, 15, 40, 80, 150 mg/dl Bilirubin neg, +, ++, +++ Blood neg, trace, moderate, hemolyzed, +, ++, +++ Urobilinogen 0.2, 1, 2, 4, 8 mg/dl Specific gravity 1.000, 1.005, 1.010, 1.015, 1.020, 1.025, 1.030 Colour reference card is water resistant, wet-wipeable Primary packaging – unbreakable and sealed Reading time: Approximately 30 seconds to 2 minutes Specimen or sample type to be used: Orine Accessories: Pouch, Specimen collection loops, test stand. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: Test strip, urine, 10 parameter, box/100 + Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Urine Test – strips Product code Description: Test strips for urine analysis. Parameters: Glucose neg, 100, 250, 500, 1000, 2000 mg/dl, or more Protein neg, trace, 30, 100, 300, 2000 mg/dl, or more pH 5.0, 6.0, 6.5, 7.0, 7.5, 8.0, 8.5 Leukocytes neg, trace, +, ++, +++ Nitrites neg, pos Ketone neg, 5, 15, 40, 80, 150 mg/dl Bilirubin neg, +, ++, +++ Blood neg, trace, moderate, hemolyzed, +, ++, +++ Urobilinogen 0.2, 1, 2, 4, 8 mg/dl Specific gravity 1.000, 1.005, 1.010, 1.015, 1.020, 1.025, 1.030 Colour reference card is water resistant, wet-wipeable Primary packaging – unbreakable and sealed Reading time: Approximately 30 seconds to 2 minutes Specimen or sample type to be used: Orine Accessories: Pouch, Specimen collection loops, test stand. Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: Test strip, urine, 10 parameter, box/100 + Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature Storage conditions: The test should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulatory & Conformity Requirements: CE mark conforming to Medical Device Regulation 2017/ 745 or US FDA 510(K) clearance CE certificate available for Class IIa and above devices Manufacturer’s Quality Management System is certified to ISO 13485 : 2016 Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Classified under EU MDR 2017/745 as Class I device Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
546 | In Vitro Diagnostics |  | Blood Test - Anti A | Blood Test - Anti A | Product name: Blood Test - Anti A Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, A in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Blood Test - Anti A Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, A in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelflife: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
547 | In Vitro Diagnostics |  | Blood Test - Anti B | Blood Test - Anti B | Product name: Blood Test - Anti B Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, B in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Blood Test - Anti B Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, B in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
548 | In Vitro Diagnostics |  | Blood Test - Anti A+B | Blood Test - Anti A+B | Product name: Blood Test - Anti A+B Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, AB in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Blood Test - Anti A+B Product code Type of Immunoassay/ Analyte: The reagents will cause direct agglutination (clumping) of test red cells that carry the corresponding ABO antigen, AB in this case. No agglutination generally indicates absence of the corresponding ABO antigen. Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: All population Packaging and labelling: Pack size: 10 + ml Bottle Carttrige guide Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions: Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
549 | In Vitro Diagnostics |  | Test Rhesus, anti D | Test Rhesus, anti D | Product name: Test Rhesus, anti D Product code Type of Immunoassay/ Analyte: Anti-D blood grouping reagent contains monoclonal murine IgM antibody Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: Pregnant Women Packaging and labelling: Pack size: 10 + ml Bottle Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Test Rhesus, anti D Product code Type of Immunoassay/ Analyte: Anti-D blood grouping reagent contains monoclonal murine IgM antibody Name of reagent Product code Specimen or sample type to be used: Venous whole blood Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Product code Population: Pregnant Women Packaging and labelling: Pack size: 10 + ml Bottle Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test kit should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance
550 | In Vitro Diagnostics |  | Pregnancy Test, Strips | Pregnancy Test, Strips | Product name: Pregnancy Test, Strips Product code Type of Immunoassay/ Analyte: The pregnancy test detects the presence of human chorionic gonadotropin ("HCG"), which appears in urine very early in pregnancy Name of reagent Product code Specimen or sample type to be used: Urine Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Population: ALL POPULATION Packaging and labelling: Pack size: Box/50, 100 Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use IEC 61010-2-101:2002 Safety requirements for electrical equipment for measurement, control, and laboratory use -- Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment IEC 61326-2-6:2005 Electrical equipment for measurement, control and laboratory use - EMC requirements -- Part 2-6: Particular requirements - In vitro diagnostic (IVD) medical equipment IEC 62304:2006 Medical device software - Software life-cycle processes IEC 62366:2007 Medical devices - Application of usability engineering to medical devices Environmental requirements: Rohs compliance | Product name: Pregnancy Test, Strips Product code Type of Immunoassay/ Analyte: The pregnancy test detects the presence of human chorionic gonadotropin ("HCG"), which appears in urine very early in pregnancy Name of reagent Product code Specimen or sample type to be used: Urine Accessories: Note: Accessories requirements must be specified by Suppliers whether they are included or not in the package. Name of instruments or component if applicable Population: ALL POPULATION Packaging and labelling: Pack size: Box/50, 100 Note: Pack Sizes are not limited to the ones indicated above. It will be considered bigger or smaller ones according to the UNFPA Needs and products availability. Shelf life: 24 /36 months Instructions for use Transport Storage and operating Temperature: Storage conditions: The test should be stored at 2 - 30 °C. (or as per stability studies guidelines) Do not freeze Transportation conditions Operating Temperature: use at 15-25 within 10 minutes (or as IFU guidelines) Regulation & conformity requirements: CE mark conforming to In Vitro Device Directive Regulation (EU) 2017/745 - CE self-declaration ISO 13485 certificate Note: CE marking is only obligatory for products for which EU specifications exist and require the affixing of CE marking. Classification: Self declaration for class I CE certification with notified body (Class IIa and higher) Safety & product Standards: Must comply with following standards EN 13612:2002 Performance evaluation of in vitro diagnostic medical devices EN 13640:2002 Stability testing of in vitro diagnostic reagents EN 13641:2002 Elimination or reduction of risk of infection related to in vitro diagnostic reagents EN 13975:2003 Sampling procedures used for acceptance testing of in vitro diagnostic medical devices - Statistical aspects EN 14136:2004 Use of external quality assessment schemes in the assessment of the performance of in vitro diagnostic examination procedures ISO 15193:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for content and presentation of reference measurement procedures ISO 15194:2009 In vitro diagnostic medical devices - Measurement of quantities in samples of biological origin - Requirements for certified reference materials and the content of supporting documentation ISO 17511:2003 In vitro diagnostic medical devices - Measurement of quantities in biological samples - Metrological traceability of values assigned to calibrators and control materials ISO 18113-1:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 1: Terms, definitions and general requirements ISO 18113-3:2009 In vitro diagnostic medical devices - Information supplied by the manufacturer (labelling) - Part 3: In vitro diagnostic instruments for professional use Environmental requirements: Rohs compliance

## 6. Technical Spec Other Items
Tab 6. Technical Specification Form for Other Items
Important guidance to Bidders:
1. Please fill in the blanks under green column.
2. Technical Specification Form should be submitted in excel form.
3. Complete every single row.
4. If possible, highlight all differences or deviations between offered product and UNFPA specifications with a different color font (red).
Bid Number | UNFPA Short description | UNFPA Detailed Technical Specifications | Bidders technical specifications of the offered product
565 | Bag, shoulder, UNFPA logo, vinyl, approx 360 x 230 x 610 mm | Multi-purpose bag with shoulder strip. - Water-proof. - Material: vinyl. - Two carrying functions: carried by hand and on the shoulder. - With adjustable shoulder strips attachable to the rings on both sides of the bag. - The rings for the shoulder strips are reinforced with heavy duty stitching. - With a hand handle in the top middle of the bag made of reinforced heavy duty stitching. - With a strong zipper. - Stiff and hard, extra-strong and reinforced bottom lining allowing the possibility of traditional birth attendants to carry the bag on top of their heads. - Reusable and durable enough to carry the kit contents for several months. - With UNFPA logo. - Size: approx. 360 x 230 x 610 mm. - Please see link for information on UNFPA logo and pantone colours: http://www.unfpa.org/styleguide/components/palette.htm
566 | Battery, dry cell, AA, alkaline, 1.5 V, for Photometer | Battery, dry cell, AA, alkaline, 1.5 V, for Photometer
567 | Flashlight, frontal, LED, battery-operated, including batteries | High quality ultra bright white light frontal lamp with 7 LED. - High intensity and durable lifetime. - High impact plastic handle. - Waterproof. - The frontal lamp can be attached to a forehead of a user by head straps threaded through strap holders on the rear of the frontal lamp. - Handy switch point for easier operation. - Tilting lamp. - With adjustable head straps. - Battery operated with AA batteries. - With adapter for alternative use with electrical socket. - Operation life span: minimum 100,000 hours. - Supplied with: - 1 pair of head straps. - 1 adapter. - AA batteries required to operate the frontal lamp. - CE mark.
568 | Flashlight, LED, re-chargeable | High quality ultra bright white light touch with 7 LED. - High intensity and durable lifetime. - High impact plastic handle. - Waterproof. - Shock proof. - Handy switch point for easier operation. - With a hand strap, easy to carry. - 30 seconds press or rotation = 15 minutes of light. - 60 seconds press or rotation = 30 minutes of light. - Dynamo or battery operated. - Light visible for minimum 1.5km. - Operation life span: minimum 100,000 hours. - Dimensions: approximately 15 x 5 x 5cm. - Eco-friendly. - Supplied with batteries. - CE mark.
569 | Kerosene storm-lamp + extra socks | Kerosene storm-lamp, feuhand hurricane, non-pressure for emergency lighting use. - Complete with wick and clear globe (Suprax or similar heat resistant). - Capacity: approx. 340ml. - Supplies with 1 spare globe and 3 spare wicks. - Certificate of conformity with ISO 9001.
570 | Poncho, wet weather, reusable, heavy duty | Heavy duty Poncho used as rain coat. - Lightweight and oversized to slip on easily. - Size: Adult. - Re-usable. - Certificate of conformity with ISO 105-C06:2010, ISO 105-X12:2002, ISO 81, ASTM E96-00 and ISO 6330:2000/ Amd 1:2008. - CE mark.
571 | Soap, hand, bar, 110g, wrapped | Bar of hand toilet soap. - Composed of minimum Total Fatty Matter (TFM) 70%, water and volatiles maximum 24%, glycerine 1% and pH alkaline ingredients 5%. - Unperfumed or unscented. - Hypoallergenic. - Individually wrapped. - Weight: 110 to 125g. - Certificate of conformity with : BS 1914: 1990; or ASTM D501-03 (2009)
572 | Tape-measure, vinyl-coated, 1.5 m/5 feet | Tape-measure, vinyl-coated, 1.5 m/5 feet
573 | Towel, huck, 430 x 500 mm | Towel, huck, 430 x 500 mm
575 | Leaflet, Fistula Kits, English | Leaflet, A4, 2-sided leaflet printed for inclusion in kit (1 paper, 2 sides), printed on 200grams mc paper full colour, English
576 | Leaflet, Fistula Kits, French | Leaflet, A4, 2-sided leaflet printed for inclusion in kit (1 paper, 2 sides), printed on 200grams mc paper full colour, French
578 | Binder with lever, 31 x 29 cm, various colours | Binder with lever, 31 x 29 cm, various colours
579 | Marker, erasable, red | Marker, erasable, red
580 | Marker, flipchart, black, tip 4.5 mm | Marker, flipchart, black, tip 4.5 mm
581 | Marker, flipchart, red, tip 4.5 mm | Marker, flipchart, red, tip 4.5 mm
582 | Paper, A4, 21 x 29.7cm, 80g, white | Paper, A4, 21 x 29.7cm, 80g, white
583 | Pen, ballpoint, black | Pen, ballpoint, black
584 | Pen, ballpoint, blue | Pen, ballpoint, blue
585 | Pen, ballpoint, red | Pen, ballpoint, red
586 | Pencil sharpener, tapered | Pencil sharpener, tapered
587 | Pencil, HB | Pencil, HB
588 | Scissors, blended, 170 mm | Scissors, blended, 170 mm
589 | Board, white, adhesive, washable, roll, 67 x 100cm | Board, white, adhesive, washable, roll, 67 x 100cm
590 | Sticker, UNFPA, 11 x 22cm | Sticker, UNFPA, 11 x 22cm
591 | Sticker, UNFPA, 6 x 12cm | Sticker, UNFPA, 6 x 12cm
592 | Tape, adhesive, 19mm x 33m, clear, roll | Tape, adhesive, 19mm x 33m, clear, roll
593 | Book, exercise, 17 x 22 cm, 5mm-squared, 100 pages | Book, exercise, 17 x 22 cm, 5mm-squared, 100 pages
594 | Book, exercise, 21 x 29.7 cm, 5mm-squared, 100 pages | Book, exercise, 21 x 29.7 cm, 5mm-squared, 100 pages
595 | Book, exercise, A4, ruled, 100 pages | Book, exercise, A4, ruled, 100 pages
596 | Calculator, solar-powdered | Calculator, solar-powdered
597 | Glue, stick, large | Glue, stick, large
598 | Marker, erasable, black | Marker, erasable, black
599 | Marker, erasable, blue | Marker, erasable, blue

## 7. Price Kits IARH kits
Tab 7. Prices for the IARH kits
Important guidance to Bidders:
1. Please note that the prices for pharmaceuticals and medical devices are automaticall transferred from Form 1.
2. Manufacturer, Country of Origin and Price of "Other Items" to be added at the end of the tab in rows 546 to 587.
3. Weighted scoring will be used during the commercial and financial evaluation, where the weight score is given per each item based on the past procurement levels. The total score for each Bidder will be the sum of weighted scores for all quoted items, which will identify the lowest-priced bids that will be passed to further Technical Evaluation.
Item No | Module | Kit Name |  | FCA Price/Kit | Lead Time, weeks | Kitting Services, Cost/Kit | Warehousing Cost/Kit | Cost of reconfiguration/Kit | Total Price/Kit | Weighting | Weighted price
1 | Comp. IARH Kit | CC Kit 0 - Administration and Training Kit |  | 0 |  |  |  |  | 0 | 0.06% | 0
2 | Core IARH Kit | Kit 1A - Male condoms |  | 0 |  |  |  |  | 0 | 0.02% | 0
3 | Comp. IARH Kit | CC Kit 1B - Female Condoms |  | 0 |  |  |  |  | 0 | 0.02% | 0
4 | Core IARH Kit | Kit 2A - Clean Delivery Kit: Individual Delivery |  | 0 |  |  |  |  | 0 | 16.72% | 0
5 | Core IARH Kit | Kit 2B - Clean Delivery Kit: Equipment for Birth Attendants |  | 0 |  |  |  |  | 0 | 1.16% | 0
6 | Core IARH Kit | Kit 3 - Post-rape treatment |  | 0 |  |  |  |  | 0 | 13.04% | 0
7 | Core IARH Kit | Kit 4 - Oral and injectable contraception |  | 0 |  |  |  |  | 0 | 2.96% | 0
8 | Core IARH Kit | Kit 5 - Treatment of STIs |  | 0 |  |  |  |  | 0 | 9.09% | 0
9 | Core IARH Kit | Kit 6A - Clinical Delivery Assistance - Reusable Equipment |  | 0 |  |  |  |  | 0 | 7.93% | 0
10 | Core IARH Kit | Kit 6B - Clinical Delivery Assistance – Midwifery Supplies: Drugs and Disposable Equipment |  | 0 |  |  |  |  | 0 | 10.88% | 0
11 | Comp. IARH Kit | CC Kit 7A - Intrauterine Devices (IUD) |  | 0 |  |  |  |  | 0 | 2.21% | 0
12 | Comp. IARH Kit | CC Kit 7B - Contraceptive Implant |  | 0 |  |  |  |  | 0 | 2.21% | 0
13 | Core IARH Kit | Kit 8 - Management of complications of miscarriage or abortion |  | 0 |  |  |  |  | 0 | 4.28% | 0
14 | Core IARH Kit | Kit 9 - Repair of cervical and vaginal tears |  | 0 |  |  |  |  | 0 | 2.99% | 0
15 | Core IARH Kit | Kit 10 - Assisted delivery with vacuum extraction |  | 0 |  |  |  |  | 0 | 2.86% | 0
16 | Core IARH Kit | Kit 11A - Obstetric surgery and severe obstetric complications: Reusable Equipment |  | 0 |  |  |  |  | 0 | 2.72% | 0
17 | Core IARH Kit | Kit 11B - Obstetric surgery and severe obstetric complications: Drugs and Disposable Equipment |  | 0 |  |  |  |  | 0 | 16.33% | 0
18 | Core IARH Kit | Kit 12 - Blood Transfusion |  | 0 |  |  |  |  | 0 | 3.89% | 0
19 | Comp. IARH Kit | CC Kit Non-Pneumatic Anti-Shock Garment |  | 0 |  |  |  |  | 0 | 0.16% | 0
20 | Comp. IARH Kit | CC Kit Oxytocin |  | 0 |  |  |  |  | 0 | 0.02% | 0
21 | Comp. IARH Kit | CC Kit Mifepristone |  | 0 |  |  |  |  | 0 | 0.03% | 0
22 | Comp. IARH Kit | CC Kit Hand-Held Vacuum Assisted Delivery System |  | 0 |  |  |  |  | 0 | 0.29% | 0
23 | Comp. IARH Kit | CC Kit Chlorhexidine Di-gluconate |  | 0 |  |  |  |  | 0 | 0.02% | 0
24 | Comp. IARH Kit | CC Kit Misoprostol, 300 tabs |  | 0 |  |  |  |  | 0 | 0.09% | 0
25 | Comp. IARH Kit | CC Kit Medroxyprogesterone acetate (DMPA-SC) |  | 0 |  |  |  |  | 0 | 0.02% | 0
27 | Medical Kit | MVA kitting service |  | 0 |  |  |  |  | 0 |  | 0
ITEMIZED OVERVIEW OF KITS
Comp. IARH Kit |  | CC Kit 0 - Administration and Training Kit
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 596 | Calculator, solar-powdered |  | 0 | 0 | 1 | 1 | 0 | 0
2 | 589 | Board, white, adhesive, washable, roll, 67 x 100cm |  | 0 | 0 | 2 | 1 | 0 | 0
3 | 578 | Binder with lever, 31 x 29 cm, various colours |  | 0 | 0 | 5 | 1 | 0 | 0
4 | 593 | Book, exercise, 17 x 22 cm, 5mm-squared, 100 pages |  | 0 | 0 | 20 | 1 | 0 | 0
5 | 594 | Book, exercise, 21 x 29.7 cm, 5mm-squared, 100 pages |  | 0 | 0 | 3 | 1 | 0 | 0
6 | 582 | Paper, A4, 21 x 29.7cm, 80g, white |  | 0 | 0 | 1000 | 1 | 0 | 0
7 | 597 | Glue, stick, large |  | 0 | 0 | 3 | 1 | 0 | 0
8 | 591 | Sticker, UNFPA, 6 x 12cm |  | 0 | 0 | 20 | 1 | 0 | 0
9 | 590 | Sticker, UNFPA, 11 x 22cm |  | 0 | 0 | 20 | 1 | 0 | 0
10 | 583 | Pen, ballpoint, black |  | 0 | 0 | 20 | 1 | 0 | 0
11 | 585 | Pen, ballpoint, red |  | 0 | 0 | 20 | 1 | 0 | 0
12 | 598 | Marker, erasable, black |  | 0 | 0 | 12 | 1 | 0 | 0
13 | 579 | Marker, erasable, red |  | 0 | 0 | 12 | 1 | 0 | 0
14 | 599 | Marker, erasable, blue |  | 0 | 0 | 12 | 1 | 0 | 0
15 | 581 | Marker, flipchart, red, tip 4.5 mm |  | 0 | 0 | 12 | 1 | 0 | 0
16 | 580 | Marker, flipchart, black, tip 4.5 mm |  | 0 | 0 | 12 | 1 | 0 | 0
17 | 587 | Pencil, HB |  | 0 | 0 | 20 | 1 | 0 | 0
18 | 586 | Pencil sharpener, tapered |  | 0 | 0 | 2 | 1 | 0 | 0
19 | 588 | Scissors, blended, 170 mm |  | 0 | 0 | 1 | 1 | 0 | 0
20 | 592 | Tape, adhesive, 19mm x 33m, clear, roll |  | 0 | 0 | 4 | 1 | 0 | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | RH KIT 1A - Male condoms
Item No | Bid No | Item Short Description |  | Qty (as per primary UOM) | Sales Pack Size | Unit of Measure | Sales Pack Size | Total Price
1 | 556 | Male condom 53 mm, standard, natural |  | 17280 | 144 | Gross | 120 | TA item
2 | 563 | Leaflet on Use of Male Condoms, English |  | 400 | 1 | Each | 400 | TA item
3 | 564 | Leaflet on Use of Male Condoms, French |  | 400 | 1 | Each | 400 | TA item
[7 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit 1B - Female Condoms
Item No | Bid No | Item Short Description |  | Qty (as per primary UOM) | Sales Pack Size | Unit of Measure | Sales Pack Size | Total Price
1 | 553 | Condomize! condom nitrile sheath inner retention ring, 175mm |  | 540 | 1 | Each | 540 | TA item
2 | 562 | Leaflet on Use of Female Condoms, Multilingual |  | 25 | 1 | Each | 25 | TA item
[7 empty] | TOTAL | 0
Core IARH Kit |  | Kit 2A - Clean Delivery Kit: Individual Delivery
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 197 | Bag (envelope), plastic, for drugs, 18x28cm, snap-lock |  | 0 | 0 | 200 | 0 | 0 | 0
2 | 571 | Soap, hand, bar, 110g, wrapped |  | 0 | 0 | 200 | 1 | 0 | 0
3 | 341 | Draw sheet, plastic, 100 x 100 cm |  | 0 | 0 | 200 | 0 | 0 | 0
4 | 456 | Scissors for umbilical cord cutting, reusable, sterile |  | 0 | 0 | 200 | 0 | 0 | 0
5 | 310 | Tape, umbilical, 3mm x 15cm, non-sterile |  | 0 | 0 | 600 | 0 | 0 | 0
6 | 335 | Cotton cloth, “tetra”, 100 x 100cm |  | 0 | 0 | 400 | 0 | 0 | 0
7 | 204 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 400 | 0 | 0 | 0
8 | N/A | PRINTING MATERIALS 6TH ED KIT 2A |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 2B - Clean Delivery Kit: Equipment for Birth Attendants
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 565 | Bag, shoulder, UNFPA logo, vinyl, approx 360 x 230 x 610 mm |  | 0 | 0 | 5 | 1 | 0 | 0
2 | 205 | Gloves, examination, latex, medium, non-sterile, P100 |  | 0 | 0 | 500 | 0 | 0 | 0
3 | 568 | Flashlight, LED, re-chargeable |  | 0 | 0 | 5 | 1 | 0 | 0
4 | 331 | Apron, protection, plastic, reusable |  | 0 | 0 | 5 | 0 | 0 | 0
5 | 570 | Poncho, wet weather, reusable, heavy duty |  | 0 | 0 | 5 | 1 | 0 | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 3 - Post-rape treatment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 552 | Levonorgestrel 1.5mg, tablet (treatment: single dose) |  | - | - | 100 | 1 | - | TA item
2 | 12 | Azithromycin 250mg, tablets |  | 0 | 0 | 220 | 0 | 0 | 0
3 | 14 | Azithromycin dihydrate, 200mg base/ 5ml suspension, 15ml bottle |  | 0 | 0 | 5 | 0 | 0 | 0
4 | 23 | Cefixime 200mg, tablet |  | 0 | 0 | 110 | 0 | 0 | 0
5 | 25 | Cefixime trihydrate 100mg/5ml powder for oral suspension, 30ml bottle |  | 0 | 0 | 10 | 0 | 0 | 0
6 | 10 | Abacavir (as sulphate) 120mg + Lamivudine 60mg, tablet |  | 0 | 0 | 600 | 0 | 0 | 0
7 | 41 | Dolutegravir, 50mg, tablet |  | 0 | 0 | 120 | 0 | 0 | 0
8 | 42 | Dolutegravir, 10mg, dispersible tablet |  | 0 | 0 | 240 | 0 | 0 | 0
9 | 92 | Tenofovir Disproxil Fumarate 300mg + Lamivudine 300mg + Dolutegravir 50mg, tablet |  | 0 | 0 | 900 | 0 | 0 | 0
10 | 550 | Test pregnancy, strip, temperature stable |  | 0 | 0 | 25 | 0 | 0 | 0
11 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 100 | 0 | 0 | 0
12 | N/A | PRINTING MATERIALS 6TH ED KIT 3 |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 4 - Oral and injectable contraception
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 557 | Ethinylestradiol, 0.03mg, plus levonorgestrel, 0.15mg, combined tablet, 1 strip for 1 cycle |  | - | - | 375 | 3 | - | TA item
2 | 552 | Levonorgestrel 1.5mg, tablet (treatment: single dose) |  | - | - | 60 | 1 | - | TA item
3 | 558 | Levonorgestrel 0.03mg, tablet, 1 strip for 1 cycle |  | - | - | 60 | 3 | - | TA item
4 | 554 | Medroxyprogesterone acetate 150mg/ml, depot injection in 1-ml vial |  | - | - | 300 | 1 | - | TA item
5 | 28 | Chlorhexidine digluconate 5% solution, 1 litre bottle |  | 0 | 0 | 3 | 0 | 0 | 0
6 | 301 | Syringe, luer, 2ml, sterile, single use |  | 0 | 0 | 300 | 0 | 0 | 0
7 | 226 | Needle, luer, 21G, sterile, single use |  | 0 | 0 | 600 | 0 | 0 | 0
8 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 4 | 0 | 0 | 0
9 | 157 | Safety box, disposal of used syringes & needles, 5 litres |  | 0 | 0 | 3 | 0 | 0 | 0
10 | N/A | PRINTING MATERIALS 6TH ED KIT 4 |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 5 - Treatment of STIs
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 15 | Benzathine benzylpenicillin 1.44G (2.4 M.I.U) in 5ml vial |  | 0 | 0 | 65 | 0 | 0 | 0
2 | 16 | Benzathine benzylpenicillin 900mg (1.2 M.I.U) in 5ml vial |  | 0 | 0 | 10 | 0 | 0 | 0
3 | 98 | Water for injection, sterile, 10-ml ampoule |  | 0 | 0 | 80 | 0 | 0 | 0
4 | 12 | Azithromycin 250mg, tablets |  | 0 | 0 | 1120 | 0 | 0 | 0
5 | 14 | Azithromycin dihydrate 200mg base/5ml suspension, 15ml bottle |  | 0 | 0 | 30 | 0 | 0 | 0
6 | 28 | Chlorhexidine digluconate solution 5% in 1 litre bottle |  | 0 | 0 | 3 | 0 | 0 | 0
7 | 23 | Cefixime 200mg, tablet |  | 0 | 0 | 470 | 0 | 0 | 0
8 | 25 | Cefixime trihydrate 100mg/5ml powder for oral suspension, 30ml bottle |  | 0 | 0 | 30 | 0 | 0 | 0
9 | 75 | Metronidazole 250mg tablet |  | 0 | 0 | 2000 | 0 | 0 | 0
10 | 35 | Clotrimazole 500mg + applicator, vaginal tablet |  | 0 | 0 | 100 | 0 | 0 | 0
13 | 556 | Male condom 53 mm, standard, natural |  | - | - | 2880 | 144 | - | TA item
14 | 553 | Condomize! condom nitrile sheath inner retention ring, 175mm |  | - | - | 90 | 1 | - | TA item
15 | 303 | Syringe, luer, 5ml, sterile, single use |  | 0 | 0 | 100 | 0 | 0 | 0
16 | 226 | Needle, luer, 21G, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
17 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 3 | 0 | 0 | 0
18 | 157 | Safety box, disposal of used syringes & needles, 5 litres |  | 0 | 0 | 4 | 0 | 0 | 0
19 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 1000 | 0 | 0 | 0
20 | 563 | Leaflet on Use of Male Condoms, English |  | - | - | 80 | 1 | - | TA item
21 | 564 | Leaflet on Use of Male Condoms, French |  | - | - | 80 | 1 | - | TA item
22 | 562 | Leaflet on Use of Female Condoms, Multilingual |  | - | - | 30 | 1 | - | TA item
23 | N/A | PRINTING MATERIALS 6TH ED KIT 5 |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 6A - Clinical Delivery Assistance - Midwifery Supplies: Reusable Equipment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 356 | Sphygmomanometer, aneroid, adult |  | 0 | 0 | 1 | 0 | 0 | 0
2 | 357 | Stethoscope, binaural, complete |  | 0 | 0 | 1 | 0 | 0 | 0
3 | 358 | Stethoscope, obstetrical, Pinard, monoaural |  | 0 | 0 | 2 | 0 | 0 | 0
4 | 319 | Basin, kidney, stainless steel, 825ml |  | 0 | 0 | 2 | 0 | 0 | 0
5 | 359 | Thermometer, clinical, digital 32°C-43°C |  | 0 | 0 | 4 | 0 | 0 | 0
6 | 220 | Brush, hand, scrubbing, plastic |  | 0 | 0 | 2 | 0 | 0 | 0
7 | 311 | Tourniquet, latex rubber, 75cm |  | 0 | 0 | 2 | 0 | 0 | 0
8 | 107 | Resuscitator, hand-operating, child/newborn |  | 0 | 0 | 1 | 0 | 0 | 0
9 | 329 | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 355 | Scale, infant, spring-type, 5kg x 25g |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 352 | Sling, for use with infant scale |  | 0 | 0 | 5 | 0 | 0 | 0
12 | 360 | Timer, for respiration, ARI |  | 0 | 0 | 3 | 0 | 0 | 0
13 | 473 | Scissors, Mayo, 14cm, curved, blunt/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
14 | 472 | Scissors, gynaecological, 20cm, curved, blunt/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
15 | 400 | Forceps, artery, Kocher, 14cm, straight |  | 0 | 0 | 7 | 0 | 0 | 0
16 | 465 | Scissors, Deaver, 14cm, curved, sharp/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
17 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
18 | 411 | Forceps, tissue, standard, 145mm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
19 | 132 | Stove, kerosene, single-burner, pressure-type |  | 0 | 0 | 1 | 0 | 0 | 0
20 | 131 | Sterilizer, steam, pressure-type, 39 litres |  | 0 | 0 | 1 | 0 | 0 | 0
21 | 133 | Timer, 60 min, mechanical |  | 0 | 0 | 1 | 0 | 0 | 0
22 | 126 | Drum, sterilizing, diameter 165mm |  | 0 | 0 | 2 | 0 | 0 | 0
23 | 127 | Drum, sterilizing, diameter 260mm |  | 0 | 0 | 2 | 0 | 0 | 0
24 | 128 | Drum, sterilizing, diameter 290mm |  | 0 | 0 | 2 | 0 | 0 | 0
25 | 567 | Flashlight, frontal, LED, battery-operated, including batteries |  | 0 | #REF! | 1 | 1 | 0 | 0
26 | 569 | Kerosene storm-lamp + extra socks |  | 0 | 0 | 1 | 1 | 0 | 0
27 | N/A | PRINTING MATERIALS 6TH ED KIT 6A |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 6B - Clinical Delivery Assistance – Midwifery Supplies: Drugs and Disposable Equipment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 6 | Ampicillin sodium 500mg, powder for injection in vial |  | 0 | 0 | 200 | 0 | 0 | 0
2 | 52 | Gentamicin sulphate 40mg base/ml for injection in 2ml ampoule |  | 0 | 0 | 100 | 0 | 0 | 0
3 | 33 | Clindamycin, injection, 600mg (as phosphate) in 4ml ampoule |  | 0 | 0 | 125 | 0 | 0 | 0
4 | 1 | Amoxicillin, powder for oral suspension, 125mg/5ml, 100 ml bottle |  | 0 | 0 | 5 | 0 | 0 | 0
5 | 80.2 | Misoprostol 200mcg, tablet |  | 0 | 0 | 60 | 0 | 0 | 0
3 | 48 | Ferrous fumarate 185mg (60mg iron)/folic acid 0.4mg, tablet |  | 0 | 0 | 5000 | 0 | 0 | 0
7 | 93 | Tetracycline hydrochloride 1%, eye-ointment, 5g tube - CANCELLED: TO BE REPLACED |  | 0 | 0 | 6 | 0 | 0 | 0
8 | 66 | Lidocaine hydrochloride 1% , 20ml ampoule USP or BP or equivalent |  | 0 | 0 | 20 | 0 | 0 | 0
9 | 84 | Oxytocin 10 IU/ml for injection in 1ml ampoule (keep cold: 2-8 °C) |  | 0 | 0 | 100 | 0 | 0 | 0
10 | 57 | Sodium lactate (Ringer’s lactate), IV infusion, 1 litre + giving set |  | 0 | 0 | 30 | 0 | 0 | 0
11 | 55 | Glucose 5%, isotonic, 1 litre + infusion set, sterile, single use |  | 0 | 0 | 20 | 0 | 0 | 0
12 | 72 | Magnesium sulphate 500mg/ml injection in 10ml ampoule |  | 0 | 0 | 10 | 0 | 0 | 0
13 | 59 | Hydralazine hydrochloride 20mg/2ml for injection in 2ml ampoule |  | 0 | 0 | 60 | 0 | 0 | 0
14 | 21 | Calcium gluconate 100mg base/ml for injection in 10ml ampoule |  | 0 | 0 | 20 | 0 | 0 | 0
15 | 85 | Paracetamol 500mg, tablet |  | 0 | 0 | 1000 | 0 | 0 | 0
16 | 96 | Vitamin K, injection, 2mg/0.2ml, 0.2 ml vial |  | 0 | 0 | 50 | 0 | 0 | 0
17 | 273 | Syringe, 0.5ml, with 0.01-ml increments |  | 0 | 0 | 50 | 0 | 0 | 0
18 | 89 | Sodium Chloride 0.9%, injection, 10 ml ampoule |  | 0 | 0 | 80 | 0 | 0 | 0
19 | 98 | Water for injection, sterile, 10-ml ampoule |  | 0 | 0 | 260 | 0 | 0 | 0
20 | 28 | Chlorhexidine digluconate 5% solution, 1 litre |  | 0 | 0 | 12 | 0 | 0 | 0
21 | 31 | Chlorhexidine gluconate 4% solution (Hibiscrub), 500ml bottle |  | 0 | 0 | 1 | 0 | 0 | 0
22 | 86 | Povidone-iodine 10% solution for cutaneous use, 1-litre bottle |  | 0 | 0 | 5 | 0 | 0 | 0
23 | 371 | Clamp, umbilical, 5.2cm, sterile, single use |  | 0 | 0 | 100 | 0 | 0 | 0
24 | 161 | Cannula, IV short, 20G, sterile, single use |  | 0 | 0 | 100 | 0 | 0 | 0
25 | 297 | Syringe, luer, 10 ml, sterile, single use |  | 0 | 0 | 500 | 0 | 0 | 0
26 | 301 | Syringe, luer, 2ml, sterile, single use |  | 0 | 0 | 400 | 0 | 0 | 0
27 | 299 | Syringe, luer, 1ml, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
28 | 226 | Needle, luer, 21G, sterile, single use |  | 0 | 0 | 1100 | 0 | 0 | 0
29 | 443 | Needle, luer, 25G, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
30 | 279 | Syringe, feeding, catheter tip, 50 ml, sterile, P1 |  | 0 | 0 | 10 | 0 | 0 | 0
31 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
32 | 214 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
33 | 209 | Gloves, gynaecological, medium, powder-free, sterile |  | 0 | 0 | 5 | 0 | 0 | 0
34 | 205 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 300 | 0 | 0 | 0
35 | 251 | Suture, absorbable, DEC3 (2-0), 3/8, 30mm, round, sterile |  | 0 | 0 | 36 | 0 | 0 | 0
36 | 199 | Extractor, mucus, 20ml, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
37 | 312 | Tube, suction, CH10, 50 cm long, conical tip, sterile |  | 0 | 0 | 10 | 0 | 0 | 0
38 | 313 | Tube, suction, CH14, 50 cm long, conical tip, sterile |  | 0 | 0 | 10 | 0 | 0 | 0
39 | 174 | Catheter, urethral, CH12, sterile |  | 0 | 0 | 20 | 0 | 0 | 0
40 | 187 | Gauze, compress, 10 x 10cm, sterile |  | 0 | 0 | 300 | 0 | 0 | 0
41 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 3 | 0 | 0 | 0
42 | 305 | Tape, adhesive, zinc oxide, 2.5cm x 5m |  | 0 | 0 | 5 | 0 | 0 | 0
43 | 571 | Soap, hand, bar, 110g, wrapped |  | 0 | 0 | 30 | 1 | 0 | 0
44 | 157 | Safety box, disposal of used syringes & needles, 5 litres |  | 0 | 0 | 11 | 0 | 0 | 0
45 | 343 | Glasses, safety, regular size, disposable |  | 0 | 0 | 2 | 0 | 0 | 0
46 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 1000 | 0 | 0 | 0
47 | 331 | Apron, protection, plastic, reusable |  | 0 | 0 | 2 | 0 | 0 | 0
48 | 340 | Draw sheet, plastic, 90 x 180cm, reusable |  | 0 | 0 | 2 | 0 | 0 | 0
49 | 545 | Test urinary protein, strip |  | 0 | 0 | 200 | 0 | 0 | 0
50 | 122 | Indicator, TST control, spot, adhesive |  | 0 | 0 | 300 | 0 | 0 | 0
51 | 155 | Bag, biohazard, yellow, 50 liters |  | 0 | 0 | 50 | 0 | 0 | 0
52 | 595 | Book, exercise, A4, ruled, 100 pages |  | 0 | 0 | 2 | 1 | 0 | 0
53 | 584 | Pen, ballpoint, blue |  | 0 | 0 | 10 | 1 | 0 | 0
54 | 565 | Bag, shoulder, UNFPA logo, vinyl, approx 360 x 230 x 610 mm |  | 0 | 0 | 1 | 1 | 0 | 0
55 | N/A | PRINTING MATERIALS 6TH ED KIT 6B |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit 7A - Intrauterine Devices (IUD)
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 555 | Intra-Uterine Device (IUD), Cu-T 380A |  | - | - | 90 | 1 | - | TA item
2 | 12 | Azithromycin 250mg, tablets |  | 0 | 0 | 240 | 0 | 0 | 0
3 | 23 | Cefixime 200mg, tablet |  | 0 | 0 | 170 | 0 | 0 | 0
4 | 86 | Povidone-iodine 10% solution for cutaneous use, 1 litre bottle |  | 0 | 0 | 6 | 0 | 0 | 0
5 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 100 | 0 | 0 | 0
6 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
7 | 214 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
8 | 187 | Gauze, compress, 10 x 10cm, sterile |  | 0 | 0 | 1000 | 0 | 0 | 0
9 | 497 | Speculum, vaginal, Graves, 95 x 35mm |  | 0 | 0 | 2 | 0 | 0 | 0
10 | 496 | Speculum, vaginal, Graves, 115 x 35mm |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 387 | Forceps, artery, Rochester-Pean, 22cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
13 | 316 | Sound, uterine, Martin, 32cm |  | 0 | 0 | 3 | 0 | 0 | 0
14 | 382 | Forceps, uterine, Duplay, 28cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
15 | 472 | Scissors, gynaecological, 20cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
16 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 1 | 0 | 0 | 0
17 | 319 | Basin, kidney, stainless steel, 825ml |  | 0 | 0 | 2 | 0 | 0 | 0
18 | 330 | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit 7B - Contraceptive Implant
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 87 | Povidone iodine 10% solution for cutaneous use, 500-ml bottle |  | 0 | 0 | 2 | 0 | 0 | 0
2 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 2 | 0 | 0 | 0
3 | 215 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
4 | 218 | Gloves, surgical, size 7.5, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
5 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
6 | 66 | Lidocaine hydrochloride 1% , 20ml ampoule USP or BP or equivalent |  | 0 | 0 | 60 | 0 | 0 | 0
7 | 289 | Syringe, luer, 5ml, with needle, 23G, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
8 | 239 | Adhesive bandage, wound plaster, waterproof |  | 0 | 0 | 50 | 0 | 0 | 0
9 | 244 | Gauze, 8cm x 4m, roll, non-sterile |  | 0 | 0 | 10 | 0 | 0 | 0
10 | 363 | Blade, scalpel, sterile, single use, no.10 |  | 0 | 0 | 100 | 0 | 0 | 0
11 | 305 | Tape, adhesive, zinc oxide, 2.5cm x 5m |  | 0 | 0 | 10 | 0 | 0 | 0
12 | 561 | Levonorgestrel 75mg×2 |  | - | - | 50 | - | - | TA item
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 8 - Management of complications of miscarriage or abortion
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 52 | Gentamicin sulphate 40mg base/ml for injection in 2ml ampoule |  | 0 | 0 | 100 | 0 | 0 | 0
2 | 6 | Ampicillin sodium 500mg, powder for injection in vial |  | 0 | 0 | 1500 | 0 | 0 | 0
3 | 80.2 | Misoprostol 200mcg, tablet |  | 0 | 0 | 180 | 0 | 0 | 0
4 | 39 | Disinfectant tablet for water containing 1.67g of sodium dichloroisocyanurate (NaDCC) |  | 0 | 0 | 400 | 0 | 0 | 0
5 | 62 | Ibuprofen 400mg tablet |  | 0 | 0 | 1000 | 0 | 0 | 0
6 | 66 | Lidocaine hydrochloride 1% , 20ml ampoule USP or BP or equivalent |  | 0 | 0 | 60 | 0 | 0 | 0
7 | 9 | Atropine sulphate 1mg/ml for injection in 1-ml ampoule |  | 0 | 0 | 30 | 0 | 0 | 0
8 | 98 | Water for injection, sterile, 10-ml ampoule |  | 0 | 0 | 500 | 0 | 0 | 0
9 | 31 | Chlorhexidine gluconate 4% solution (Hibiscrub), 500-ml bottle |  | 0 | 0 | 3 | 0 | 0 | 0
10 | 86 | Povidone-iodine 10% solution for cutaneous use, 1 litre bottle |  | 0 | 0 | 6 | 0 | 0 | 0
11 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
12 | 215 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
13 | 205 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 100 | 0 | 0 | 0
14 | 297 | Syringe, luer, 10 ml, sterile, single use |  | 0 | 0 | 1000 | 0 | 0 | 0
15 | 301 | Syringe, luer, 2ml, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
16 | 226 | Needle, luer, 21G, sterile, single use |  | 0 | 0 | 1300 | 0 | 0 | 0
17 | 187 | Gauze, compress, 10 x 10cm, sterile |  | 0 | 0 | 1200 | 0 | 0 | 0
18 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 100 | 0 | 0 | 0
19 | 157 | Safety box, disposal of used syringes & needles, 5 litres |  | 0 | 0 | 15 | 0 | 0 | 0
20 | 155 | Bag, Biohazard, yellow, 50 litres |  | 0 | 0 | 50 | 0 | 0 | 0
21 | 433 | Manual Vacuum Aspiration (MVA) with 2cc silicone |  | 0 | 0 | 4 | 0 | 0 | 0
22 | 422 | Accesory kit for MVA, including (1) O-ring and (1) collar stop |  | 0 | 0 | 2 | 0 | 0 | 0
23 | 434 | Silicone, 2ml |  | 0 | 0 | 60 | 0 | 0 | 0
24 | 427 | Cannulae, Easygrip, 6mm, integrated base, SR6 |  | 0 | 0 | 8 | 0 | 0 | 0
25 | 428 | Cannulae, Easygrip, 7mm, integrated base, SR7 |  | 0 | 0 | 8 | 0 | 0 | 0
26 | 429 | Cannulae, Easygrip, 8mm, integrated base, SR8 |  | 0 | 0 | 8 | 0 | 0 | 0
27 | 430 | Cannulae, Easygrip, 9mm, integrated base, SR9 |  | 0 | 0 | 4 | 0 | 0 | 0
28 | 423 | Cannulae, Easygrip, 10mm, integrated base, SR10 |  | 0 | 0 | 4 | 0 | 0 | 0
29 | 424 | Cannulae, Easygrip, 12mm, integrated base, SR12 |  | 0 | 0 | 4 | 0 | 0 | 0
30 | 431 | Dilators, Denniston, polymer |  | 0 | 0 | 20 | 0 | 0 | 0
31 | 124 | Basket, sterilizing, approx. 120 x 250 x 60 mm |  | 0 | 0 | 1 | 0 | 0 | 0
32 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 1 | 0 | 0 | 0
33 | 404 | Forceps, uterine, Museux, 24cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
34 | 393 | Forceps, uterine evacuation, ovum, Bierer, large |  | 0 | 0 | 1 | 0 | 0 | 0
35 | 392 | Forceps, uterine evacuation, ovum, Bierer, small |  | 0 | 0 | 1 | 0 | 0 | 0
36 | 452 | Retractor, vaginal, Doyen, 8.5 x 4.5cm |  | 0 | 0 | 1 | 0 | 0 | 0
37 | 489 | Speculum, vaginal, Auvard, retractor, weighed, 8 x 3.8cm |  | 0 | 0 | 1 | 0 | 0 | 0
38 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 1 | 0 | 0 | 0
39 | N/A | PRINTING MATERIALS 6TH ED KIT 8 |  |  |  |  | 1
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 9 - Repair of cervical and vaginal tears
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 86 | Povidone-iodine 10% solution for cutaneous use, 1 litre bottle |  | 0 | 0 | 15 | 0 | 0 | 0
2 | 31 | Chlorhexidine gluconate 4% solution (Hibiscrub), 500 ml bottle |  | 0 | 0 | 3 | 0 | 0 | 0
3 | 560 | Vaginal lubricant jelly/exploration gel, approx. 100-g tube |  | - | - | 1 | - | - | TA item
4 | 259 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile |  | 0 | 0 | 72 | 0 | 0 | 0
5 | 248 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile |  | 0 | 0 | 60 | 0 | 0 | 0
6 | 253 | Suture, absorbable, DEC3 (2-0), 3/8, 50mm, round, sterile |  | 0 | 0 | 72 | 0 | 0 | 0
7 | 187 | Gauze, compress, 10 x 10cm, sterile |  | 0 | 0 | 400 | 0 | 0 | 0
8 | 217 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 40 | 0 | 0 | 0
9 | 214 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 40 | 0 | 0 | 0
10 | 205 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 200 | 0 | 0 | 0
11 | 474 | Scissors, Mayo, 17cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
13 | 452 | Retractor, vaginal, Doyen, 8.5 x 4.5cm |  | 0 | 0 | 2 | 0 | 0 | 0
14 | 498 | Speculum, vaginal, Graves, 75 x 20mm |  | 0 | 0 | 1 | 0 | 0 | 0
15 | 497 | Speculum, vaginal, Graves, 95 x 35mm |  | 0 | 0 | 1 | 0 | 0 | 0
16 | 496 | Speculum, vaginal, Graves, 115 x 35mm |  | 0 | 0 | 1 | 0 | 0 | 0
17 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 2 | 0 | 0 | 0
18 | 330 | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 10 - Assisted delivery with vacuum extraction
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 502 | Vacuum extractor, Bird, anterior and posterior cups, manual, set |  | 0 | 0 | 1 | 0 | 0 | 0
2 | N/A | PRINTING MATERIALS 6TH ED KIT 10 |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 11A - Obstetric surgery and severe obstetric complications: Reusable Equipment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 125 | Basket, instruments, for sterilization, wired, 400 x 200 x 90 mm |  | 0 | 0 | 1 | 0 | 0 | 0
2 | 499 | Clamp, towel, Backhaus, 120 mm |  | 0 | 0 | 4 | 0 | 0 | 0
3 | 398 | Forceps, artery, Kelly, 14cm, curved |  | 0 | 0 | 10 | 0 | 0 | 0
4 | 403 | Forceps, artery, Halsted-Mosquito, 12.5cm, curved |  | 0 | 0 | 6 | 0 | 0 | 0
5 | 400 | Forceps, artery, Kocher, 14 cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
6 | 386 | Forceps, artery, Rochester-Pean, 20 cm, curved |  | 0 | 0 | 2 | 0 | 0 | 0
7 | 388 | Forceps, artery, Rochester-Pean, 24 cm, curved |  | 0 | 0 | 2 | 0 | 0 | 0
8 | 402 | Forceps, artery, Mixter, 23cm |  | 0 | 0 | 1 | 0 | 0 | 0
9 | 381 | Forceps, dressing, standard, 145 mm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 397 | Forceps, dressing, standard, 250 mm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 368 | Forceps, intestinal clamp, Doyen, 23cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 2 | 0 | 0 | 0
13 | 385 | Forceps, uterine, Phaneuf, 21.5 cm, curved |  | 0 | 0 | 2 | 0 | 0 | 0
14 | 382 | Forceps, uterine, Duplay, 28cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
15 | 417 | Forceps, tissue, Allis, 4x5 teeth, 15 cm |  | 0 | 0 | 2 | 0 | 0 | 0
16 | 413 | Forceps, tissue, Babcock, 20 cm |  | 0 | 0 | 1 | 0 | 0 | 0
17 | 420 | Forceps, tissue, Duval, 23cm |  | 0 | 0 | 2 | 0 | 0 | 0
18 | 411 | Forceps, tissue, standard, 145 mm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
19 | 415 | Forceps, tissue, standard, 250 mm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
20 | 322 | Bowl, stainless steel, 500 ml |  | 0 | 0 | 1 | 0 | 0 | 0
21 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
22 | 451 | Retractor, abdominal, Collin, 3 blades |  | 0 | 0 | 1 | 0 | 0 | 0
23 | 450 | Retractor, abdominal, Balfour, 3 blades |  | 0 | 0 | 1 | 0 | 0 | 0
24 | 448 | Retractor, double-ended, Farabeuf, 15 cm, pair |  | 0 | 0 | 1 | 0 | 0 | 0
25 | 461 | Scalpel handle no. 4, blade holder, 13 cm |  | 0 | 0 | 1 | 0 | 0 | 0
26 | 478 | Scissors, Metzenbaum/Nelson, 18cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
27 | 479 | Scissors, Metzenbaum/Nelson, 23cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
28 | 474 | Scissors, Mayo, 17cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
29 | 476 | Scissors, Mayo, 23cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
30 | 362 | Spatula, Ribbon retractor, malleable, 27 mm x 250 mm |  | 0 | 0 | 2 | 0 | 0 | 0
31 | 245 | Tube, suction, Yankauer, 28cm |  | 0 | 0 | 1 | 0 | 0 | 0
32 | 373 | Cranioclast, Braun, 420 mm |  | 0 | 0 | 1 | 0 | 0 | 0
33 | 445 | Perforator, Smellie, 25cm |  | 0 | 0 | 1 | 0 | 0 | 0
34 | 472 | Scissors, gynaecological, 20cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
35 | 378 | Hook, decapitation, Braun, 31cm |  | 0 | 0 | 1 | 0 | 0 | 0
36 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 1 | 0 | 0 | 0
37 | 112 | Pump, suction, foot-operated, aspirator |  | 0 | 0 | 1 | 0 | 0 | 0
38 | 107 | Resuscitator, hand-operating, child/newborn |  | 0 | 0 | 1 | 0 | 0 | 0
39 | 108 | Resuscitator, hand-operating, neonate, set |  | 0 | 0 | 1 | 0 | 0 | 0
40 | 109 | Resuscitator, hand-operated, adult, set |  | 0 | 0 | 1 | 0 | 0 | 0
41 | N/A | PRINTING MATERIALS 6TH ED KIT 11A |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 11B - Obstetric surgery and severe obstetric complications: Drugs and Disposable Equipment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 6 | Ampicillin sodium 500mg, powder for injection in vial |  | 0 | 0 | 400 | 0 | 0 | 0
2 | 33 | Clindamycin, injection, 600mg (as phosphate), 4ml ampoule |  | 0 | 0 | 300 | 0 | 0 | 0
3 | 1 | Amoxicillin, powder for oral suspension, 125mg/5ml, 100 ml bottle |  | 0 | 0 | 5 | 0 | 0 | 0
4 | 52 | Gentamicin sulphate 40mg base/ml for injection in 2ml ampoule |  | 0 | 0 | 150 | 0 | 0 | 0
5 | 36 | Cloxacillin, powder for solution (IV/IM), 500 mg vial |  | 0 | 0 | 90 | 0 | 0 | 0
6 | 85 | Paracetamol 500mg, tablet |  | 0 | 0 | 2000 | 0 | 0 | 0
7 | 93 | Tetracycline hydrochloride 1%, eye ointment, 5g tube - CANCELLED: TO BE REPLACED |  | 0 | 0 | 30 | 0 | 0 | 0
8 | 54 | Glucose 10%, injection, 10ml ampoule |  | 0 | 0 | 120 | 0 | 0 | 0
9 | 98 | Water for injection, sterile, 10-ml ampoule |  | 0 | 0 | 500 | 0 | 0 | 0
10 | 89 | Sodium Chloride 0.9%, injection, 10 ml ampoule |  | 0 | 0 | 100 | 0 | 0 | 0
11 | 84 | Oxytocin 10 IU/ml for injection in 1ml ampoule (keep cold: 2–8°C) |  | 0 | 0 | 200 | 0 | 0 | 0
12 | 80.1 | Misoprostol 200mcg, tablet |  | 0 | 0 | 60 | 0 | 0 | 0
13 | 94 | Tranexamic acid, injection, 100mg/ml, 10 ml ampoule |  | 0 | 0 | 10 | 0 | 0 | 0
14 | 21 | Calcium gluconate 100mg base/ml for injection in 10ml ampoule |  | 0 | 0 | 40 | 0 | 0 | 0
15 | 72 | Magnesium sulphate 500mg/ml injection in 10ml ampoule |  | 0 | 0 | 100 | 0 | 0 | 0
16 | 59 | Hydralazine hydrochloride 20mg/2ml powder for injection in 2ml ampoule |  | 0 | 0 | 60 | 0 | 0 | 0
17 | 66 | Lidocaine hydrochloride 1% , 20ml ampoule USP or BP or equivalent |  | 0 | 0 | 160 | 0 | 0 | 0
18 | 46 | Epinephrine (Adrenaline), injection, 1mg/ml, 1ml ampoule |  | 0 | 0 | 10 | 0 | 0 | 0
19 | 18 | Bupivacaine hydrochloride (as anhydrous) 0.5%, Intrathecal injection, 5g/min, 10 ml ampoule |  | 0 | 0 | 100 | 0 | 0 | 0
20 | 9 | Atropine sulphate 1mg/ml for injection in 1-ml ampoule |  | 0 | 0 | 20 | 0 | 0 | 0
21 | 45 | Ephedrine hydrochloride, 30mg/ml for injection in 1-ml ampoule |  | 0 | 0 | 150 | 0 | 0 | 0
22 | 63 | Ketamine hydrochloride 50mg base/ml for injection in 10-ml vial |  | 0 | 0 | 50 | 0 | 0 | 0
23 | 97 | Vitamin K-1, injection, 1mg/1ml, 1ml vial |  | 0 | 0 | 110 | 0 | 0 | 0
24 | 273 | Syringe, 0.5ml, permanently attached needle, sterile, single use |  | 0 | 0 | 110 | 0 | 0 | 0
25 | 55 | Glucose, solution for infusion, 5% (iso-osmotic), 1-litre bag + Infusion-giving set, sterile, single use |  | 0 | 0 | 320 | 0 | 0 | 0
26 | 57 | Sodium lactate (Ringer’s lactate), IV infusion, 1 litre+giving set |  | 0 | 0 | 210 | 0 | 0 | 0
27 | 39 | Disinfectant tablet for water containing 1.67g of sodium dichloroisocyanurate (NaDCC) |  | 0 | 0 | 400 | 0 | 0 | 0
28 | 31 | Chlorhexidine gluconate 4% solution (Hibiscrub), 500-ml bottle |  | 0 | 0 | 21 | 0 | 0 | 0
29 | 28 | Chlorhexidine digluconate 5% solution, 1 litre bottle |  | 0 | 0 | 30 | 0 | 0 | 0
30 | 86 | Povidone iodine 10% solution for cutaneous use, 1 litre bottle |  | 0 | 0 | 15 | 0 | 0 | 0
31 | 312 | Tube, suction, CH10, 50 cm long, conical tip, sterile |  | 0 | 0 | 60 | 0 | 0 | 0
32 | 313 | Tube, suction, CH14, 50 cm long, conical tip, sterile |  | 0 | 0 | 60 | 0 | 0 | 0
33 | 162 | Cannula, IV short, 20G, sterile, single use |  | 0 | 0 | 500 | 0 | 0 | 0
34 | 160 | Cannula, IV short, 18G, sterile, single use |  | 0 | 0 | 100 | 0 | 0 | 0
35 | 371 | Clamp, umbilical, 5.2cm, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
36 | 299 | Syringe, luer, 1ml, sterile, single use |  | 0 | 0 | 400 | 0 | 0 | 0
37 | 301 | Syringe, luer, 2ml, sterile, single use |  | 0 | 0 | 500 | 0 | 0 | 0
38 | 303 | Syringe, luer, 5ml, sterile, single use |  | 0 | 0 | 500 | 0 | 0 | 0
39 | 297 | Syringe, luer, 10 ml, sterile, single use |  | 0 | 0 | 900 | 0 | 0 | 0
40 | 293 | Syringe, luer lock, 20ml, sterile,single use |  | 0 | 0 | 100 | 0 | 0 | 0
41 | 226 | Needle, luer, 21G, sterile, single use |  | 0 | 0 | 3500 | 0 | 0 | 0
42 | 227 | Needle, luer, 23G, sterile, single use |  | 0 | 0 | 100 | 0 | 0 | 0
43 | 443 | Needle, luer, 25G, sterile, single use |  | 0 | 0 | 500 | 0 | 0 | 0
44 | 444 | Needle, scalp vein, butterfly, 25G, sterile, single use |  | 0 | 0 | 200 | 0 | 0 | 0
45 | 220 | Brush, hand, scrubbing, plastic |  | 0 | 0 | 10 | 0 | 0 | 0
46 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 400 | 0 | 0 | 0
47 | 215 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 400 | 0 | 0 | 0
48 | 209 | Gloves, gynaecological, medium, powder-free, sterile |  | 0 | 0 | 10 | 0 | 0 | 0
49 | 205 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 500 | 0 | 0 | 0
50 | 259 | Suture, absorbable, DEC4 (1), 3/8, 36mm, triangular, sterile |  | 0 | 0 | 288 | 0 | 0 | 0
51 | 248 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile |  | 0 | 0 | 288 | 0 | 0 | 0
52 | 271 | Suture, non-absorbable, DEC 3 (2-0), 3/8, triangular, 30mm, sterile |  | 0 | 0 | 144 | 0 | 0 | 0
53 | 172 | Catheter, urethral, Foley, CH14, sterile |  | 0 | 0 | 150 | 0 | 0 | 0
54 | 314 | Bag, urine, collecting, 2 litres |  | 0 | 0 | 150 | 0 | 0 | 0
55 | 337 | Drape, surgical woven, 100 x 150cm |  | 0 | 0 | 6 | 0 | 0 | 0
56 | 232 | Needle, spinal, 22G, sterile, single use |  | 0 | 0 | 120 | 0 | 0 | 0
57 | 187 | Gauze, compress, 10 x 10cm, sterile |  | 0 | 0 | 1000 | 0 | 0 | 0
58 | 305 | Tape, adhesive, zinc oxide, 2.5cm x 5m |  | 0 | 0 | 50 | 0 | 0 | 0
59 | 307 | Tape, adhesive, zinc oxide, perforated, 10cmx5m |  | 0 | 0 | 5 | 0 | 0 | 0
60 | 367 | Blade, scalpel, sterile, single use, no. 22 |  | 0 | 0 | 100 | 0 | 0 | 0
61 | 157 | Safety box, disposal of used syringes & needles, 5 litres |  | 0 | 0 | 35 | 0 | 0 | 0
62 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 1000 | 0 | 0 | 0
63 | 155 | Bag, Biohazard, yellow, 50 liters |  | 0 | 0 | 150 | 0 | 0 | 0
64 | 343 | Glasses, safety, regular size, disposable |  | 0 | 0 | 2 | 0 | 0 | 0
65 | 550 | Test pregnancy, strip, temperature stable |  | 0 | 0 | 100 | 0 | 0 | 0
66 | N/A | PRINTING MATERIALS 6TH ED KIT 11B |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Core IARH Kit |  | Kit 12 - Blood Transfusion
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 546 | Test blood-group, anti-A, 10-ml bottle, fl.dropper |  | 0 | 0 | 1 | 0 | 0 | 0
2 | 548 | Test blood-group, anti-A+B, 10-ml bottle, fl.dropper |  | 0 | 0 | 1 | 0 | 0 | 0
3 | 547 | Test blood-group, anti-B, 10-ml bottle, fl.dropper |  | 0 | 0 | 1 | 0 | 0 | 0
4 | 549 | Test Rhesus, anti-D, 10-ml bottle, dropper |  | 0 | 0 | 1 | 0 | 0 | 0
5 | 528 | Test HIV 1+2, rapid (Chase buffer included) |  | 0 | 0 | 100 | 0 | 0 | 0
6 | 533 | Test Hepatitis B surface antigen (HBsAg), rapid |  | 0 | 0 | 100 | 0 | 0 | 0
7 | 534 | Test Hepatitis C virus (HCV), rapid |  | 0 | 0 | 100 | 0 | 0 | 0
8 | 542 | Test Syphilis, rapid plasma regain (RPR) |  | 0 | 0 | 100 | 0 | 0 | 0
9 | 346 | Haemoglobin photometer handheld with accessories |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 348 | Capillary tubes for HemoCue Hb 301 |  | 0 | 0 | 200 | 0 | 0 | 0
11 | 240 | Pressure cuff, infusion/transfusion, for 500 to 1000 ml bags |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 566 | Battery, dry cell, AA, alkaline, 1.5 V, for photometer |  | 0 | 0 | 8 | 1 | 0 | 0
13 | 191 | Blood bag + CDPA-1 (citrate-phosphate-dextrose-adenine), 250ml, sterile |  | 0 | 0 | 50 | 0 | 0 | 0
14 | 193 | Blood bag + CDPA-1 (citrate-phosphate-dextrose-adenine), 350ml, sterile |  | 0 | 0 | 40 | 0 | 0 | 0
15 | 195 | Blood bag + CDPA-1 (citrate-phosphate-dextrose-adenine), 450ml, sterile |  | 0 | 0 | 40 | 0 | 0 | 0
16 | 156 | Blood-giving set |  | 0 | 0 | 100 | 0 | 0 | 0
17 | 204 | Gloves, examination, latex, medium, non-sterile |  | 0 | 0 | 500 | 0 | 0 | 0
18 | 344 | Tile, porcelain, with depressions for blood grouping |  | 0 | 0 | 1 | 0 | 0 | 0
19 | N/A | PRINTING MATERIALS 6TH ED KIT 12 |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Non-Pneumatic Anti-Shock Garment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 154 | Non-Pneumatic Anti-Shock Garment |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Oxytocin
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 84 | Oxytocin 10 I.U. base/ml injection in 1ml ampoule. (Store between 2-8 °C) |  | 0 | 0 | 300 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Mifepristone
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 79 | Mifepristone 200mg tablet |  | 0 | 0 | 70 | 0 | 0 | 0
2 | N/A | PRINTING MATERIALS 6TH ED KIT IARHKITMIFEP |  |  |  | 1 | - |  | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Hand-Held Vacuum Assisted Delivery System
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 504 | Vacuum extractor, Vacca OmniCup, hand-held, 1 person, reusable |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Chlorhexidine Di-gluconate
Item No | Bid No | UNFPA ITEM ID |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 29 | Chlorhexidine Di-gluconate 7.1%, gel, tube, 20g |  | 0 | 0 | 200 | 0 | 0 | 0
2 | N/A | PRINTING MATERIALS 6TH ED KIT IARHKITCHG |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Misoprostol, 300 tabs
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 80.2 | Misoprostol 200mcg, tablet |  | 0 | 0 | 300 | 0 | 0 | 0
2 | N/A | PRINTING MATERIALS 6TH ED KIT IARHKITMISO |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Comp. IARH Kit |  | CC Kit Medroxyprogesterone acetate (DMPA-SC)
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 559 | Medroxyprogesterone acetate 104mg/0.65ml, subcutaneous inj. |  | - | - | 200 | 1 |  | TA item
2 | N/A | PRINTING MATERIALS 6TH ED KIT IARHKITDMPASC |  |  |  | 1 |  |  | 0
[8 empty] | TOTAL | 0
Medical Kit |  | MVA kitting service
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per primary UOM) | Total Price
1 | 433 | MVA with 2cc silicone |  | 0 | 0 | 1 | 0 | 0 | 0
2 | 425 | Cannulae, Easygrip, 4mm, integrated base, SR4 |  | 0 | 0 | 1 | 0 | 0 | 0
3 | 426 | Cannulae, Easygrip, 5mm, integrated base, SR5 |  | 0 | 0 | 1 | 0 | 0 | 0
4 | 427 | Cannulae, Easygrip, 6mm, integrated base, SR6 |  | 0 | 0 | 1 | 0 | 0 | 0
5 | 428 | Cannulae, Easygrip, 7mm, integrated base, SR7 |  | 0 | 0 | 1 | 0 | 0 | 0
6 | 429 | Cannulae, Easygrip, 8mm, integrated base, SR8 |  | 0 | 0 | 1 | 0 | 0 | 0
7 | 430 | Cannulae, Easygrip, 9mm, integrated base, SR9 |  | 0 | 0 | 1 | 0 | 0 | 0
8 | 423 | Cannulae, Easygrip, 10mm, integrated base, SR10 |  | 0 | 0 | 1 | 0 | 0 | 0
9 | 424 | Cannulae, Easygrip, 12mm, integrated base, SR12 |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 431 | Dilators, Denniston, polymer |  | 0 | 0 | 5 | 0 | 0 | 0
11 | 422 | Accesory kit for MVA, including (1) O-ring, (1) collar stop, |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 434 | Silicone, 2cc, P10 |  | 0 | 0 | 10 | 0 | 0 | 0
*Note: Service to pack more than one item (1-4 specified above) into an individual bag as a kit. Requestors have the ability to choose which, and the amount of items to be included in one kit
Other Items
Item No | Bid No | Category | Item Short Description | Manufacturer | Country of Origin | Unit of Measure | Sales Pack Size | Price per Sales Pack
2 | 552 | EmergencyContraceptive | Levonorgestrel 1.5mg, tablet (treatment: single dose) | - | - | Tablet | 1 | TA item
3 | 553 | FemaleCondom | Condomize! condom nitrile sheath inner retention ring, 175mm | - | - | Each | 1 | TA item
4 | 554 | InjectableContraceptive | Medroxyprogesterone acetate 150mg/ml, depot injection in 1-ml vial | - | - | Vial | 1 | TA item
5 | 555 | IUD | Intra-Uterine Device (IUD), Cu-T 380A | - | - | Each | 1 | TA item
6 | 556 | MaleCondom | Male condom 53 mm, standard, natural | - | - | Each | 144 | TA item
7 | 557 | OralContraceptive | Ethinylestradiol, 0.03mg, plus levonorgestrel, 0.15mg, combined tablet, 1 strip for 1 cycle | - | - | Cycle | 3 | TA item
8 | 558 | OralContraceptive | Levonorgestrel 0.03mg, tablet, 1 strip for 1 cycle | - | - | Cycle | 3 | TA item
9 | 559 | Injectable Contraceptives | Medroxyprogesterone acetate 104mg/0.65ml,subcutaneous injection | - | - | Vial | 1 | TA item
10 | 560 | Personal Lubricants | Vaginal lubricant jelly/exploration gel, approx. 100-g tube | - | - | Tube | 1 | TA item
11 | 561 | Implantable Contraceptives | Levonorgestrel 75mg×2 | - | - | Rods | 1 | TA item
12 | 562 | Printing | Leaflet on Use of Female Condoms, Multilingual | - | - | Each | 1 | TA item
13 | 563 | Printing | Leaflet on Use of Male Condoms, English | - | - | Each | 1 | TA item
14 | 564 | Printing | Leaflet on Use of Male Condoms, French | - | - | Each | 1 | TA item
15 | 565 | Other | Bag, shoulder, UNFPA logo, vinyl, approx 360 x 230 x 610 mm |  |  | Each | 1
16 | 566 | Other | Battery, dry cell, AA, alkaline, 1.5 V, for Photometer |  |  | Each | 1
17 | 567 | Other | Flashlight, frontal, LED, battery-operated, including batteries |  |  | Each | 1
18 | 568 | Other | Flashlight, LED, re-chargeable |  |  | Set | 1
19 | 569 | Other | Kerosene storm-lamp + extra socks |  |  | Each | 1
20 | 570 | Other | Poncho, wet weather, reusable, heavy duty |  |  | Each | 1
21 | 571 | Other | Soap, hand, bar, 110g, wrapped |  |  | Each | 1
24 | 578 | Stationary | Binder with lever, 31 x 29 cm, various colours |  |  | Each | 1
25 | 579 | Stationary | Marker, erasable, red |  |  | Each | 1
26 | 580 | Stationary | Marker, flipchart, black, tip 4.5 mm |  |  | Each | 1
27 | 581 | Stationary | Marker, flipchart, red, tip 4.5 mm |  |  | Pair | 1
28 | 582 | Stationary | Paper, A4, 21 x 29.7cm, 80g, white |  |  | Each | 1
29 | 583 | Stationary | Pen, ballpoint, black |  |  | Each | 1
30 | 584 | Stationary | Pen, ballpoint, blue |  |  | Each | 1
31 | 585 | Stationary | Pen, ballpoint, red |  |  | Each | 1
32 | 586 | Stationary | Pencil sharpener, tapered |  |  | Each | 1
33 | 587 | Stationary | Pencil, HB |  |  | Each | 1
34 | 588 | Stationary | Scissors, blended, 170 mm |  |  | Each | 1
35 | 589 | Stationary | Board, white, adhesive, washable, roll, 67 x 100cm |  |  | Each | 1
36 | 590 | Stationary | Sticker, UNFPA, 11 x 22cm |  |  | Each | 1
37 | 591 | Stationary | Sticker, UNFPA, 6 x 12cm |  |  | Each | 1
38 | 592 | Stationary | Tape, adhesive, 19mm x 33m, clear, roll |  |  | Each | 1
39 | 593 | Stationary | Book, exercise, 17 x 22 cm, 5mm-squared, 100 pages |  |  | Each | 1
40 | 594 | Stationary | Book, exercise, 21 x 29.7 cm, 5mm-squared, 100 pages |  |  | Each | 1
41 | 595 | Stationary | Book, exercise, A4, ruled, 100 pages |  |  | Each | 1
42 | 596 | Stationary | Calculator, solar-powdered |  |  | Each | 1
43 | 597 | Stationary | Glue, stick, large |  |  | Each | 1
44 | 598 | Stationary | Marker, erasable, black |  |  | Each | 1
45 | 599 | Stationary | Marker, erasable, blue |  |  | Each | 1
Bidders Full Name:
Bidder's Warehouse full address for FCA deliveries:
Please indicate any cumulative volume percentage discount levels that increase as the cumulative value/volume increases throughout the validity of the LTA:
Please indicate any early payment percentage discount within a specified period of time faster than UNFPA's standard payment of net 30 days:
Name and Title
Signature
Date
Note on C500: 1) Protocols for medical abortion, Ipas, EN 2) Protocols for medical abortion, Ipas, FR 3) Protocols for medical abortion, Ipas, AR 4) Protocols for medical abortion, Ipas, SP 5) Mifepristone PLUS Misoprostol or Misoprostol Alone for Treatment of Intrauterine Fetal Death 12-24 weeks' LMP, Gynuity Health Projects, EN 6) Mifepristone PLUS Misoprostol or Misoprostol Alone for Treatment of Intrauterine Fetal Death 12-24 weeks' LMP, Gynuity Health Projects, FR 7) Mifepristone PLUS Misoprostol or Misoprostol Alone for Treatment of Intrauterine Fetal Death 12-24 weeks' LMP, Gynuity Health Projects, AR 8) Mifepristone PLUS Misoprostol or Misoprostol Alone for Treatment of Intrauterine Fetal Death 12-24 weeks' LMP, Gynuity Health Projects, SP ======

## 8. Price Other Medical Kits
Tab 8. Prices for the Other Medical Kits
Important guidance to Bidders:
1. Please note that the prices for pharmaceuticals and medical devices are automaticall transferred from Tab 1.
2. Manufacturer, Country of Origin and Price of "Other Items" to be added at the end of the tab in rows 220 to 228.
3. Weighted scoring will be used during the commercial and financial evaluation, where the weight score is given per each item based on the past procurement levels. The total score for each Bidder will be the sum of weighted scores for all quoted items, which will identify the lowest-priced bids that will be passed to further Technical Evaluation.
Item No | Module | Kit Name |  | FCA Price/Kit | Lead Time, weeks | Kitting Services, Cost/Kit | Warehousing Cost/Kit | Cost of reconfiguration/Kit | Total Price/Kit | Weighting | Weighted price
28 | Medical Kit | Midwifery kit – Equipment |  | 0 |  |  |  |  | 0 | 26% | 0
29 | Medical Kit | Midwifery kit – Renewable |  | 0 |  |  |  |  | 0 | 16% | 0
30 | Medical Kit | Fistula Repair Kit 1 - Surgical Instruments |  | 0 |  |  |  |  | 0 | 4% | 0
31 | Medical Kit | Fistula Repair Kit 2 - Supplementary Items |  | 0 |  |  |  |  | 0 | 17% | 0
32 | Medical Kit | Intrauterine Devices (IUD) Insertion and Removal kit |  | 0 |  |  |  |  | 0 | 12.47% | 0
33 | Medical Kit | Contraceptive Implant Insertion and Removal kit |  | 0 |  |  |  |  | 0 | 12.47% | 0
34 | Medical Kit | Minilaparotomy Kit |  | 0 |  |  |  |  | 0 | 12.47% | 0
ITEMIZED OVERVIEW OF KITS
Medical Kit |  | Midwifery kit – Equipment
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 331 | Apron, protection, plastic, reusable |  | 0 | 0 | 2 | 0 | 0 | 0
2 | 319 | Basin, kidney, stainless steel, 825ml |  | 0 | 0 | 2 | 0 | 0 | 0
3 | 595 | Book, exercise, A4, ruled, 100 pages |  | 0 | 0 | 20 | 1 | 0 | 0
4 | 323 | Bowl, polypropylene, 6L |  | 0 | 0 | 2 | 0 | 0 | 0
5 | 321 | Bowl, stainless steel, 4L |  | 0 | 0 | 2 | 0 | 0 | 0
6 | 322 | Bowl, stainless steel, 500 ml |  | 0 | 0 | 2 | 0 | 0 | 0
7 | 220 | Brush, hand, scrubbing, plastic, P1 |  | 0 | 0 | 2 | 0 | 0 | 0
8 | 340 | Draw sheet, plastic, 90 x 180cm, reusable |  | 0 | 0 | 2 | 0 | 0 | 0
9 | 400 | Forceps, artery, Kocher, 14cm, straight |  | 0 | 0 | 4 | 0 | 0 | 0
10 | 400 | Forceps, artery, Kocher, 14cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
11 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 2 | 0 | 0 | 0
12 | 411 | Forceps, tissue, standard, 145 mm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
13 | 324 | Jar, forceps, 180 mm, no cover |  | 0 | 0 | 2 | 0 | 0 | 0
14 | 325 | Jar, thermometer, plastic, 11cm, w/o cover |  | 0 | 0 | 2 | 0 | 0 | 0
15 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
16 | 584 | Pen, ballpoint, blue |  | 0 | 0 | 10 | 1 | 0 | 0
17 | 446 | Probe, surgical, double-ended, 14.5 cm |  | 0 | 0 | 2 | 0 | 0 | 0
18 | 355 | Scale, infant, spring-type, 5kg x 25g |  | 0 | 0 | 1 | 0 | 0 | 0
19 | 461 | Scalpel handle no. 4, blade holder, 13 cm |  | 0 | 0 | 2 | 0 | 0 | 0
20 | 465 | Scissors, Deaver, 14cm, curved, sharp/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
21 | 472 | Scissors, gynaecological, 20cm, curved, blunt/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
22 | 473 | Scissors, Mayo, 14cm, curved, blunt/blunt |  | 0 | 0 | 2 | 0 | 0 | 0
23 | 352 | Sling, for use with infant scale |  | 0 | 0 | 1 | 0 | 0 | 0
24 | 356 | Sphygmomanometer, aneroid, adult |  | 0 | 0 | 1 | 0 | 0 | 0
25 | 357 | Stethoscope, binaural, complete |  | 0 | 0 | 1 | 0 | 0 | 0
26 | 358 | Stethoscope, obstetrical, Pinard, monoaural |  | 0 | 0 | 2 | 0 | 0 | 0
27 | 572 | Tape-measure, vinyl-coated, 1.5 m/5 feet |  | 0 | 0 | 2 | 1 | 0 | 0
28 | 359 | Thermometer, clinical, digital 32°C-43°C |  | 0 | 0 | 10 | 0 | 0 | 0
29 | 311 | Tourniquet, latex rubber, 75cm |  | 0 | 0 | 2 | 0 | 0 | 0
30 | 573 | Towel, huck, 430 x 500 mm |  | 0 | 0 | 5 | 1 | 0 | 0
31 | 328 | Tray, dressing, stainless steel, 30 x 20 x 3cm |  | 0 | 0 | 2 | 0 | 0 | 0
32 | 329 | Tray, instruments, stainless steel, 22.5 x 12.5 x 5cm, with cover |  | 0 | 0 | 4 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Midwifery kit – Renewable
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 196 | Bag (envelope), plastic, for drugs, 10 x 15cm |  | 0 | 0 | 500 | 0 | 0 | 0
2 | 367 | Blade, scalpel, sterile, single use, no. 22, P100 |  | 0 | 0 | 100 | 0 | 0 | 0
3 | 334 | Blanket, survival, 220 x 140cm |  | 0 | 0 | 5 | 0 | 0 | 0
4 | 162 | Cannula, IV short, 20G, sterile, single use, P50 |  | 0 | 0 | 50 | 0 | 0 | 0
5 | 174 | Catheter, urethral, CH12, sterile, P1 |  | 0 | 0 | 50 | 0 | 0 | 0
6 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 5 | 0 | 0 | 0
7 | 199 | Extractor, mucus, 20ml, sterile, single use, P1 |  | 0 | 0 | 50 | 0 | 0 | 0
8 | 184 | Gauze, compress, 10 x 10cm, non-sterile, P100 |  | 0 | 0 | 500 | 0 | 0 | 0
9 | 187 | Gauze, compress, 10 x 10cm, sterile, P5 |  | 0 | 0 | 600 | 0 | 0 | 0
10 | 205 | Gloves, exam, latex, medium, non-sterile, P100 |  | 0 | 0 | 200 | 0 | 0 | 0
11 | 209 | Gloves, gynaecological, medium, powder-free, sterile, 1 pair |  | 0 | 0 | 25 | 0 | 0 | 0
12 | 214 | Gloves, surgical, size 7, powder-free, sterile, single use, 1 pair |  | 0 | 0 | 25 | 0 | 0 | 0
13 | 217 | Gloves, surgical, size 8, powder-free, sterile, single use, 1 pair |  | 0 | 0 | 25 | 0 | 0 | 0
14 | 225 | Needle, luer, 19G, sterile, P100 |  | 0 | 0 | 100 | 0 | 0 | 0
15 | 226 | Needle, luer, 21G, sterile, P100 |  | 0 | 0 | 100 | 0 | 0 | 0
16 | 229 | Needle, scalp vein, 21G, sterile, single use, P50 |  | 0 | 0 | 50 | 0 | 0 | 0
17 | 158 | Safety box, disposal of used syringes & needles, 5L, P25 |  | 0 | 0 | 25 | 0 | 0 | 0
18 | 571 | Soap, hand, bar, 110g, wrapped |  | 0 | 0 | 36 | 1 | 0 | 0
19 | 247 | Suture, absorb., DEC 2 (3-0), 3/8, 26mm, triangular, sterile, P36 |  | 0 | 0 | 36 | 0 | 0 | 0
20 | 260 | Suture, absorb., DEC 4 (1), 3/8, 50mm, round, sterile, P36 |  | 0 | 0 | 36 | 0 | 0 | 0
21 | 279 | Syringe, feeding, catheter tip, 50 ml, sterile, P1 |  | 0 | 0 | 10 | 0 | 0 | 0
22 | 297 | Syringe, luer, 10 ml, w/o needle, sterile, single use, P100 |  | 0 | 0 | 100 | 0 | 0 | 0
23 | 301 | Syringe, luer, 2ml, w/o needle, sterile, single use, P100 |  | 0 | 0 | 100 | 0 | 0 | 0
24 | 305 | Tape, adhesive, zinc oxide, 2.5cm x 5m, P1 |  | 0 | 0 | 12 | 0 | 0 | 0
25 | 309 | Tape, umbilical, 3mmx50m, non-sterile, single use, P1 |  | 0 | 0 | 2 | 0 | 0 | 0
26 | 312 | Tube, suction, CH10, 50 cm long, conical tip, sterile, P1 |  | 0 | 0 | 50 | 0 | 0 | 0
27 | 313 | Tube, suction, CH14, 50 cm long, conical tip, sterile, P1 |  | 0 | 0 | 25 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Fistula Repair Kit 1 - Surgical Instruments
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 575 | Leaflet, Fistula Kits, English |  | - | - | 1 | 1 | 0 | 0
2 | 576 | Leaflet, Fistula Kits, French |  | - | - | 1 | 1 | 0 | 0
3 | 317 | Basin, kidney, stainless steel, 1400ml |  | 0 | 0 | 2 | 0 | 0 | 0
4 | 490 | Speculum, vaginal, Auvard, retractor, weighed, 10 x 4 cm |  | 0 | 0 | 1 | 0 | 0 | 0
5 | 492 | Speculum, vaginal, Sims, retractor, double-ended, medium |  | 0 | 0 | 2 | 0 | 0 | 0
6 | 491 | Speculum, vaginal, Sims, retractor, double-ended, large |  | 0 | 0 | 1 | 0 | 0 | 0
7 | 487 | Scissors, Thorek, 18cm, curved 90 degrees |  | 0 | 0 | 1 | 0 | 0 | 0
8 | 466 | Scissors, fistula, 20cm curved |  | 0 | 0 | 1 | 0 | 0 | 0
9 | 470 | Scissors, Boyd, 17cm, double curved |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 480 | Scissors, Metzenbaum/Nelson, 24 cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 442 | Needle holder, Mayo-Hegar, 20cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
13 | 457 | Scalpel handle no.7, blade holder, 15 cm |  | 0 | 0 | 1 | 0 | 0 | 0
14 | 461 | Scalpel handle no.4, blade holder, 13 cm |  | 0 | 0 | 1 | 0 | 0 | 0
15 | 414 | Forceps,tissue,standard,dissecting,1x2 teeth,20cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
16 | 396 | Forceps, dressing, standard, fine serrated jaws, 20cm |  | 0 | 0 | 1 | 0 | 0 | 0
17 | 482 | Scissors, suture stich, 18cm, straight, sharp/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
18 | 484 | Scissors, Thorek, 18cm, curved, sharp/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
19 | 447 | Probe, surgical, with eye, malleable, 20 cm |  | 0 | 0 | 1 | 0 | 0 | 0
20 | 501 | Sound, uterine, Martin, malleable, 30 cm |  | 0 | 0 | 1 | 0 | 0 | 0
21 | 174 | Catheter, urethral, CH12, sterile, P1 |  | 0 | 0 | 1 | 0 | 0 | 0
22 | 449 | Retractor, Langenbeck, single-ended, medium, 21cm |  | 0 | 0 | 2 | 0 | 0 | 0
23 | 404 | Forceps, uterine, Museux, 24cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
24 | 435 | Needle, aneurysm, Deschamps, 20cm, sharp, left hand |  | 0 | 0 | 1 | 0 | 0 | 0
25 | 436 | Needle, aneurysm, Deschamps, 20cm, sharp, right hand |  | 0 | 0 | 1 | 0 | 0 | 0
26 | 402 | Forceps, artery, Mixter, 23cm |  | 0 | 0 | 2 | 0 | 0 | 0
27 | 410 | Forceps, tissue, Allis, 3x4 teeth, 20cm |  | 0 | 0 | 4 | 0 | 0 | 0
28 | 418 | Forceps, tissue, Allis, 3x4 teeth, 15cm |  | 0 | 0 | 2 | 0 | 0 | 0
29 | 454 | Safety Pin, forceps holder, Mayo, 11.5cm |  | 0 | 0 | 2 | 0 | 0 | 0
30 | 455 | Safety Pin, forceps holder, Schaedel, 9cm |  | 0 | 0 | 6 | 0 | 0 | 0
31 | 408 | Forceps, sponge-holding, Foerster, 24 cm, straight |  | 0 | 0 | 2 | 0 | 0 | 0
32 | 403 | Forceps, artery, Halsted-Mosquito, 12.5cm, curved |  | 0 | 0 | 10 | 0 | 0 | 0
33 | 405 | Forceps, artery, Spencer-Wells, 20.5 cm, curved |  | 0 | 0 | 4 | 0 | 0 | 0
34 | 500 | Clamp, towel, Backhaus, 90 mm |  | 0 | 0 | 4 | 0 | 0 | 0
35 | 499 | Clamp, towel, Backhaus, 120 mm |  | 0 | 0 | 6 | 0 | 0 | 0
36 | 379 | Dilator, uterine, Hegar, set of 16 dilators, size 3-18 |  | 0 | 0 | 1 | 0 | 0 | 0
37 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 2 | 0 | 0 | 0
38 | 327 | Ruler, metal, in centimeters, 30cm |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Fistula Repair Kit 2 - Supplementary Items
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 575 | Leaflet, Fistula Kits, English |  | - | - | 1 | 1 | 0 | 0
2 | 576 | Leaflet, Fistula Kits, French |  | - | - | 1 | 1 | 0 | 0
3 | 178 | Catheter, urethral, CH5, with metal guide wire, sterile |  | 0 | 0 | 6 | 0 | 0 | 0
4 | 179 | Catheter, urethral, size CH6, with metal guide wire, sterile |  | 0 | 0 | 6 | 0 | 0 | 0
5 | 315 | Bag, urine, collecting, 2 litres, tap for emptying |  | 0 | 0 | 25 | 0 | 0 | 0
6 | 172 | Catheter, urethral, Foley, CH14, sterile, P1 |  | 0 | 0 | 1 | 0 | 0 | 0
7 | 180 | Catheter, urethral, Foley, CH16, sterile, P1 |  | 0 | 0 | 10 | 0 | 0 | 0
8 | 173 | Catheter, urethral, Foley, adult, CH18, sterile, P1 |  | 0 | 0 | 20 | 0 | 0 | 0
9 | 181 | Catheter, urethral, Foley, CH20, sterile, P1 |  | 0 | 0 | 5 | 0 | 0 | 0
10 | 364 | Blade, scalpel, sterile, single use, no. 11, P1 |  | 0 | 0 | 30 | 0 | 0 | 0
11 | 366 | Blade, scalpel, sterile, single use, no. 15, P1 |  | 0 | 0 | 5 | 0 | 0 | 0
12 | 295 | Bladder syringe 60-100 mls (with long nozzle, not luer lock) |  | 0 | 0 | 25 | 0 | 0 | 0
13 | 233 | Needle, spinal, 22G, sterile, single use, P25 |  | 0 | 0 | 25 | 0 | 0 | 0
14 | 224 | Needle, spinal, 25G, sterile, single use |  | 0 | 0 | 25 | 0 | 0 | 0
15 | 183 | Colostomy bag, transparent, drainable, w/filter, P20 |  | 0 | 0 | 20 | 0 | 0 | 0
16 | 264 | Suture, absorbable, DEC 3.5 (0), with needle, fistula repair |  | 0 | 0 | 36 | 0 | 0 | 0
17 | 263 | Suture, absorbable, DEC 3 (2/0), with needle, fistula repair |  | 0 | 0 | 36 | 0 | 0 | 0
18 | 262 | Suture, absorbable, DEC 2 (3/0), with needle, fistula repair |  | 0 | 0 | 36 | 0 | 0 | 0
19 | 261 | Suture, absorbable, DEC 1.5 (4/0), with needle, fistula repair |  | 0 | 0 | 12 | 0 | 0 | 0
20 | 265 | Suture, absorbable, DEC 4 (1), with needle, fistula repair |  | 0 | 0 | 36 | 0 | 0 | 0
21 | 269 | Suture, non-absorbable PA, DEC 4 (1), 1/2 circle, round |  | 0 | 0 | 36 | 0 | 0 | 0
22 | 268 | Suture, non-absorbable PA, DEC 3 (2/0), 1/2 circle, round |  | 0 | 0 | 36 | 0 | 0 | 0
23 | 249 | Suture, absorbable, DEC 3 (2-0), 1/2, 30mm, round, sterile, P36 |  | 0 | 0 | 36 | 0 | 0 | 0
24 | 235 | Surgical needle, 1/2 circle, round, spring eye, size14, sterile |  | 0 | 0 | 12 | 0 | 0 | 0
25 | 237 | Needle, surgical, 1/2, 63mm, round, spring eye, P24 |  | 0 | 0 | 48 | 0 | 0 | 0
26 | 236 | Needle, surgical, 1/2, 27mm, round, spring eye, P24 |  | 0 | 0 | 48 | 0 | 0 | 0
27 | 83 | Oxybutynin hydrochloride, 5mg, tablet |  | 0 | 0 | 1000 | 0 | 0 | 0
28 | 20 | Bupivacaine hydrochloride 0.5% + glucose or dextrose 7.5%, 4ml, ampoule, injection for spinal anaesthesia (heavy or hyperbaric solution) |  | 0 | 0 | 20 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Intrauterine Devices (IUD) Insertion and Removal kit
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 496 | Speculum, vaginal, Graves, 115 x 35mm |  | 0 | 0 | 1 | 0 | 0 | 0
2 | 380 | Forceps, dressing, Cheron, 25cm |  | 0 | 0 | 2 | 0 | 0 | 0
3 | 387 | Forceps, artery, Rochester-Pean, 22cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
4 | 316 | Sound, uterine, Martin, 32cm |  | 0 | 0 | 1 | 0 | 0 | 0
5 | 382 | Forceps, uterine, Duplay, 28cm, curved |  | 0 | 0 | 1 | 0 | 0 | 0
6 | 472 | Scissors, gynaecological, 20cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
7 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 1 | 0 | 0 | 0
8 | 319 | Basin, kidney, stainless steel, 825ml |  | 0 | 0 | 1 | 0 | 0 | 0
9 | 330 | Tray, instruments, stainless steel, 32 x 20 x 8cm, with cover |  | 0 | 0 | 1 | 0 | 0 | 0
10 | 476 | Scissors, Mayo, 23cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 600 | Implant card |  | - | - | 1 | 1 | 0 | 0
12 | 601 | Counselling material |  | - | - | 1 | 1 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Minilaparotomy Kit
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 320 | Bowl, stainless steel, 180ml |  | 0 | 0 | 2 | 0 | 0 | 0
2 | 460 | Scalpel handle no. 3, blade holder, 13 cm |  | 0 | 0 | 2 | 0 | 0 | 0
3 | 505 | Forceps, Dressing, Standard Pattern, 12.7 cm |  | 0 | 0 | 1 | 0 | 0 | 0
4 | 506 | Forceps, Intestinal, Allis, Delicate, 15 cm (5x6 teeth) |  | 0 | 0 | 2 | 0 | 0 | 0
5 | 474 | Scissors, Mayo, 17cm, curved, blunt/blunt |  | 0 | 0 | 1 | 0 | 0 | 0
6 | 507 | Retractor, double-ended, Farabeuf, 120 mm, pair |  | 0 | 0 | 2 | 0 | 0 | 0
7 | 508 | Army-Navy retractors, double-ended (for subumbilical procedure) |  | 0 | 0 | 2 | 0 | 0 | 0
8 | 509 | Forceps, Artery, Kelly, Straight, 14 cm |  | 0 | 0 | 2 | 0 | 0 | 0
9 | 510 | Forceps mosquito, delicate, curved, 14 cm |  | 0 | 0 | 2 | 0 | 0 | 0
10 | 511 | Scissors, Tonsil, Metzenbaum, Curved, 18 cm |  | 0 | 0 | 1 | 0 | 0 | 0
11 | 411 | Forceps, tissue, standard, 145mm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
12 | 413 | Forceps, tissue, Babcock, 20 cm |  | 0 | 0 | 2 | 0 | 0 | 0
13 | 408 | Forceps, sponge-holding, Foerster, 24 cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
14 | 512 | Ramathobodi tubal hook |  | 0 | 0 | 1 | 0 | 0 | 0
15 | 440 | Needle holder, Mayo-Hegar, 18cm, straight |  | 0 | 0 | 1 | 0 | 0 | 0
16 | 497 | Speculum, vaginal, Graves, 95 x 35mm |  | 0 | 0 | 1 | 0 | 0 | 0
17 | 515 | Forceps, sponge, Foerster, straight, 24 cm |  | 0 | 0 | 1 | 0 | 0 | 0
18 | 514 | Forceps, Schroeder-Braun uterine tenaculum, 24.8 cm |  | 0 | 0 | 1 | 0 | 0 | 0
19 | 513 | Ramathobodi uterine elevator, 28 cm |  | 0 | 0 | 1 | 0 | 0 | 0
20 | 516 | Jackson vaginal retractor, 38mm X 76mm (deep blade) (for alternative method for inserting the uterine elevator) |  | 0 | 0 | 1 | 0 | 0 | 0
[8 empty] | TOTAL | 0
Medical Kit |  | Contraceptive Implant Insertion and Removal kit
Item No | Bid No | Item Short Description |  | Manufacturer | Country of Origin | Qty (as per UOM) | Sales Pack Size | Price per Unit (as per UOM) | Total Price
1 | 87 | Povidone iodine 10% solution for cutaneous use, 500-ml bottle |  | 0 | 0 | 2 | 0 | 0 | 0
2 | 188 | Cotton wool, 500g, roll, non-sterile |  | 0 | 0 | 2 | 0 | 0 | 0
3 | 214 | Gloves, surgical, size 7, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
4 | 218 | Gloves, surgical, size 7,5, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
5 | 216 | Gloves, surgical, size 8, powder-free, sterile, single use |  | 0 | 0 | 50 | 0 | 0 | 0
6 | 66 | Lidocaine hydrochloride 1%, 20ml ampoule USP or BP or equivalent |  | 0 | 0 | 60 | 0 | 0 | 0
7 | 289 | Syringe, luer, 5ml, with needle,23G,sterile,single use |  | 0 | 0 | 50 | 0 | 0 | 0
8 | 239 | Adhesive bandage, wound plaster, waterproof |  | 0 | 0 | 50 | 0 | 0 | 0
9 | 241 | Gauze, 8cm x 4m, roll, non-sterile |  | 0 | 0 | 10 | 0 | 0 | 0
10 | 363 | Blade, scalpel, sterile, single use, no.10 |  | 0 | 0 | 100 | 0 | 0 | 0
11 | 305 | Tape, adhesive, zinc oxide, 2.5cm x 5m |  | 0 | 0 | 10 | 0 | 0 | 0
12 | 600 | Implant card |  | - | - | 25 | 1 | 0 | 0
13 | 601 | Counselling material |  | - | - | 25 | 1 | 0 | 0
[8 empty] | TOTAL | 0
Other Items
Item No | Bid No | UNFPA ITEM ID | Item Short Description | Manufacturer | Country of Origin | Unit of Measure | Sales Pack Size | Price per Sales Pack Size
1 | 571 | Other | Soap, hand, bar, 110g, wrapped |  |  | Each | 1
2 | 572 | Other | Tape-measure, vinyl-coated, 1.5 m/5 feet |  |  | Each | 1
3 | 573 | Other | Towel, huck, 430 x 500 mm |  |  | Each | 1
4 | 575 | Printing | Leaflet, Fistula Kits, English | - | - | Each | 1
5 | 576 | Printing | Leaflet, Fistula Kits, French | - | - | Each | 1
6 | 584 | Stationary | Pen, ballpoint, blue |  |  | Each | 1
7 | 595 | Stationary | Book, exercise, A4, ruled, 100 pages |  |  | Each | 1
8 | 600 | Printing | Implant Card | - | - | Each | 1
9 | 601 | Printing | Counseling material | - | - | Each | 1
Bidders Full Name:
Bidder's Warehouse full address for FCA deliveries:
Please indicate any cumulative volume percentage discount levels that increase as the cumulative value/volume increases throughout the validity of the LTA:
Please indicate any early payment percentage discount within a specified period of time faster than UNFPA's standard payment of net 30 days:
Name and Title
Signature
Date

## 9. Price DataLogger & ColdBoxes
Tab 9. Prices for Data Loggers and Cold Boxes
TEMPERATURE MONITORING DEVICES
Model | WHO Reference | Manufacturer |  | Price in USD per Unit | Bidder's Comment
Data Logger Model 1 |  |  |  | 0
Data Logger Model 2 (Alternative) |  |  |  | 0
KEEP COLD BOXES FOR OXYTOCIN SHIPMENT BY AIR
Model | Duration (in hours) | Quantity of ampoules per Box | Price in USD per Unit | Additional price per pack size offered | Bidder's Comment
Keep Cold Box 1 |  |  | 0 | 0
Keep Cold Box 2 (Alternative) |  |  | 0 | 0
UkraineKyiv