## Lot 1. Blood centrifuge
UNOPS-Ukraine-UPTF-2023-G-003
Amendment 1
Amendment 2
Brand
Model
Lot | Description | Technical specification | Value | Is Bid Compliant? Bidder to complete (Y / N) | Document and page where the technical specification or condition can be reviewed | Details of the offered goods. Bidder to complete
1 | Blood bank centrifuge: Blood bank refrigerated centrifuge specially designed for the separation of blood components in blood bank units. Quantity: 1 unit. UNOPS reserves the right to increment quantity up to 2 units depending on available budget. | 1. General requirements | 1.1 Floor standing centrifuge.
1.2 Microprocessor speed control.
1.3 Compatible with 250 ml, 450 ml, and 500 ml blood bags.
1.4 Capacity: at least 12 blood bags of 250-500 ml.
1.5 Maximum relative centrifugal force (with installed rotor for 12 blood bags): no less than 6.400 x g.
1.6 Maximum rotation rate (with installed rotor for 12 blood bags): no less than 4400 rpm
1.7 Adjustable speed.
1.8 Adjustable temperature range: -10 to 40 °C.
1.9 Adjustable timer range: up to 99h 59m.
1.10 Digital display for operational parameters and system messages. The equipment shall display, at least: temperature (°C), time, speed, acceleration and deceleration.
1.11 Option to program at least 10 acceleration/deceleration steps.
1.12 Program memory for at least 100 89 profiles.
|  | 2. Mechanical requirements | 2.1 Adjustable feet.
2.2 Noise level ≤ 65dB.
2.3 HCFC and CFC-free refrigeration compressor.
|  | 3. Safety requirements | 3.1 Electronic lid lock with emergency opening feature, visual and audio alarms.
3.2 Shut-off in case of imbalance.
3.3 At least the following visual/audible alarms:
- Imbalance.
- Over-speed.
- Over-temperature.
- Lid sensor fault.
- Error messages.
|  | 4. Electrical requirements | 4.1 Power supply: 210~240 V 50 Hz, single phase, plug compatible with Ukrainian standards. Power supply: 220 ±10% V 50 Hz, single phase, plug C type or 380V±10% 50 Hz, three-phase
4.2 Maximum power with aerodynamic fairing (cowling): no more than 6 kVA.
|  | 5. Accessories included in the offered price | 5.1 Swing-out rotor with aerodynamic fairing (cowling) for at least 12 250-500 ml bags.
5.2 Blood bag buckets for at least 12 250-500 ml blood bags.
5.3 Necessary adapters for items 5.1. and 5.2. to centrifuge 500mL blood bags.
5.4 Necessary adapters for items 5.1. and 5.2. to centrifuge 450mL blood bags.
5.5 Necessary adapters for items 5.1. and 5.2. to centrifuge 250mL blood bags
| Documents, certificates and standards required, to be presented by the bidder with the offer: | 6. Regulatory approval and marketing authorization in Ukraine, considering the Letter of the MOH No. 24-04/5396/2-22 dated 25 February 2022 and Resolution of the Cabinet of Ministers of Ukraine "On the List of Products of Critical Imports" No. 153 dated 24 February 2022, in accordance with at least one of the following: | 6.1 Valid and current product registration issued by the State Service of Ukraine on Medicines and Drug Control (SMDC) in conformance with the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable.
6.2 Compliance declaration and certification or documents proving the possibility of using the medical products following the procedure to assess compliance with the requirements of the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable and their updates.
6.3 A copy of the relevant statement by the Ministry of Healthcare on introducing certain medical products into circulation and/or use, for which the requirements of technical regulations have not been met, but whose use is justified from a healthcare perspective, for the batch of items in question.
|  | 7. Regulatory approval and marketing authorization issued by a Notified Body in at least one of the GHTF Founding Member countries, in accordance to the following list: | 7.1 European Union: Regulation (EU) 2017/745 on medical devices or Regulation (EU) 2017/746 on in-vitro diagnostic medical devices, as applicable
7.2 United States: FDA authorization
7.3 Canada: SOR/98-282
7.4 Australia: TGA conformity certification
7.5 Japan: PMDA pre-market approval
|  | 8. Certificate of Quality Management System | 8.1 According to ISO 13485 in the name of the Manufacturer for the medical device offered. The Certificate must be issued by an independent Certifying Body/Agency.
|  | 9. Letter of Authorization from Manufacturer | 9.1 Issued by the Manufacturer of the equipment authorizing the bidder to sell/distribute the offered equipment, or Authorization from the Official Distributor. Please consider that the authorization text must also indicate that the authorized Distributor is jointly committed with the bidder to provide the required post-sales services, guarantees, and training requested. Furthermore, authorization must be accompanied by the document that certifies that the Distributor has official authorization from the Manufacturer to sell/distribute the offered equipment.
|  | 10. User manual and technical brochure | 10.1 In Ukrainian and English
|  | 11. Guarantee conditions affidavit | 11.1 Affidavit of the bidder accepting the guarantee conditions established in the bidding documents and committing to deliver together with the equipment a guarantee certificate in the name of the beneficiary.
|  | 12. Spare parts affidavit | 12.1 Affidavit from the bidder guaranteeing the availability of spare parts and accessories for at least 5 years in case of award.
|  | 13. Maintenance affidavit | 13.1 Affidavit from the manufacturer stating the address in Ukraine of the company that will conduct maintenance of the equipment
|  | 14. Warranty, preventive and corrective maintenance affidavit | 14.1 Affidavit from the company that will carry out the maintenance (be it the supplier, manufacturer or a third party) committing to execute the guarantee and the preventive and corrective maintenance in accordance with the conditions indicated in the bidding documents
|  | 15. Personnel affidavit | 15.1 Copy of the training certificates of the engineer(s) that will conduct the installation and maintenance or affidavit for their training from the manufacturer
| Post Sale Services | 16. Delivery, installation, testing and commissioning: | 16.1 On-site delivery and installation to the recipient institution is included. Installation shall include connection to the Laboratory Information System when available.
16.2 The supplier shall transport the equipment inside the recipient institution to the installation site/room, open the packages, assemble and install it according to the manufacturer’s recommendations. Any damage to the equipment or accessories that occurred during transportation to the facility will be covered by insurance at the supplier's expense.
16.3 The supplier shall perform on-site testing, calibration and commissioning of the equipment with certification of conformity to standards and manufacturer’s recommendation. After installation, testing and calibration completion, the equipment is operational and ready to use. All equipment must be functioning with no physical damage and/or defect, prior to the acceptance by the end-user.
16.4 The supplier shall deliver any original software license in the name of the recipient hospital, when applicable, together with the equipment.
16.5 A complete set of commissioning forms with the entire set of tests performed and the results obtained shall be given to the recipient institution and to UNOPS once installation and commissioning is completed.
16.6 The Beneficiary will assist the personnel of the Supplier during their travels and on-site permanence to improve their safety although the Supplier shall be aware of the safety risks of the current situation and of its possible evolution on the field.
|  | 17. Training: | 17.1 The supplier must provide training on the use/operation of the equipment and all its accessories for at least one (1) to the end-users. The training shall include at least: 1) Operation instructions in normal functioning, 2) Daily and periodic user-related cleaning and maintenance, 3) Errors and malfunctions, 4) Main risks in the use of the equipment.
17.2 The supplier must provide training on the preventive and corrective maintenance of the equipment and all its accessories for at least one (1) calendar day to the maintenance staff.
17.3 The location of the training course shall be the location where the equipment is delivered and installed.
|  | 18. Warranty | 18.1 Twentyfour (24) months of manufacturer warranty. The warranty certificate will be in the name of the final Beneficiary.
18.2 Twentyfour (24) months of supplier warranty extending the Manufacturer warranty in scope not in duration. The scope will include: All maintenance shall be free of charge for the recipient institution throughout the warranty period. The warranty certificate will be in the name of the final Beneficiary. At least 95% of one single year of full functioning, i.e. 347 days out of 365 days, will be guaranteed by the Supplier within the warranty period. In the event that the item supplied has been malfunctioning for more than five percent (5%) of one single year of the warranty period, i.e. more than 18 natural days in one single year, the Supplier shall extend the warranty period for a duration of six times of the time duration when the equipment was malfunctioning. For this the supplier will make available on-line remote support for the maintenance of the supplied goods. This support shall be with a person trained by the manufacturer or with the manufacturer and available at least during business days. In case that with the on-line remote suport the issue cannot be solved the Supplier will send the trained technician on-site at the Beneficiary premise, being any cost of equipment transport or technician traveling at Supplier/Installer charge and included in the offered price. The time elapsed between the communication about the broken equipment and the intervention on site will be, inside the warranty period, not longer than 5 business days including in these 5 days all the remote support.
18.3 Twentyfour 24 months extended warranty shall include preventive maintenance and calibration of the equipment according to the manufacturers’ recommendations. The schedule of the visit will be presented to the Beneficiary before the final acceptance of the equipment. During any on-site maintenance and calibration visits a short user training update will be carried out by the Supplier.
18.4 The warranty period shall start from the date of acceptance of the goods by the end-user after training, testing and commissioning.
18.5 The equipment shall be serviced by an engineer properly certified by the equipment manufacturer (a copy of the service engineer’s certificate shall be provided in the offer).
18.6 The company providing the service shall have a presence in Ukraine (a guarantee letter from the equipment manufacturer shall be provided in the offer to this end, stating the service’s address in Ukraine).
18.7 After corrective and preventive maintenance visits, a report describing the technical procedures and results performed on the equipment shall be provided to the recipient institution and proper calibration shall be performed.
18.8 The bidder shall have the responsibilities and accountabilities that upon evaluation of equipment defects/faults, the equipment are appropriately repaired or replaced and shall be in good working condition thereafter.
|  | 19. Manuals | 19.1 The supplier must provide the end-users an electronic copy of the User manual in the Ukrainian language and the Service manual in the Ukrainian or English language for the equipment and accessories.

## Lot 2. Whole blood processing p
UNOPS-Ukraine-UPTF-2023-G-003
Brand
Model
Lot | Description | Technical specification | Value | Is Bid Compliant? Bidder to complete (Y / N) | Document and page where the technical specification or condition can be review | Details of the offered goods. Bidder to complete
2 | Whole Blood Processing Platform: System that integrates and automatizes the functions of centrifugation, sedimentation, separation, and sealing of whole blood units, resulting in multiple blood components. Quantity: 1 unit UNOPS reserves the right to increment quantity up to 2 units depending on available budget. | 1. General requirements | 1.1 Automates the process for obtaining RBC, leukocyte, platelet and plasma units from a single whole blood unit of 450mL.
1.2 Capacity for processing at least 1 blood unit in a single run.
1.3 Throughput – The system controls the time required to complete the procedure
1.4 Uses one to four disposable sets to collect and process whole blood doses.
1.5 Touch-screen user interface.
1.6 Manual data entry and via barcode.
1.7 Ethernet connectivity via RJ45 port.
1.8 Preconfigured and customizable processes.
1.9 Recording and traceability of operational and production data. It should include, at least: number of blood units processed, duration of the process, number of components produced (with segregation by type), volume of each component.
1.10 Includes software license to transfer process data to the Laboratory Information System.
1.11 Generates automatic reports with process and output information that can be exported through Ethernet port or via USB
1.12 Noise level < 65dB.
1.13 With four wheels, at least two with breaks.
|  | 2. Safety requirements | 2.1. Stop process emergency button.
2.2. Fluid leak detection sensor.
2.3. Automatic tube sealing after procedure is complete
2.4. Minimum audio/visual alerts:
2.4.1. Data loss.
2.4.2. System temperature too high/low.
2.4.3. Bag mismatch.
2.4.4. Leak detected.
2.4.5. Hydraulic pressure too high/low.
2.4.6. Centrifugation speed out of range.
2.4.7. High vibration.
2.4.8. Plasma and RBC detection problem.
2.4.9. Line not detected.
|  | 3. Electrical requirements | 3.1 Power supply: 220 ±10% V 50 Hz, single phase, plug compatible with Ukrainian standards.
|  | 4. Accessories included in the offered price | 4.1 barcode scanner.
4.2 counterbalance bags with appropriate buckets or adapters if needed.
4.3. Any accessory required for the normal operation of the device.
|  | 5. Consumables | 5.1. 6000 blood bag sets that include: RBC bag, platelet bag, plasma bag, residual leukocyte bag and sample bag.
5.2. 500 Platelet Concentrate set with sample bag.
| Documents, certificates and standards required, to be presented by the bidder with the offer: | 6. Regulatory approval and marketing authorization in Ukraine, considering the Letter of the MOH No. 24-04/5396/2-22 dated 25 February 2022 and Resolution of the Cabinet of Ministers of Ukraine "On the List of Products of Critical Imports" No. 153 dated 24 February 2022, in accordance with at least one of the following: | 6.1 Valid and current product registration issued by the State Service of Ukraine on Medicines and Drug Control (SMDC) in conformance with the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable.
6.2 Compliance declaration and certification or documents proving the possibility of using the medical products following the procedure to assess compliance with the requirements of the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable and their updates.
6.3 A copy of the relevant statement by the Ministry of Healthcare on introducing certain medical products into circulation and/or use, for which the requirements of technical regulations have not been met, but whose use is justified from a healthcare perspective, for the batch of items in question.
|  | 7. Regulatory approval and marketing authorization issued by a Notified Body in at least one of the GHTF Founding Member countries, in accordance to the following list: | 7.1 European Union: Regulation (EU) 2017/745 on medical devices or Regulation (EU) 2017/746 on in-vitro diagnostic medical devices, as applicable
7.2 United States: FDA authorization
7.3 Canada: SOR/98-282
7.4 Australia: TGA conformity certification
7.5 Japan: PMDA pre-market approval
|  | 8. Certificate of Quality Management System | 8.1 According to ISO 13485 in the name of the Manufacturer for the medical device offered. The Certificate must be issued by an independent Certifying Body/Agency.
|  | 9. Letter of Authorization from Manufacturer | 9.1 Issued by the Manufacturer of the equipment authorizing the bidder to sell/distribute the offered equipment, or Authorization from the Official Distributor. Please consider that the authorization text must also indicate that the authorized Distributor is jointly committed with the bidder to provide the required post-sales services, guarantees, and training requested. Furthermore, authorization must be accompanied by the document that certifies that the Distributor has official authorization from the Manufacturer to sell/distribute the offered equipment.
|  | 10. User manual and technical brochure | 10.1 In Ukrainian and English
|  | 11. Guarantee conditions affidavit | 11.1 Affidavit of the bidder accepting the guarantee conditions established in the bidding documents and committing to deliver together with the equipment a guarantee certificate in the name of the beneficiary.
|  | 12. Spare parts affidavit | 12.1 Affidavit from the bidder guaranteeing the availability of spare parts and accessories for at least 5 years in case of award.
|  | 13. Maintenance affidavit | 13.1 Affidavit from the manufacturer stating the address in Ukraine of the company that will conduct maintenance of the equipment
|  | 14. Warranty, preventive and corrective maintenance affidavit | 14.1 Affidavit from the company that will carry out the maintenance (be it the supplier, manufacturer or a third party) committing to execute the guarantee and the preventive and corrective maintenance in accordance with the conditions indicated in the bidding documents
|  | 15. Personnel affidavit | 15.1 Copy of the training certificates of the engineer(s) that will conduct the installation and maintenance or affidavit for their training from the manufacturer
| Post Sale Services | 16. Delivery, installation, testing and commissioning: | 16.1 On-site delivery and installation to the recipient institution is included. Installation shall include connection to the Laboratory Information System when available.
16.2 The supplier shall transport the equipment inside the recipient institution to the installation site/room, open the packages, assemble and install it according to the manufacturer’s recommendations. Any damage to the equipment or accessories that occurred during transportation to the facility will be covered by insurance at the supplier's expense.
16.3 The supplier shall perform on-site testing, calibration and commissioning of the equipment with certification of conformity to standards and manufacturer’s recommendation. After installation, testing and calibration completion, the equipment is operational and ready to use. All equipment must be functioning with no physical damage and/or defect, prior to the acceptance by the end-user.
16.4 The supplier shall deliver any original software license in the name of the recipient hospital, when applicable, together with the equipment.
16.5 A complete set of commissioning forms with the entire set of tests performed and the results obtained shall be given to the recipient institution and to UNOPS once installation and commissioning is completed.
16.6 The Beneficiary will assist the personnel of the Supplier during their travels and on-site permanence to improve their safety although the Supplier shall be aware of the safety risks of the current situation and of its possible evolution on the field.
|  | 17. Training: | 17.1 The supplier must provide training on the use/operation of the equipment and all its accessories for at least one (1) to the end-users. The training shall include at least: 1) Operation instructions in normal functioning, 2) Daily and periodic user-related cleaning and maintenance, 3) Errors and malfunctions, 4) Main risks in the use of the equipment.
17.2 The supplier must provide training on the preventive and corrective maintenance of the equipment and all its accessories for at least one (1) calendar day to the maintenance staff.
17.3 The location of the training course shall be the location where the equipment is delivered and installed.
|  | 18. Warranty | 18.1 Twentyfour (24) months of manufacturer warranty. The warranty certificate will be in the name of the final Beneficiary.
18.2 Twentyfour (24) months of supplier warranty extending the Manufacturer warranty in scope not in duration. All maintenance shall be free of charge for the recipient institution throughout the warranty period. The warranty certificate will be in the name of the final Beneficiary. At least 95% of one single year of full functioning, i.e. 347 days out of 365 days, will be guaranteed by the Supplier within the warranty period. In the event that the item supplied has been malfunctioning for more than five percent (5%) of one single year of the warranty period, i.e. more than 18 natural days in one single year, the Supplier shall extend the warranty period for a duration of six times of the time duration when the equipment was malfunctioning. For this the supplier will make available on-line remote support for the maintenance of the supplied goods. This support shall be with a person trained by the manufacturer or with the manufacturer and available at least during business days. In case that with the on-line remote suport the issue cannot be solved the Supplier will send the trained technician on-site at the Beneficiary premise, being any cost of equipment transport or technician traveling at Supplier/Installer charge and included in the offered price. The time elapsed between the communication about the broken equipment and the intervention on site will be, inside the warranty period, not longer than 5 business days including in these 5 days all the remote support.
18.3 Twentyfour 24 months extended warranty shall include preventive maintenance and calibration of the equipment according to the manufacturers’ recommendations. The schedule of the visit will be presented to the Beneficiary before the final acceptance of the equipment. During any on-site maintenance and calibration visits a short user training update will be carried out by the Supplier.
18.4 The warranty period shall start from the date of acceptance of the goods by the end-user after training, testing and commissioning.
18.5 The equipment shall be serviced by an engineer properly certified by the equipment manufacturer (a copy of the service engineer’s certificate shall be provided in the offer).
18.6 The company providing the service shall have a presence in Ukraine (a guarantee letter from the equipment manufacturer shall be provided in the offer to this end, stating the service’s address in Ukraine).
18.7 After corrective and preventive maintenance visits, a report describing the technical procedures and results performed on the equipment shall be provided to the recipient institution and proper calibration shall be performed.
18.8 The bidder shall have the responsibilities and accountabilities that upon evaluation of equipment defects/faults, the equipment are appropriately repaired or replaced and shall be in good working condition thereafter.
|  | 19. Manuals | 19.1 The supplier must provide the end-users an electronic copy of the User manual in the Ukrainian language and the Service manual in the Ukrainian or English language for the equipment and accessories.

## Lot 3. Automated microplates EL
UNOPS-Ukraine-UPTF-2023-G-003
Brand
Model
Lot | Description | Technical specification | Value | Is Bid Compliant? Bidder to complete (Y / N) | Document and page where the technical specification or condition can be review | Details of the offered goods. Bidder to complete
3 | Automated 2 Microplates ELISA Analyzer: Photometric microplate analyzer used for enzyme linked immunoadsorbent assay (ELISA). | 1. General requirements | 1.1 Fully automated microplate analysis equipment that contains a pipetting station, a spectrophotometer, an incubation station with shaker, and washing station.
1.2 The equipment must be open system compatible with different brands of reagents, without exclusive use of reagents produced by a single supplier.
1.3 Capacity: at least two 96-well microplates.
1.4 Light source: halogen, LED or xenon.
1.5 Visual or audible alarms for at least: lamp power supply fault, filter settings fault, wash extent fault, low liquid level, waste container full, and system failure.
1.6 Data management:
1.6.1 External (computer based)
1.6.2 Connectivity:
1.6.2.1 Bidirectional ASTM protocol.
1.6.2.2 Possibility to connect the analyzer with the existing Laboratory Information System (LIS) of the beneficiary.
1.7 Software:
1.7.1 User interface in English or Ukrainian language.
1.7.2 Automatic management (control) of all system elements.
1.7.3 Management of various subsystem interfaces.
1.7.4 Calibration and automatic system self-diagnostics.
1.7.5 Interpretation, presentation, and printing of results.
1.7.6 Customizable operational protocols.
1.7.7 Fully automatic protocol management.
1.7.8 Quantitative outputs, with at least the following interpolation methods: Single point, point-to-point, quadratic, linear regression, cubic spline Lin/Log and Log/Log.
1.7.9 Printing protocols, work sessions, and test results.
1.8 Sample and Dispensing unit:
1.8.1 Disposable tips automatically managed by software
1.8.2 Pipetting area: handling at least 2 microplates.
1.8.3 Precision (sample and reagent): CV<6.0% with 10μl, CV<2.2% with 100μl.
1.8.4 Pneumatic level sensor.
1.8.5 Possibility to perform preliminary dilution.
1.8.6 Sample dispensing time: < 25 min/96 well (100 μl volume).
1.8.7 Reagent dispensing time: < 5 min/96 well (100 μl volume).
1.8.8 No carryover (disposable tips for reagents and samples).
1.8.9 At least two barcode scanning devices allowing to automatically identify samples, reagents, and control/standard substances.
1.8.10 Tube size handling: range of at least 12 - 13.5 mm diameter, 55 - 75 mm height.
1.8.10 Capacity for at least 96 sample tubes.
1.9 Incubation module
1.9.1 At least two independent chambers.
1.9.2 Operating range of at least 35 to 45ºC.
1.9.3 Stability at ± 1°C.
1.9.4 Integrated plate shaker.
1.10 Washing:
1.10.1 At least 8 aspiration channels and 8 dosing channels.
1.10.2 Washing volume of at least 50 to 2000 µl/well.
1.10.3 Soak time range of at least 1 to 120 s.
1.10.4 Automated wash cycles, managed per assay.
1.10.5 Option to program at least 400 washing cycles.
1.10.6 Waste tank level sensor.
1.11 Photometric reading:
1.11.1 At least 16 (2 x 8) channels.
1.11.2 Photometric method: vertical type with single, double, or triple beam with out-of range beam filter.
1.11.3 Wavelength from 405 nm to 700 nm.
1.11.4 8 interference filters, at least five provided: approx 405, 450, 492, 550 and 620 nm.
1.11.5 Measurement range: 0 to 3 Optical Density (OD) or higher.
1.11.6 Linearity (absorbance): ±2.0 % or better.
1.11.7 Accuracy: ± 1 % (0 to 2 OU) or ± 10 mOU (mili optical unit) or better.
1.11.8 Resolution: 0.001 OU (optical unit) or better.
|  | 2. Electrical requirements | 2.1 Power supply: 220 ±10% V 50 Hz, single phase, plug compatible with Ukrainian standards.
|  | 3. Consumables | 3.1. 96000 units of 200 µL disposable pipette tips compatible with the offered equipment (as required to run 12000 tests).
3.2. 12000 units of 1000 µL disposable pipette tips compatible with the offered equipment (as required to run 12000 tests).
3.3. Any other consumables required for the normal use of the device in the indicated productivity.
| Documents, certificates and standards required, to be presented by the bidder with the offer: | 4. Regulatory approval and marketing authorization in Ukraine, considering the Letter of the MOH No. 24-04/5396/2-22 dated 25 February 2022 and Resolution of the Cabinet of Ministers of Ukraine "On the List of Products of Critical Imports" No. 153 dated 24 February 2022, in accordance with at least one of the following: | 4.1 Valid and current product registration issued by the State Service of Ukraine on Medicines and Drug Control (SMDC) in conformance with the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable.
4.2 Compliance declaration and certification or documents proving the possibility of using the medical products following the procedure to assess compliance with the requirements of the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable and their updates.
4.3 A copy of the relevant statement by the Ministry of Healthcare on introducing certain medical products into circulation and/or use, for which the requirements of technical regulations have not been met, but whose use is justified from a healthcare perspective, for the batch of items in question.
|  | 5. Regulatory approval and marketing authorization issued by a Notified Body in at least one of the GHTF Founding Member countries, in accordance to the following list: | 5.1 European Union: Regulation (EU) 2017/745 on medical devices or Regulation (EU) 2017/746 on in-vitro diagnostic medical devices, as applicable
5.2 United States: FDA authorization
5.3 Canada: SOR/98-282
5.4 Australia: TGA conformity certification
5.5 Japan: PMDA pre-market approval
|  | 6. Certificate of Quality Management System | 6.1 According to ISO 13485 in the name of the Manufacturer for the medical device offered. The Certificate must be issued by an independent Certifying Body/Agency.
|  | 7. Letter of Authorization from Manufacturer | 7.1 Issued by the Manufacturer of the equipment authorizing the bidder to sell/distribute the offered equipment, or Authorization from the Official Distributor. Please consider that the authorization text must also indicate that the authorized Distributor is jointly committed with the bidder to provide the required post-sales services, guarantees, and training requested. Furthermore, authorization must be accompanied by the document that certifies that the Distributor has official authorization from the Manufacturer to sell/distribute the offered equipment.
|  | 8. User manual and technical brochure | 8.1 In Ukrainian and English
|  | 9. Guarantee conditions affidavit | 9.1 Affidavit of the bidder accepting the guarantee conditions established in the bidding documents and committing to deliver together with the equipment a guarantee certificate in the name of the beneficiary.
|  | 10. Spare parts affidavit | 10.1 Affidavit from the bidder guaranteeing the availability of spare parts and accessories for at least 5 years in case of award.
|  | 11. Maintenance affidavit | 11.1 Affidavit from the manufacturer stating the address in Ukraine of the company that will conduct maintenance of the equipment
|  | 12. Warranty, preventive and corrective maintenance affidavit | 12.1 Affidavit from the company that will carry out the maintenance (be it the supplier, manufacturer or a third party) committing to execute the guarantee and the preventive and corrective maintenance in accordance with the conditions indicated in the bidding documents
|  | 13. Personnel affidavit | 13.1 Copy of the training certificates of the engineer(s) that will conduct the installation and maintenance or affidavit for their training from the manufacturer
| Post Sale Services | 14. Delivery, installation, testing and commissioning: | 14.1 On-site delivery and installation to the recipient institution is included. Installation shall include connection to the Laboratory Information System when available.
15.2 The supplier shall transport the equipment inside the recipient institution to the installation site/room, open the packages, assemble and install it according to the manufacturer’s recommendations. Any damage to the equipment or accessories that occurred during transportation to the facility will be covered by insurance at the supplier's expense.
14.3 The supplier shall perform on-site testing, calibration and commissioning of the equipment with certification of conformity to standards and manufacturer’s recommendation. After installation, testing and calibration completion, the equipment is operational and ready to use. All equipment must be functioning with no physical damage and/or defect, prior to the acceptance by the end-user.
14.4 The supplier shall deliver any original software license in the name of the recipient hospital, when applicable, together with the equipment.
14.5 A complete set of commissioning forms with the entire set of tests performed and the results obtained shall be given to the recipient institution and to UNOPS once installation and commissioning is completed.
14.6 The Beneficiary will assist the personnel of the Supplier during their travels and on-site permanence to improve their safety although the Supplier shall be aware of the safety risks of the current situation and of its possible evolution on the field.
|  | 15. Training: | 15.1 The supplier must provide training on the use/operation of the equipment and all its accessories for at least one (1) to the end-users. The training shall include at least: 1) Operation instructions in normal functioning, 2) Daily and periodic user-related cleaning and maintenance, 3) Errors and malfunctions, 4) Main risks in the use of the equipment.
15.2 The supplier must provide training on the preventive and corrective maintenance of the equipment and all its accessories for at least one (1) calendar day to the maintenance staff.
15.3 The location of the training course shall be the location where the equipment is delivered and installed.
|  | 16. Warranty | 16.1 Twentyfour (24) months of manufacturer warranty. The warranty certificate will be in the name of the final Beneficiary.
16.2 Twentyfour (24) months of supplier warranty extending the Manufacturer warranty in scope not in duration. All maintenance shall be free of charge for the recipient institution throughout the warranty period. The warranty certificate will be in the name of the final Beneficiary. At least 95% of one single year of full functioning, i.e. 347 days out of 365 days, will be guaranteed by the Supplier within the warranty period. In the event that the item supplied has been malfunctioning for more than five percent (5%) of one single year of the warranty period, i.e. more than 18 natural days in one single year, the Supplier shall extend the warranty period for a duration of six times of the time duration when the equipment was malfunctioning. For this the supplier will make available on-line remote support for the maintenance of the supplied goods. This support shall be with a person trained by the manufacturer or with the manufacturer and available at least during business days. In case that with the on-line remote suport the issue cannot be solved the Supplier will send the trained technician on-site at the Beneficiary premise, being any cost of equipment transport or technician traveling at Supplier/Installer charge and included in the offered price. The time elapsed between the communication about the broken equipment and the intervention on site will be, inside the warranty period, not longer than 5 business days including in these 5 days all the remote support.
16.3 Twentyfour 24 months extended warranty shall include preventive maintenance and calibration of the equipment according to the manufacturers’ recommendations. The schedule of the visit will be presented to the Beneficiary before the final acceptance of the equipment. During any on-site maintenance and calibration visits a short user training update will be carried out by the Supplier.
16.4 The warranty period shall start from the date of acceptance of the goods by the end-user after training, testing and commissioning.
16.5 The equipment shall be serviced by an engineer properly certified by the equipment manufacturer (a copy of the service engineer’s certificate shall be provided in the offer).
16.6 The company providing the service shall have a presence in Ukraine (a guarantee letter from the equipment manufacturer shall be provided in the offer to this end, stating the service’s address in Ukraine).
16.7 After corrective and preventive maintenance visits, a report describing the technical procedures and results performed on the equipment shall be provided to the recipient institution and proper calibration shall be performed.
16.8 The bidder shall have the responsibilities and accountabilities that upon evaluation of equipment defects/faults, the equipment are appropriately repaired or replaced and shall be in good working condition thereafter.
|  | 17. Manuals | 17.1 The supplier must provide the end-users an electronic copy of the User manual in the Ukrainian language and the Service manual in the Ukrainian or English language for the equipment and accessories.

## Lot 4. X-ray blood irradiator
UNOPS-Ukraine-UPTF-2023-G-003
Brand
Model
Lot | Description | Technical specification | Value | Is Bid Compliant? Bidder to complete (Y / N) | Document and page where the technical specification or condition can be review | Details of the offered goods. Bidder to complete
4 | X-Ray irradiation device designed to irradiate human blood units to prevent transfusion-associated graft-versus-host disease (GvHD). Quantity: 1 unit. UNOPS reserves the right to increment quantity up to 2 units depending on available budget. | 1. General requirements | 1.1 Application: human blood products.
1.2 Capacity of at least 3 blood bags of 450 ml or 4 blood bags of 300 ml.
1.3 Independent equipment with self-shielding design.
1.4 Rotating canister configuration that guarantees dose homogeneity.
1.5 Dose uniformity ratio not higher than 1.5.
1.6 Central dose rate of at least 6 Gy/min (± 5%).
1.7 Time setting for central dose rate of 5 minutes or better.
1.8 Cycles per hour for 25 Gy Central: at least 10.
1.9 Minimum cycle settings: time, speed, dose (kV, mA).
1.10 Voltage settings available: 80 kV – 200 kV, 20 kV step
1.11 Amperage setting available: 5 mA – 20 mA, 5 mA step
1.12 Water cooling for X-ray generator: piped water, water flow speed no less than 5 liters/min.
1.13 Built-in touchscreen, at least 14 inches.
1.14. USB ports.
1.15. Internal memory for at least 800 extractions.
1.16. Data transfer capabilities:
1.16.1. LAN or WiFi.
1.16.2. At least 1 USB port.
1.17. Minimum 15 programs preconfigured by manufacturer.
1.18. Audiovisual alarms/messages: process failure, incorrect tube placement, incorrect bag placement on scales, press closing/opening error, press calibration error, flow regulation error, the program completed.
1.19. Scales and optoelectronic sensor reset and automatic calibration with every procedure.
|  | 2. Safety requirements | 2.1. Safety lock to prevent door opening during radiation.
2.2. Equipment management protection – two operator sign-in procedures: 1 - password; 2 - personal smart card.
|  | 3. Accessories included in the offered price | 3.1. Set of mapping phantoms that allow equipment dosimetry and calibration.
3.2. Set of 2 additional canisters.
3.3. Barcode reader and built-in label printer.
|  | 4. Electrical requirements | 4.1 Power supply: 210~240 V 50 Hz, single phase, plug compatible with Ukrainian standards. Power supply: 220 ±10% V 50 Hz, single phase, plug C type or 380V±10% 50 Hz, three-phase
| Documents, certificates and standards required, to be presented by the bidder with the offer: | 5. Regulatory approval and marketing authorization in Ukraine, considering the Letter of the MOH No. 24-04/5396/2-22 dated 25 February 2022 and Resolution of the Cabinet of Ministers of Ukraine "On the List of Products of Critical Imports" No. 153 dated 24 February 2022, in accordance with at least one of the following: | 5.1 Valid and current product registration issued by the State Service of Ukraine on Medicines and Drug Control (SMDC) in conformance with the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable.
5.2 Compliance declaration and certification or documents proving the possibility of using the medical products following the procedure to assess compliance with the requirements of the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable and their updates.
5.3 A copy of the relevant statement by the Ministry of Healthcare on introducing certain medical products into circulation and/or use, for which the requirements of technical regulations have not been met, but whose use is justified from a healthcare perspective, for the batch of items in question.
|  | 6. Regulatory approval and marketing authorization issued by a Notified Body in at least one of the GHTF Founding Member countries, in accordance to the following list: | 6.1 European Union: Regulation (EU) 2017/745 on medical devices or Regulation (EU) 2017/746 on in-vitro diagnostic medical devices, as applicable
6.2 United States: FDA authorization
6.3 Canada: SOR/98-282
6.4 Australia: TGA conformity certification
6.5 Japan: PMDA pre-market approval
|  | 7. Certificate of Quality Management System | 7.1 According to ISO 13485 in the name of the Manufacturer for the medical device offered. The Certificate must be issued by an independent Certifying Body/Agency.
|  | 8. Letter of Authorization from Manufacturer | 8.1 Issued by the Manufacturer of the equipment authorizing the bidder to sell/distribute the offered equipment, or Authorization from the Official Distributor. Please consider that the authorization text must also indicate that the authorized Distributor is jointly committed with the bidder to provide the required post-sales services, guarantees, and training requested. Furthermore, authorization must be accompanied by the document that certifies that the Distributor has official authorization from the Manufacturer to sell/distribute the offered equipment.
|  | 9. User manual and technical brochure | 9.1 In Ukrainian and English
|  | 10. Guarantee conditions affidavit | 10.1 Affidavit of the bidder accepting the guarantee conditions established in the bidding documents and committing to deliver together with the equipment a guarantee certificate in the name of the beneficiary.
|  | 11. Spare parts affidavit | 11.1 Affidavit from the bidder guaranteeing the availability of spare parts and accessories for at least 5 years in case of award.
|  | 12. Maintenance affidavit | 12.1 Affidavit from the manufacturer stating the address in Ukraine of the company that will conduct maintenance of the equipment
|  | 13. Warranty, preventive and corrective maintenance affidavit | 13.1 Affidavit from the company that will carry out the maintenance (be it the supplier, manufacturer or a third party) committing to execute the guarantee and the preventive and corrective maintenance in accordance with the conditions indicated in the bidding documents
|  | 14. Personnel affidavit | 14.1 Copy of the training certificates of the engineer(s) that will conduct the installation and maintenance or affidavit for their training from the manufacturer
| Post Sale Services | 15. Delivery, installation, testing and commissioning: | 15.1 On-site delivery and installation to the recipient institution is included. Installation shall include connection to the Laboratory Information System when available.
15.2 The supplier shall transport the equipment inside the recipient institution to the installation site/room, open the packages, assemble and install it according to the manufacturer’s recommendations. Any damage to the equipment or accessories that occurred during transportation to the facility will be covered by insurance at the supplier's expense.
15.3 The supplier shall perform on-site testing, calibration and commissioning of the equipment with certification of conformity to standards and manufacturer’s recommendation. After installation, testing and calibration completion, the equipment is operational and ready to use. All equipment must be functioning with no physical damage and/or defect, prior to the acceptance by the end-user.
15.4 The supplier shall deliver any original software license in the name of the recipient hospital, when applicable, together with the equipment.
15.5 A complete set of commissioning forms with the entire set of tests performed and the results obtained shall be given to the recipient institution and to UNOPS once installation and commissioning is completed.
15.6 The Beneficiary will assist the personnel of the Supplier during their travels and on-site permanence to improve their safety although the Supplier shall be aware of the safety risks of the current situation and of its possible evolution on the field.
|  | 16. Training: | 16.1 The supplier must provide training on the use/operation of the equipment and all its accessories for at least one (1) to the end-users. The training shall include at least: 1) Operation instructions in normal functioning, 2) Daily and periodic user-related cleaning and maintenance, 3) Errors and malfunctions, 4) Main risks in the use of the equipment.
16.2 The supplier must provide training on the preventive and corrective maintenance of the equipment and all its accessories for at least one (1) calendar day to the maintenance staff.
16.3 The location of the training course shall be the location where the equipment is delivered and installed.
|  | 17. Warranty | 17.1 Twentyfour (24) months of manufacturer warranty. The warranty certificate will be in the name of the final Beneficiary.
17.2 Twentyfour (24) months of supplier warranty extending the Manufacturer warranty in scope not in duration. All maintenance shall be free of charge for the recipient institution throughout the warranty period. The warranty certificate will be in the name of the final Beneficiary. At least 95% of one single year of full functioning, i.e. 347 days out of 365 days, will be guaranteed by the Supplier within the warranty period. In the event that the item supplied has been malfunctioning for more than five percent (5%) of one single year of the warranty period, i.e. more than 18 natural days in one single year, the Supplier shall extend the warranty period for a duration of six times of the time duration when the equipment was malfunctioning. For this the supplier will make available on-line remote support for the maintenance of the supplied goods. This support shall be with a person trained by the manufacturer or with the manufacturer and available at least during business days. In case that with the on-line remote suport the issue cannot be solved the Supplier will send the trained technician on-site at the Beneficiary premise, being any cost of equipment transport or technician traveling at Supplier/Installer charge and included in the offered price. The time elapsed between the communication about the broken equipment and the intervention on site will be, inside the warranty period, not longer than 5 business days including in these 5 days all the remote support.
17.3 Twentyfour 24 months extended warranty shall include preventive maintenance and calibration of the equipment according to the manufacturers’ recommendations. The schedule of the visit will be presented to the Beneficiary before the final acceptance of the equipment. During any on-site maintenance and calibration visits a short user training update will be carried out by the Supplier.
17.4 The warranty period shall start from the date of acceptance of the goods by the end-user after training, testing and commissioning.
17.5 The equipment shall be serviced by an engineer properly certified by the equipment manufacturer (a copy of the service engineer’s certificate shall be provided in the offer).
17.6 The company providing the service shall have a presence in Ukraine (a guarantee letter from the equipment manufacturer shall be provided in the offer to this end, stating the service’s address in Ukraine).
17.7 After corrective and preventive maintenance visits, a report describing the technical procedures and results performed on the equipment shall be provided to the recipient institution and proper calibration shall be performed.
17.8 The bidder shall have the responsibilities and accountabilities that upon evaluation of equipment defects/faults, the equipment are appropriately repaired or replaced and shall be in good working condition thereafter.
|  | 18. Manuals | 18.1 The supplier must provide the end-users an electronic copy of the User manual in the Ukrainian language and the Service manual in the Ukrainian or English language for the equipment and accessories.

## Lot 5. Automatic blood separato
UNOPS-Ukraine-UPTF-2023-G-003
Brand
Model
Lot | Description | Technical specification | Value | Is Bid Compliant? Bidder to complete (Y / N) | Document and page where the technical specification or condition can be review | Details of the offered goods. Bidder to complete
3 | Automated Blood Separator: Medical device designed to automate the separation of blood components (RBC, plasma, platelet) from previously centrifuged blood units. Quantity: 1 unit. UNOPS reserves the right to increment quantity up to 2 units depending on available budget. | 1. General requirements | 1.1 Automatic blood components extractor to prepare quality blood components from centrifuged whole blood, plug compatible with Ukrainian standards.
1.2 Able to automatically seal the blood bag tubing at the end of the process.
1.3 Compatible with at least conventional Top&Top and Top&Bottom bags.
1.4 Average separation time of no more than 2 minutes and 30 seconds.
1.5 Press with servo motor system.
1.6 System consists of one main and side press.
1.7 Controlled by microcomputer.
1.8 Minimum 5 adjustable sealing heads.
1.9 Minimum 4 integrated scales.
1.10 Control of extracted component weight in real time.
1.11 Flow regulation.
1.12 At least 10 optical sensors for fine volume control.
1.13 Automatic air removal.
1.14. Touch-screen digital interface.
1.15. Internal memory for at least 800 extractions.
1.16. Data transfer capabilities:
1.16.1. LAN or WiFi.
1.16.2. At least 1 USB port.
1.17. Minimum 15 programs preconfigured by manufacturer.
1.18. Audiovisual alarms/messages: process failure, incorrect tube placement, incorrect bag placement on scales, press closing/opening error, press calibration error, flow regulation error, the program completed.
1.19. Scales and optoelectronic sensor reset and automatic calibration with every procedure.
|  | 2. Electrical requirements | 2.1 Power supply: 100-240 V, 50 Hz, single phase, plug compatible with Ukrainian standards.
|  | 3. Accessories included in the offered price | 3.1. Protective shield against liquid splashing.
3.2. Barcode scanner.
| Documents, certificates and standards required, to be presented by the bidder with the offer: | 4. Regulatory approval and marketing authorization in Ukraine, considering the Letter of the MOH No. 24-04/5396/2-22 dated 25 February 2022 and Resolution of the Cabinet of Ministers of Ukraine "On the List of Products of Critical Imports" No. 153 dated 24 February 2022, in accordance with at least one of the following: | 4.1 Valid and current product registration issued by the State Service of Ukraine on Medicines and Drug Control (SMDC) in conformance with the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable.
4.2 Compliance declaration and certification or documents proving the possibility of using the medical products following the procedure to assess compliance with the requirements of the Technical Regulation on medical devices approved by Resolution of Cabinet of Ministers of Ukraine No. 753 of 02/10/2013 or Technical regulation on medical devices for in vitro diagnostics approved by Resolution of Cabinet of Ministers of Ukraine No. 754 of 02/10/2013, as applicable and their updates.
4.3 A copy of the relevant statement by the Ministry of Healthcare on introducing certain medical products into circulation and/or use, for which the requirements of technical regulations have not been met, but whose use is justified from a healthcare perspective, for the batch of items in question.
|  | 5. Regulatory approval and marketing authorization issued by a Notified Body in at least one of the GHTF Founding Member countries, in accordance to the following list: | 5.1 European Union: Regulation (EU) 2017/745 on medical devices or Regulation (EU) 2017/746 on in-vitro diagnostic medical devices, as applicable
5.2 United States: FDA authorization
5.3 Canada: SOR/98-282
5.4 Australia: TGA conformity certification
5.5 Japan: PMDA pre-market approval
|  | 6. Certificate of Quality Management System | 6.1 According to ISO 13485 in the name of the Manufacturer for the medical device offered. The Certificate must be issued by an independent Certifying Body/Agency.
|  | 7. Letter of Authorization from Manufacturer | 7.1 Issued by the Manufacturer of the equipment authorizing the bidder to sell/distribute the offered equipment, or Authorization from the Official Distributor. Please consider that the authorization text must also indicate that the authorized Distributor is jointly committed with the bidder to provide the required post-sales services, guarantees, and training requested. Furthermore, authorization must be accompanied by the document that certifies that the Distributor has official authorization from the Manufacturer to sell/distribute the offered equipment.
|  | 8. User manual and technical brochure | 8.1 In Ukrainian and English
|  | 9. Guarantee conditions affidavit | 9.1 Affidavit of the bidder accepting the guarantee conditions established in the bidding documents and committing to deliver together with the equipment a guarantee certificate in the name of the beneficiary.
|  | 10. Spare parts affidavit | 10.1 Affidavit from the bidder guaranteeing the availability of spare parts and accessories for at least 5 years in case of award.
|  | 11. Maintenance affidavit | 11.1 Affidavit from the manufacturer stating the address in Ukraine of the company that will conduct maintenance of the equipment
|  | 12. Warranty, preventive and corrective maintenance affidavit | 12.1 Affidavit from the company that will carry out the maintenance (be it the supplier, manufacturer or a third party) committing to execute the guarantee and the preventive and corrective maintenance in accordance with the conditions indicated in the bidding documents
|  | 13. Personnel affidavit | 13.1 Copy of the training certificates of the engineer(s) that will conduct the installation and maintenance or affidavit for their training from the manufacturer
| Post Sale Services | 14. Delivery, installation, testing and commissioning: | 14.1 On-site delivery and installation to the recipient institution is included. Installation shall include connection to the Laboratory Information System when available.
14.2 The supplier shall transport the equipment inside the recipient institution to the installation site/room, open the packages, assemble and install it according to the manufacturer’s recommendations. Any damage to the equipment or accessories that occurred during transportation to the facility will be covered by insurance at the supplier's expense.
14.3 The supplier shall perform on-site testing, calibration and commissioning of the equipment with certification of conformity to standards and manufacturer’s recommendation. After installation, testing and calibration completion, the equipment is operational and ready to use. All equipment must be functioning with no physical damage and/or defect, prior to the acceptance by the end-user.
14.4 The supplier shall deliver any original software license in the name of the recipient hospital, when applicable, together with the equipment.
14.5 A complete set of commissioning forms with the entire set of tests performed and the results obtained shall be given to the recipient institution and to UNOPS once installation and commissioning is completed.
14.6 The Beneficiary will assist the personnel of the Supplier during their travels and on-site permanence to improve their safety although the Supplier shall be aware of the safety risks of the current situation and of its possible evolution on the field.
|  | 15. Training: | 15.1 The supplier must provide training on the use/operation of the equipment and all its accessories for at least one (1) to the end-users. The training shall include at least: 1) Operation instructions in normal functioning, 2) Daily and periodic user-related cleaning and maintenance, 3) Errors and malfunctions, 4) Main risks in the use of the equipment.
15.2 The supplier must provide training on the preventive and corrective maintenance of the equipment and all its accessories for at least one (1) calendar day to the maintenance staff.
15.3 The location of the training course shall be the location where the equipment is delivered and installed.
|  | 16. Warranty | 16.1 Twentyfour (24) months of manufacturer warranty. The warranty certificate will be in the name of the final Beneficiary.
16.2 Twentyfour (24) months of supplier warranty extending the Manufacturer warranty in scope not in duration. All maintenance shall be free of charge for the recipient institution throughout the warranty period. The warranty certificate will be in the name of the final Beneficiary. At least 95% of one single year of full functioning, i.e. 347 days out of 365 days, will be guaranteed by the Supplier within the warranty period. In the event that the item supplied has been malfunctioning for more than five percent (5%) of one single year of the warranty period, i.e. more than 18 natural days in one single year, the Supplier shall extend the warranty period for a duration of six times of the time duration when the equipment was malfunctioning. For this the supplier will make available on-line remote support for the maintenance of the supplied goods. This support shall be with a person trained by the manufacturer or with the manufacturer and available at least during business days. In case that with the on-line remote suport the issue cannot be solved the Supplier will send the trained technician on-site at the Beneficiary premise, being any cost of equipment transport or technician traveling at Supplier/Installer charge and included in the offered price. The time elapsed between the communication about the broken equipment and the intervention on site will be, inside the warranty period, not longer than 5 business days including in these 5 days all the remote support.
16.3 Twentyfour 24 months extended warranty shall include preventive maintenance and calibration of the equipment according to the manufacturers’ recommendations. The schedule of the visit will be presented to the Beneficiary before the final acceptance of the equipment. During any on-site maintenance and calibration visits a short user training update will be carried out by the Supplier.
16.4 The warranty period shall start from the date of acceptance of the goods by the end-user after training, testing and commissioning.
16.5 The equipment shall be serviced by an engineer properly certified by the equipment manufacturer (a copy of the service engineer’s certificate shall be provided in the offer).
16.6 The company providing the service shall have a presence in Ukraine (a guarantee letter from the equipment manufacturer shall be provided in the offer to this end, stating the service’s address in Ukraine).
16.7 After corrective and preventive maintenance visits, a report describing the technical procedures and results performed on the equipment shall be provided to the recipient institution and proper calibration shall be performed.
16.8 The bidder shall have the responsibilities and accountabilities that upon evaluation of equipment defects/faults, the equipment are appropriately repaired or replaced and shall be in good working condition thereafter.
|  | 17. Manuals | 17.1 The supplier must provide the end-users an electronic copy of the User manual in the Ukrainian language and the Service manual in the Ukrainian or English language for the equipment and accessories.