## L1 Laser photocoagulation syste
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 1: Laser photocoagulation system
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
| Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Working wavelength: 810 nm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Working hours : | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Continuous: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Pulse-periodic: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Micropulse: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Single pulse: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Pulse duration (depending on the mode) in the range not less than: 10 – 9000 ms. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Radiation power (depending on the method of delivery of laser radiation) in the range not less than: 50 mW – 3.0 W | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Cooling system – thermoelectric and/or air: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Colour touch control panel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Automatic detection of attached devices with information output to the display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Automatic shutdown in case of discrepancy between the parameters of the laser radiation power: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Emergency shutdown button: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Slit lamp adapter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Slit lamp: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Full integration of the laser beam delivery system into the optical path of the slit lamp: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Increase at least: 5х, 8х, 12х, 20х, 32х | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Slit width in the range at least: 0 – 14 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Slit height in the range at least: 1 – 14 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | The angle of inclination of the slit at least: ± 90° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Filters: cobalt blue, green, filter for protection against UV radiation, heat-absorbing filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Light source – halogen lamp (at least 6 V, 20 W) or LED: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Power supply:
23 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Network cable, Eurostandard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Components, spare parts, consumables:
25 | Specialised table for ophthalmic equipment, with electric drive (electric lifting table): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | The dimensions and load capacity of the table must correspond to the dimensions and weight of the proposed model: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Surgeon's elbow stands: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Indirect laser headband ophthalmoscope: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Probe for delivery of laser radiation: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Double Aspheric Lens Kit Corresponding to Anti-Reflective Coating (60D, 78D, 90D, 120D): 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Protective filter for use with an operating microscope: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | Foot pedal control: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Dust cover: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
37 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
D | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
38 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
39 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8,ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for theManufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
40 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
E | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
F | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
G | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
H | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
I | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
J | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
K | Documentation and Licences
41 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
42 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
43 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
L | Installation and Support
44 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
M | Remote Support
47 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L2 Video ophthalmoscope
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 2: Video ophthalmoscope
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Descriptions: Indirect video ophthalmoscope is an apparatus for detailed diagnosis of eye pathology, analysis of the retina, fundus vessels and optic nerve with video recording of the procedure. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical specifications:
3 | Image recording function: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | The ability to connect to various peripheral devices: monitor, printer, computer, with the ability to capture a video image: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Innovative optical system and high-sensitivity camera sensor for high-resolution clear images: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Fully integrated camera system: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Automatic light measurement in the centre of the image: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Lowlight function: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Waterproof camera head. Possibility of treatment with a disinfectant solution: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | High-resolution camera: CCD 470,000 pixels, 460 lines. Sensor size — 1/2'' CCD, colour or compatible: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Focus adjustment for any working distance in the approximate range: ot 250 to 800 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Auto White Balance: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Camera Control: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Image with S-VHS monitor or similar: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Outputs FBAS (composite) and Y/C (SVHS): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | PAL or NTSC formats optional: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
17 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
19 | Video ophthalmoscope (optics on the head helmet with C-mount)::1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | A-cam camera: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Processor: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Video cable: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | S-VHS or compatible cable: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Scleral depressors large and small: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Lines for ophthalmoscopy A.R. 20D 50 мм: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Transformer: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Spare lamp: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Transport case: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
30 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
31 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
32 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
33 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
34 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
35 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
36 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
37 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Remote Support
40 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L3 Non-mydriatic fundus-camera
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 3: Non-mydriatic fundus-camera
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Descriptions: Non-mydriatic fundus camera used in non-contact fundus examinations, as well as treatment. The fundus is illuminated by an optical installation system. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specification:
3 | Maintain high quality images: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Software that allows you to optimise diagnostic studies, conduct a course of treatment, the ability to optimally control the digital installation, process images, as well as register them: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | The ability to compare various studies, calculate diagnostic indicators, analyse colour RGB channels and perform other actions: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Special mode of photography of the anterior segment of the eye: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Autoexposition mode for shooting: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | The camera interface supports the DICOM standard: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Sensitivity: ISO 1600, 3200/6400 (low flash mode) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Flash brightness: auto exposition | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Software: RX Capture | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Digital camera not less than: 15 Mpx | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | The number of pixels of the matrix at least: 15.1mln | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Image registration: specialised digital camera | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Pupil diameter: at least ø 4.0 mm and at least ø 3.3 mm (in SP mode) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Angle of view: 45° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Shooting modes: Colour photography, photo in red light (digital filter red free), photo using cobalt filter (digital filter cobalt) or (Exciter, barrier, green, ICG-exciter, ICG-barrier):: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Camera type - non-midriatic: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Magnification: 2x (digital) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Operating distance in the range at least: 30 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Light Source: surveillance - IR LEDs; registration - white LEDs | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Special joystick for various functions: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | High Quality Focusing: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Movement of the product body in any direction: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Registration of received images: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Switching from the anterior eye segment to the retina: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
27 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
29 | Equipment table : 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | LapTop not less than 17" or PC with monitor at least 19": 1pc | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Program module "mosaic" or analogue: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
33 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
34 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
35 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
36 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
37 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
38 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
39 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
40 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Remote Support
43 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L4 Digital wide field retinal c
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 4: Digital wide field retinal chamber, paediatric
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Descriptions: Compact, portable, and easy-to-use mydriatic wide-field digital imaging system used for the photo documentation of paediatric ocular diseases that manifest in an infant’s eyes. | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
2 | Operational Requirements: The ergonomically designed lightweight handpiece allows for single-handed operation and captures high-resolution images of the paediatric retina. The camera is a specialty device used for paediatric retinal imaging. | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
3 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
B | Technical Specifications:
4 | Retinopathy of prematurity: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
5 | Retinoblastoma (retinal tumour): | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
6 | Shakenbaby syndrome (retinal haemorrhage): | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
7 | Anterior segment of the eye and external pathologies in children and adults: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
8 | Interchangeable lens nozzle (handle) containing the camera | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
9 | High Resolution Digital Camera: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
10 | Real-time image acquisition: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
11 | Get freeze frames from a video without losing resolution: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
12 | Frame-by-frame viewing of the video to exclude the possibility of missed pathologies: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
13 | Video recording at least: 120 sec | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
14 | Examination of patients in a horizontal position: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
15 | Photo and video shooting with the help of a nozzle (handle) without the use of additional equipment and any other accessories: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
16 | Lens for retinopathy of prematurity with a field of view at least: 120° | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
17 | The maximum field of view of the fundus per picture at least: 120° | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
18 | The width of the lenses in the part of contact with the eye at most: 9 mm | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
19 | Fibro-optical light source: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
20 | Freeze frame function to capture the image: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
21 | Black balance adjustment function: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
22 | Control the light intensity and focus of the camera during video and photo shooting using the pedal: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
23 | Preservation, comparison of images obtained before and after treatment to identify the changes that have occurred: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
24 | Connection harness of the nozzle, including three separate cables (fibre optic cable for object lighting; cable controlling the camera; focus motor cable): | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
25 | Control pedal: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
26 | Laptop with installed software (Windows 7 at least, diagonal of at least 15"; RAM at least 4 GB; hard disk capacity of at least 500 GB; DVR-RW): | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
27 | Specialised trolley for placing equipment: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
28 | Transport handle: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
29 | 4 wheels, two of which with a braking device: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
30 | Compartment for storage of accessories: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
C | Power supply:
31 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
32 | Network cable, Eurostandard connector: | ☐ Yes ☐ No | As per Section III.2a attach the Manufacturer’s Authorisation
D | Components, spare parts, consumables:
33 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
34 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
35 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
36 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8,ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
37 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
38 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
39 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
40 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
41 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Remote Support
44 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L5 Specialised chemotherapy cha
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 5: Specialised chemotherapy chair
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Description: Haemodialysis, Chemotherapy, Blood transfusion. Upholstery and armrests are critical to patient comfort. Depending on age and health status, patient access in the seated and other positions. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Operational Requirements: Multifunctional automated easy controllable chair | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specifications:
4 | Haemodialysis: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Chemotherapy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Blood transfusion: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Minimum 4 and maximum up to 9 electric motors: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | 4 wheels with brake: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Adjustable headrest: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Side cushions: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Armrests | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Paper roll holder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Adjustable dropper tripod: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Multimedia equipment with a "touch screen" screen or remote control: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | 3 chair positions: "0" position, "bed" position for heart massage, Trendelenburg position: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Motors that prevent pressure sores (optional): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Electric drive: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
18 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
20 | Ncollar tripod for dropper: 1pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Paper roll holder: 1pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Multimedia equipment (optional): 1pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Multifunctional Stand: 1pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Motors for vibration (optional): 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Side cushions: 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Double pockets for personal belongings: 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Reading lamp: 1pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
29 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
30 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
31 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
32 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
33 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
34 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
35 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
36 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Remote Support
39 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L6 ENT Workstation
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 6: ENT Workstation
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Description: Designed to integrally diagnose and treat ENT patients, using the microscopic/endoscopic visual system, X-ray film view box, surgical instrument tray, warmer, a wide range of medical bottles | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specification:
C | Equipment:
3 | Specialised modular console: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Vacuum aspirator: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Endoscopic video camera: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Specialized Medical LCD Monitor: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Patient's chair: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Doctor's chair: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Observation lamp, shadowless: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Cold Light Source: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Headlamp lighting lamp: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Tools, components, spare parts, consumables: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | General description of the specialised modular console:
13 | Floor modular design (at least 2 modules): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | The frame is made of double-walled panels (material - galvanised steel) with a seal: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | The surface is painted with high-quality powder paint by spraying with thermal firing: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | All work surfaces should be easy to clean and be resistant to disinfectants: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | The tool compartment is covered with a transparent acrylic lid: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | The dimensions of the console should allow the free arrangement of equipment and tools: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Each module has 4 rollers for easy movement: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Modular console equipment:
20 | Stainless steel retractable table top: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Drawer, deep, flat (at least 4 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Garbage can be opened with a foot pedal: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Tool tray (at least 6 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Removable tray for used tools with a special silicone mat to protect tools from mechanical damage: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Removable plastic containers for disinfection of endoscopes (at least 3 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Heater of otorhinolaryngological mirrors: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Instrument box with UV sterilisation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Ear washing and drainage system for the external auditory canal: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Eustachian tube purge system: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Aspiration tank: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Liquid medicine sprayer, direct (at least 3 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Liquid medicine sprayer, curved: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Powdered Drug Dispenser: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | The system of washing and cleaning hoses with a disinfectant solution: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Capacity for dosage forms (not less than 6 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Slide-out keyboard panel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | System of connection to water supply and sewerage: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Electronic dashboard: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | Monitor mount, rotatable (2 pcs.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | Holder for headlamp lighting lamp: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | Camera head holder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | Heated endoscope holder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | Shelf-gallery for tools and accessories: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
44 | Pendant bracket for vacuum aspirator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | General description of the vacuum aspirator:
45 | Maximum vacuum at least: 60 kPa | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | Aspiration tank with a volume at least: 1500 ml | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
47 | Tank material - glass, polycarbonate, polysulfone: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
48 | The tank cover is autoclavable, with an overflow protection mechanism: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
49 | Automatic switching on when the suction hose is removed: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
50 | Automatic tank emptying system with flushing function: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
51 | Vacuum regulator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
52 | Bacterial filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | General description of the endoscopic video camera:
53 | Resolution at least: 750 x 580 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
54 | Digital zoom at least: 2,5x | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
55 | Auto White Balance: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
56 | Electronic filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
57 | Freeze frame function: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
58 | Picture-in-picture (PIP) function: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
59 | Archiving photos and videos to a flash card: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
60 | Adaptation to endoscopes of different manufacturers (specify manufacturers): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
61 | Various input/output signals (specify): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
62 | Foot pedal for saving images: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
H | General description of a specialised medical LCD monitor:
63 | Diagonal at least: 22” | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
64 | Resolution at least: 1920 x 1080 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
65 | Aspect ratio: 16:9 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
66 | Contrast ratio at least: 600:1 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
67 | Viewing angle (vertical/horizontal) at least: 1780 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
I | General description of the patient's chair:
68 | Drive - mechanic-hydraulic: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
69 | Height adjustment in the range at least: 55 – 75 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
70 | The angle of inclination of the backrest of the chair at least: 90° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
71 | Rotation of the chair at least: 360° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
72 | Base diameter at least: 70 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
73 | Upholstery material - synthetic leather: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
74 | Adjustable headrest and armrests: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
75 | Fixed footrest: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
J | General description of the doctor's stool:
76 | Material – chrome-plated steel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
77 | Pyatiluche's reason: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
78 | Seat diameter at least: 40 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
79 | Braking mechanism: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
80 | Height adjustment in the range at least: 55 – 75 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
81 | Adjustment by gas spring: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
K | General description of the cold light source:
82 | Light source – LED: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
83 | Automatic overheating protection system: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
84 | Light intensity adjustment: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
85 | The service life of the LED at least: 30 000 hours | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
L | General description of the headlamp/lighting lamp:
86 | Adjustable headband with leather lining: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
87 | Flexible light guide with a length of at least 220 cm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
88 | Focusing the light spot: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
89 | Fixators for the light guide on the headband: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
90 | Adjustable head tape size: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
M | Power supply:
91 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
92 | Network cable, Eurostandard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Components, spare parts, consumables:
93 | Otoscope (0°/2.7 mm/50-60 mm): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
94 | Sinusscope (0°/4.0 mm/175-180 mm): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
95 | Laryngoscope (0°/4.0-5.0 mm/175-180 mm): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
96 | Nasal cannula for adults: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
97 | Negatoscope 600х400 mm: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
98 | Bacterial filter: 50 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
99 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
100 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
O | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
101 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
102 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
103 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
P | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
Q | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
R | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
S | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
T | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
U | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
V | Documentation and Licences
104 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
105 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
106 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
W | Installation and Support
107 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
108 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
109 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
X | Remote Support
110 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
Y | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L7 Laser therapy system for ENT
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 7: Laser therapy system for ENT
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Description: Specialised system for use in surgical ENT practice: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Operational Requirements: Benign and malignant neoplasms of the oral and nasal cavities, pharynx, larynx, trachea, ear canal, auricle: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specifications:
4 | Tonsillectomy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Papillomatosis: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Subglottic laryngeal stenosis and tracheal stenosis: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Stapedectomy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Cholesteatoma: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Myringotomy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Laser type – CO2: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Wavelength: 10 600 nm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Operating modes – continuous, pulse, superpulse: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Maximum output power in continuous mode at least: 40 W | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Maximum output power in pulse mode at least:35 W | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Maximum output power in superpulse mode at least: 15 W | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Special manipulator with scanner for delivery of laser energy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | At least 6 scanning shapes: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Control of the incision length, area and depth of exposure to laser energy: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | The system shall ensure maximum preservation of the attached healthy tissues: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Control panel - TOUCH LCD display (diagonal not less than 10"): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Internal Memory - Patient Database and Treatment Protocols: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | English and/or Russian language support in the user interface: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
23 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
25 | Flexible optical cable for laser energy delivery (compatible with endoscopic instruments from different manufacturers): 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Protective glasses, adjustable (for personnel): 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Protective glasses, adjustable (for the patient): 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
29 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
30 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
31 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
32 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
33 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
34 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
35 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
36 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Remote Support
39 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L8 Heart lung machine
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 8: Heart lung machine
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered Bidder to complete
A | General description:
1 | Description of Function: Heart Lung Machine is an apparatus through which blood is temporarily diverted, during heart surgery, to oxygenate it and pump it throughout the body, thus maintaining circulation until the heart and lungs are able to return to normal functioning. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specification :
3 | Movable console for 5 pump modules: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | The pumps should show accurate measurements and the machine should be used for performing paediatric and neonate surgery. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Pumps should preferably have a direct drive system and maintenance free. All the Pumps should have pulsatile mode in build. Each Module should work its own. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Vertical racks for equipment placement: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Horizontal shelf or rod: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Availability of additional connectors for connecting external devices: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Roller pump with central occlusion adjustment mechanism: 3 pcs | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Double indication (rev/min, l/min): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Rotation speed in the range not less than: 0 – 250 rpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Flow at 1/2 inch tube in a range not less than: 0 – 10 l/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Direction of rotation in and counterclockwise direction: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Ability to operate two pumps in the "master – slave" mode: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Reflection of all pump parameters on the display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Power from the console: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Small roller pump: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Double indication (rev/min, l/min): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Rotation speed in the range not less than: 0 – 250 rpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Flow rate at 1/4 inch tube in the range not less than: 0 – 1,5 l/min. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Reflection of all pump parameters on the display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Power from the console: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Centrifugal pump:
24 | Flow in the range not less than: 0 – 9 l/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Rotation speed in the range not less than: 0 – 3500 rpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Ability to work at a given flow rate: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Integrated bubble sensor, level sensor and flow sensor: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Availability of emergency drive: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Power from the console: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Reflection of all pump parameters on the display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Control board for placing at least 6 display modules: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Pressure Monitor:
33 | Dual pressure control with two independent channels: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | Pressure adjustment by automatically changing the pump speed: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Monitoring of pressure channels: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Pressure control in the range not less than: - 250 – + 800 mm Hg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | Acoustic and visual alarm; | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Temperature Monitor:
39 | Measurements in the range not less than: 0°С - + 50°С | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | The accuracy of temperature determination is not more than: ± 0,3°C | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | Acoustic and visual alarm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | Machine should show all warnings and alarms in complete text messages and different audible tones with different colour coding. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | Oxygenator Blood Level Monitor/Bubble Detector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
44 | Regulation of pump speed depending on blood level: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | Servoregulation of the arterial pump by blood level: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | Touch sensors 1/2,1/4, 3/8 inch: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
47 | Acoustic and visual alarm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
48 | Pulsating flow control unit: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
49 | Cardioplegia control unit: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
50 | Timer: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
51 | Gas faucet (for adults): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
52 | Venous clamp:
53 | Venous clamp with the possibility of fine and rough adjustment of the venous flow: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
54 | Unit for monitoring the level of charging of batteries and voltage in the power supply network: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
55 | Complete software package: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
56 | Notebook or Tablet PC for data management: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
57 | Thermostatic control device on at least 2 water circuits: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
58 | The thermostatic control device (heater/cooler machine should be compatible to get integrated into Heart Lung machine). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
59 | The unit shall be capable of operating continuously in ambient temperature of 2 - 40 degree celsius. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
60 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
61 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
62 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
63 | Technical documentation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
64 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
65 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
66 | Built-in rechargeable battery, which ensures the operation of the device for at least 1 hour: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
67 | Automatic battery discharge when connecting the device to the centralized power supply network: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
68 | Console lamp: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
69 | Replacement lamp: 4 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
70 | Pressure sensor: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
71 | Pressure sensor holder: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
72 | Pressure sensor cap: 50 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
73 | Pressure sensor cable: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
74 | Perfusate level sensor in the oxygenator: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
75 | Perfusate level sensor holder: 50 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
76 | Venous clamp holder: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
77 | Connector curved (90 degrees): 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
78 | Ice container with the possibility of fastening to the console: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
79 | Blanket for warming/cooling (for adults): 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
80 | Adapter to blanket with connecting tubes: 2 sets | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
81 | Necessary devices (connectors, circuits, lines, etc.) for connecting a thermostatic device: 2 sets | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
82 | Disinfectant for thermoregulatory device: 20 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
83 | Device for connecting a gas mixer: 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
84 | Device for attaching the oxygenator to the AIC: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
85 | Temperature sensors:
86 | Cutaneous: 10 pcs | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
87 | To measure esophageal and rectal temperature: 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
88 | To measure the temperature in the lines (tubes): 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
89 | For direct temperature measurement in the oxygenator: 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
90 | Aortic cannulas:
91 | 14 Fr: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
92 | 16 Fr: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
93 | 18 Fr: 3 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
94 | 20/21 Fr: 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
95 | 22 Fr: 15 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
96 | 24 Fr: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
97 | Venous cannulas:
98 | 22 Fr: 2 pcs | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
99 | 24 Fr: 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
100 | 26 Fr:15 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
101 | 28 Fr: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
102 | 30 Fr: 15 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
103 | 32 Fr: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
104 | 34 Fr: 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
105 | Venous cannulas are two-stage:
106 | 32/34 - 39/40 Fr: 10 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
107 | 36 - 50/51 Fr: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
108 | Cannulas cardioplegic retrograde: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
109 | Cannulas cardioplegic antegrade:
110 | 9 Fr: 15 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
111 | 11 Fr: 15 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Oxygenator for adults: 50 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
112 | Basic requirements:
113 | Type membrane, hollow fibre: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
114 | Blood flow at least:7 l/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
115 | Primary filling volume not more than: 260 ml | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
116 | The surface area of the gas exchange membrane shall not exceed: 2,5 м2 +/- 5% | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
117 | The capacity of the venous tank is not less than: 4000 ml | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
118 | Holder for oxygenator: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
119 | Hemoconcentrator for adults: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
120 | Disposable centrifugal pump: 20 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
121 | Compressor for AIC:1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
122 | By the time of delivery to the customs territory of the Republic of Uzbekistan, the shelf life of consumables (with a limited shelf life) must be at least 80% of the established one. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
123 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
124 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
F | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
125 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
126 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
127 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
G | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
H | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
I | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
J | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
K | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
L | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
M | Documentation and Licences
128 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
129 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
130 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
N | Installation and Support
131 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
132 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
133 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
O | Remote Support
134 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
P | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L9 Slit Lamp
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 9: Slit Lamp
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Increase not less than: 5-6,5х, 10х, 16х, 25х | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Enlargement of the eyepiece at least: 12,5x | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Adjustment of the interpupillary distance in the range not less than: 49 – 75 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | The range of adjustment of diopter eyepieces is at least: - 5D - + 5D | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Slit width in the range at least: 0 – 14 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Slit length in the range at least: 1 – 14 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Lighting field diameters at least: 8/5/3/2/1/0.2 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | The angle of inclination of the slit is at least: 90° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Filters: cobalt blue, green, filter for protection against UV radiation, heat-absorbing filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Light source – halogen lamp (at least 6 V, 20 W) or LED: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Movement of the device: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Longitudinal at least: 72 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Transverse at least: 100 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Vertical at least 30 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Vertical movement of the chin stop at least: 60 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | External fixation label: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Power supply:
17 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Network cable, Eurostandard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Components, spare parts, consumables:
19 | Compact electronic tonometer for measuring IOP according to Goldman:1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Specialised table for ophthalmic equipment, with electric drive (electric lifting table): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
| Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | The dimensions and load capacity of the table must correspond to the dimensions and weight of the proposed model: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Halogen lamp (at halogen light source): 2 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Dust cover: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
27 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
D | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
28 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
29 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
30 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
E | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
F | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
G | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
H | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
I | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
J | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
K | Documentation and Licences
31 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
32 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
33 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
L | Installation and Support
34 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
M | Remote Support
37 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L10 Ophthalmic Workstation
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 10: Ophthalmic Workstation
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Description: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Equipment:
2 | Automatic phoropter (digital refractor): 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Digital scoreboard: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Floor stand: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Patient's chair: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Automatic perimeter: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Portable ophthalmic tonometer: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Ophthalmic diaphanoscope: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Automatic phoropter (digital refractor):
9 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Country of origin: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Technical documentation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Fully computerised system for determining refractive error in the most comfortable conditions for the patient and the doctor: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Measurement range:
14 | Sphere in the range at least: - 28 D - + 26 D | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Measurement step: 0.25/0.5/1.0 D or wider range | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Cylinder in the range at least: - 6 D - + 6 D | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Measurement step: 0.25/1.0 D or wider range | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Cylinder axis in the range at least: 0° - 180° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Measurement step: 1°/5° or more wide range | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Prism in the range at least: 0 D – 20 D | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Measurement step: 0,1 D/1,0 D or more wide range | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Adjustment of the interpupillary distance in the range at least: 48 – 80 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Adjustment step: 0,5m/1,0 m | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Built-in additional lenses:
24 | Auto Cross Cylinder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Occluder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Point diaphragm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Red Maddox prism (horizontal/vertical): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Red/Green Filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Polarizing filter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Retinoscopic lens: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Touch monitor control (at least 8" diagonal): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Built-in thermographic printer: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Digital scoreboard:
33 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | TYPE – LCD: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Diagonal not less than: 19” | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | Mounting – wall mounted: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | Circular polarization: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | Distance for testing in the range at least: 3 – 6 m | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | Full Remote Control: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | A wide range of optotypes, including children's: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
44 | The bidder must specify all available optotypes: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | Full compatibility with the offered automatic phoropter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Floor stand:
46 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
47 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
48 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
49 | Technical documentation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
50 | Floor stand on a solid base for fastening the foropter: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
51 | Maximum load on the phoropter holder at least: 10 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
52 | Free horizontal rotation of the phoropter holder: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
53 | Lighting lamp (LED): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
54 | Swivel stand for touch monitor: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
H | Patient's chair:
55 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
56 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
57 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
58 | Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
59 | The chair should be made of durable material, resistant to the effects of various disinfectant solutions: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
60 | The coating of the chair of dark shades (black, grey), resistant to the effects of various disinfectant solutions: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
61 | Seat rotation at least: 120° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
62 | Height adjustment in the range at least: 450 – 600 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
63 | Foot pedal adjustment: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
64 | Maximum load at least: 150 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
65 | Folding armrests: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
66 | Footrest: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
I | Automatic perimeter:
67 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
68 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
69 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
70 | Technical documentation: present | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
71 | Measurement modes: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
72 | Static perimetry: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
73 | Kinetic perimetry: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
74 | Colour perimetry: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
75 | Binocular perimetry: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
76 | Detection of glaucoma at the initial stage (test "blue on yellow"): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
77 | Fixation Monitoring: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
78 | Digital eye tracking with a camera (automated eye tracking): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
79 | Blind spot monitoring by Hejil-Krakau: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
80 | Measuring area: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
81 | Radius at least: 300 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
82 | Measured angle at least: 90° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
83 | Parameters of the stimulus: I, III, V according to Goldman: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
84 | The colour of the stimulation spot is white, blue, red: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
85 | The maximum brightness of the stimulus at least: 10 000 Asb | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
86 | Standard research tests: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
87 | General at least : 30° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
88 | Glaucoma at least: 50° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
89 | Macula at least: 10° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
90 | Patient database: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
91 | Number of programmed visual fields at least: 7 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
92 | Number of research methods at least: 7 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
93 | Workstation: LCD monitor (colour, specialised medical) with a diagonal of at least 24" (resolution not less than 1920 x 1080), RAM of at least 8 GB, hard disk capacity of at least 1 TB, network card of at least 1000 Mbps, CD/DVD-RW, colour laser printer, UPS (online) with a working time of at least 10 minutes: 1 set | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
J | Portable ophthalmic tonometer:
94 | Model (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
95 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
96 | Country of origin: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
97 | Technical documentation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
98 | Portable, handheld, lightweight model: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
99 | Measurement modes: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
100 | IOP (true) according to Goldman: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
101 | IOP (tonometric) according to Maklakov: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
102 | Measurement of Goldman IOP in the range not less than: 7 – 50 mm Hg Art. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
103 | Limits of error in measuring IOP according to Goldman: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
104 | From 7 to 23 mm Hg. Art. at most: ± 2 mmHg Art. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
105 | Over 23 mmHg. Art. at most: ± 5 mmHg Art. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
106 | Limits of error when measuring IOP according to Maklakov: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
107 | From 15 to 28 mm Hg. Art. at most: ± 2 mmHg Art. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
108 | Over 28 mm Hg. Art. at most: ± 5 mmHg Art. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
109 | IOP measurement time at most: 5 sec. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
110 | Measuring method - non-contact: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
111 | Indication of parameters on the LCD display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
112 | Symbolic indication (modes, battery charge, readiness for measurement, errors, etc.): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
113 | Storage in volatile memory of the last mode used: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
114 | Power supply – rechargeable battery or non-rechargeable batteries (AA type, 1.5 – 3.0 V): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
K | Ophthalmic diaphanoscope:
115 | Material of manufacture – stainless steel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
116 | Light source – halogen or xenon lamp: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
117 | Bending angle at least: 45° | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
118 | Lighting power adjustment: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
119 | Special handle for ease of use: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
120 | Power supply – rechargeable battery or non-rechargeable batteries (AA type, 1.5 – 3.0 V): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
L | Power supply:
121 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
122 | Network cable, Eurostandard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
M | Components, spare parts, consumables:
123 | Paper for automatic foropter thermal printer: 10 rolls | ☐ Yes ☐ No | Attach Supporting Statement
124 | Dustproof case for automatic foropter: 1 pc. | ☐ Yes ☐ No | Attach Supporting Statement
125 | Dust cover for automatic perimeter: 1 pc. | ☐ Yes ☐ No | Attach Supporting Statement
126 | Specialised table for ophthalmic equipment, with electric drive (electric lifting table):1 pc. | ☐ Yes ☐ No | Attach Supporting Statement
127 | Model (full and exact name): | ☐ Yes ☐ No | Attach Supporting Statement
128 | Manufacturer (full and exact name): | ☐ Yes ☐ No | Attach Supporting Statement
129 | Country of origin: | ☐ Yes ☐ No | Attach Supporting Statement
130 | Technical documentation: present | ☐ Yes ☐ No | Attach Supporting Statement
131 | The dimensions and carrying capacity of the table must correspond to the dimensions and weight of the proposed model of the automatic perimeter: | ☐ Yes ☐ No | Attach Supporting Statement
132 | Case (case) for storing a portable ophthalmic tonometer: 1 pc. | ☐ Yes ☐ No | Attach Supporting Statement
133 | Case (case) for storing the ophthalmic diaphanoscope: 1 pc. | ☐ Yes ☐ No | Attach Supporting Statement
134 | Spare halogen or xenon lamp for ophthalmic diaphanoscope: 2 pcs. | ☐ Yes ☐ No | Attach Supporting Statement
135 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
136 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
N | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
137 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
138 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
139 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
O | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
P | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
Q | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
R | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
S | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
T | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
U | Documentation and Licences
140 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
141 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
142 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
V | Installation and Support
143 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
144 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
149 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
W | Remote Support
150 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
X | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L11 Semi-automated Sealing Mac
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 11: Semi-automated Sealing Machine
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Descriptions: Universal semi-automatic seaming machine for capping glass and plastic bottles with aluminium and plastic caps. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specifications:
3 | The presence of replaceable sleeves for various types of bottles, bottles and caps: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | High-quality closure with smooth and screw caps: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | The presence of 2 duplicate start buttons on the common panel and next to the closure post to simplify the operator's work: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | The presence of foot pedal to increase productivity and reduce the load of the operator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | The presence of a protective fence of the capping head, to ensure the safety of the operator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | All components and parts must be made of stainless steel, caprolon or have an anti-corrosion decorative coating: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Easy maintenance - it is enough to install a bottle with a cap in the sleeve on the turntable and wait for the bottle to enter the seaming zone: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Performance at least: 3000 pcs/time | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | The diameter of the body of the vials at least: 9-90 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | The diameter of the lids at least: 7-49 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | The height of the bottles at least: 40-450 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Number of capping heads at least: 1 pc | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power supply:
15 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
17 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
18 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
19 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
20 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
21 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
22 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
23 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
24 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
25 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L12 Manual Machine for Sealing
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 12: Manual Machine for Sealing of Aluminium Caps
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Description: The device is designed for crimping aluminium caps when sealing glass bottles for blood and blood substitutes, as well as medical vials made of glass and a dry for medicines. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Link to the proposed model: The participant must provide a link to the proposed model on the Manufacturer's website in English. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Technical Specifications:
3 | Compact, portable device of the desktop type: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Weight at most: 10 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | In the possibility of quickly changing crimp heads and stands: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Reconfiguration to different types of bottles and vials: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Drive type: manual or automated | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Devices for crimping caps for glass bottles for blood and blood substitutes according to GOST 10782-85 or similar standard for following volumes - 50, 100, 250, 450, 500 ml: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Devices for crimping caps for vials for canned blood, previously produced according to TU 21-01 215-69 or similar standard for following volumes - 50, 100, 250, 450, 500 ml: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Devices for crimping the caps of glass bottles for medicines according to TU 64-2-10-77 or similar standard for following volumes - 20, 30 ml: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Types of crimped caps according to OST 64-7-85-79 or similar standard for following ampoules size - K-2, K-3, K-6: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Components, spare parts, consumables:
12 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
13 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
D | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
14 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
15 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
16 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
E | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
F | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
G | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
H | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
I | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
J | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
K | Documentation and Licences
17 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
18 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
19 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
L | Installation and Support
20 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
M | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L13 Lab Animals Vertical Venti
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 13: Lab Animals Vertical Ventilation System
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | System of individually ventilated cells for keeping laboratory animals: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | General requirements:
1 | Vertical ventilated caging system (including AHU) with at least 8 shelves/racks with capacity to hold at least 90 or more small size cages: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Control system with feedback of airflows and pressure: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Display of temperature and humidity values of supply and exhaust inside the cells on the control panel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Modular caging system with possibility to connect racks with cages on both sides of the ventilation unit so that cages can be added: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Can be used with different combinations of cells/types of racks i.e from Rats to Hamster: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Possibility of connection to the external ventilation system through the probe: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Possibility of data transfer via USB or bluetooth etc: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Must supply a tablet and appropriate software for remotely controlling of equipment via bluetooth channel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Technical Specifications:
9 | Working hours: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Top mounted Air handling unit with a touch screen control pannel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Plenums shall be removable and autoclavable to maintain sterility: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Mode of operation with positive pressure: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Mode of operation with negative pressure: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Pressure drop at the level of "microisolator" / "indoor shrinking medium", approximate range: '+ 20...+ 50 / -20...-50 Pa': | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Dimension: (10% variation will be accepted): 1400х490х1850 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Weight not more than: 110 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Support construction - 4 swivel wheels, 2 of them with locks: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Product materials :
18 | Case material - steel with polymer coating: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Duct material - non rusting steel: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Filters:
20 | Supply air pre-filter: Class G4 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Pre-filter of exhaust air: Class G4 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Supply/exhaust air filter type: HEPA (class H13 and/ or H14) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Replacement of the exhaust air filter without contact with the environment: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Control panel - colour touch screen or control buttons: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Control Panel interface in Russian and English:
26 | Shall have operators manually built in and accessible from the control panel. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Monitoring sensors: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Flow volume control (supply and exhaust air): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Flow rate sensor:2 pcs | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Display of volumetric flow rate and air exchange multiplicity on the display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Temperature sensor: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Humidity sensor: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Autonomous operation of the equipment in case of power outage: 30 min to 3 h | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | Maximum noise level not more than: 55 dB | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Maximum electrical power not more than: 500 W | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Power supply cable with a length of at least: 2,5 m | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Power supply:
37 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Network cable, European standard connector: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Components, spare parts, consumables:
39 | Cages for rodents: fully loaded (e.g: 90 units if the frame capacity is of 90 cells) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
42 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
H | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
43 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
44 | ISO 9001 and/or any of applicable ISO standards (ISO 13485, ISO 7176-8, ISO 21882, ISO 15004, ISO 11145, ISO 10939, ISO 11658) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
45 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
I | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
J | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
K | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
L | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
M | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
N | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
O | Documentation and Licences
46 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
47 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
48 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
P | Installation and Support
49 | The equipment must be installed, tested and commissioned by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
50 | The supplier must offer a complete solution with all the necessary cables, accessories, parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
51 | The supplier (if necessary) must provide the necessary spare parts (the cost of which must be included in the tender offer) for full functioning during the warranty period. A list of spare parts to be submitted in the bid. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
Q | Remote Support
52 | Remote access for on-line technical support/consultation for solving technical and application-related problems without delaying the workflow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
R | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L14 Defibrillator Monitor
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 14: Defibrillator Monitor
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | Description:
1 | Portable defibrillator with vital signs monitoring | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Configuration:
2 | Light weight and portable: less than 5 kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Shall provide Biphasic energy waveform: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | The unit shall operate in manual mode and semi-automatic defibrillation (AED mode): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | The unit shall be designed to withstand dust/water, shock, drops and vibration: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Built-in LCD display: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | With built in pacemaker (demand and fixed rate): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Modes:
8 | Synchronized and a synchronised defibrillation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Non-invasive pacing: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | ECG mode: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Energy:
11 | User selectable, 1 to 200 J (increment according to range): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | 200 Joule recharging time: < 7 seconds | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Display:
13 | Russian language support in the user interface: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | ECG/Heart rate monitoring: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Minimum 3-lead ECG cable: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Alarm/self test system:
16 | With advanced automatic self-test at switch on: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Audiovisual warning in case of errors and measured parameters out of limit: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Accessories:
18 | Defibrillation/pacing Paddles: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | ECG cables: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Printer:
20 | High-resolution thermal printer: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Annotation: Time, date, ECG lead, heart rate, defibrillation/pacing settings | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Printer paper: 5 rolls | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
H | Power Requirement:
23 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Electrical outlet (direct AC operation): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | With internal rechargeable batteries, minimum of 30 shocks at maximum energy or monitoring time: 3 hours (minimum) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | With battery status indicator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | With internal memory and ability to have summary: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | With lead fault indicator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
I | Spare parts, consumables:
29 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
30 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
J | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
31 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
32 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
33 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
K | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
L | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
M | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
N | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
O | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
P | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
Q | Documentation and Licences
34 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
35 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
36 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
R | Installation and Support
37 | The equipment must be installed and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
S | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L15 Ventilator
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 15: Ventilator
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
1 | Suitable and certified for emergency and transport | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Flow-triggered ventilation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Adult and paediatric patients | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Invasive and non-invasive ventilation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Microprocessor controlled. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Integrated Turbine or piston based mechanical ventilator. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Integrated air/oxygen mixer and FiO2 control with integrated alarms. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Ventilation Modes:
8 | Controlled and assisted ventilation modes, at least: PSV, VCV, PCV, SIMV, CPAP, BIPAP, PEEP, NIV: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Assist-control ventilation (A/C): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Volume Control Ventilation (VCV): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Pressure controlled ventilation (Pressure Control Ventilation, PCV): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Synchronised intermittent mandatory ventilation (SIMV): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Synchronised intermittent mandatory ventilation + pressure support SIMV-PS: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Bilevel positive airway pressure (BIPAP): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Continuous positive airway pressure (CPAP): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Positive End-Expiratory Pressure (PEEP): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Non-invasive ventilation (NIV): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Measurement Parameters:
18 | Tidal volume (50 ml – 1500 ml): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Inspiration time (0.3 – 5.0 sec): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Respiratory rate (5 – 60 bpm): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Peak flow (up to 120 Litres/min): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Minimum inspired flow (not less than 20 ml/min): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Flow trigger: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | PEEP/CPAP (0 – 15 cmH2O): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Pressure support (0 – 50 cmH2O): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Oxygen concentration (21 – 100 %): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Inspiration/Expiration ratio (1:3 – 2:1): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Display:
28 | LCD/ TFTcolor graphic display of at least 8 inches: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Display patient data with curves, trends and alphanumeric forms visualisation: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Monitoring of at least: I/E time, O2 concentration, Flow, volume, pressure and functional lung parameters: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Backup battery indicators: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Alarms:
32 | Alarms at least: Peak airway pressure, Minimum expiratory pressure, HighRespiratory rate, Low/High Tidal Volume, low O2 concentration, I:E Ratio limit exceeded, Power failure: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Allow the use of disposable patient circuits, and bacterial filters in inspiratory and expiratory circuits; or/and single branch patient circuits with bacterial filter connection at the expiratory valve outlet: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Components:
34 | Sterilizable/disinfectable inspiratory block: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
35 | Hoses, colour and connection coded, DIN standard: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Oxygen and air plugs/connectors, DIN standard: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
37 | Shock and weather proof carry case: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Accessories to be included:
38 | Expiratory valve (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | Testing bag, 1 litre | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
40 | Oxygen sensor (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | Forty (40) pieces of disposable Bacterial/Viral filters | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | Twenty (20) sets of disposable Breathing circuits (inspiratory and expiratory limbs) pieces | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | Twenty (20) pieces of disposable masks (small, medium and large), textured for easyfit | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
H | Power Requirements:
44 | 110/220V, 50/60Hz mono-phase electrical source | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | Built-in rechargeable battery with at least 4 hours autonomy and automatic recharge | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | 12V alternative source for ambulance use, included vehicular plug | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
I | Spare parts, consumables:
47 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
48 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
J | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
49 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
50 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
51 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
K | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
L | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
M | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
N | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
O | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
P | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
Q | Documentation and Licences
52 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
53 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
54 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
R | Installation and Support
55 | The equipment must be installed and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
S | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L16 Oxygen Concentrator
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 16: Oxygen Concentrator
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | Applications:
1 | Emergency medical care | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Intensive care unit and semi-intensive care unit: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Possibility of use in ambulances (inverter connection for ambulance vehicle) : | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Features:
4 | The number of oxygen outlets with adjustable flow rate is not less than: 2 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Weight not more than: 35 Kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Maximum outlet pressure: not less than 80 KPa | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Maximum flow not less than 10 L/min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | The oxygen concentrator shall be capable of generating at least 55 kPa at all flows, up to the maximum rated flow | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Oxygen purity in range between 82% and 95% or better | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Equipped with flowmeters with continuous flow-rate control from 0 LPM to the maximum rated flow-rate, at a minimum of 0.5 LPM intervals. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | The oxygen concentrator shall be prevented from providing a flow rate greater than the maximum rated flow rate | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Time needed to obtain a concentration of more than 90% not more than: 3 min | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Noise level not more than 52 dB | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Display:
14 | Timer (shows the total number of hours of operation of the device): | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Oxygen concentration indicator: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Alarms:
16 | High temperature alarm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Power supply failure | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | High and low oxygen pressure | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Low battery | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Low oxygen concentration (< 82%) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Outlets:
21 | Oxygen outlet for direct attachment of oxygen delivery tubing | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | Outlets shall be robust barbed fittings | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Components:
23 | Oxygen humidifier (Qty 20) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Nasal cannula with connectors (Qty 40) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Extension tube (Qty 6) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Adapter tube (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Inlet air filter (Qty 10) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Bacterial filter (Qty 4) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Humidifier stand/holder | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Power requirements:
30 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Built-in rechargeable battery: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Battery capacity (new, fully charged) not less than: 1 hour of full operation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Inverter for connection to ambulance power supply: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Spare parts, consumables:
34 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
35 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
36 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
37 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
38 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
39 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
40 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
41 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
42 | The equipment must be installed and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L17 Wheel chair
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 17: Wheel chair
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Chair mounted on wheels, intended for indoor and outdoor transportation of individuals with walking disability: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Foldable design: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Stainless steel or powder or chrome coating or equivalent frame: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Frame shall be rust and chip-resistant: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Anti-bacterial, washable and incombustible seat, backrest and armrest: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Folding armrests and footrests: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Armrests and footrests are adjustable in height: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Two tubeless lockable rear wheels with a diameter of approximately 600 mm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Two front swivel casters with approximate diameter of 200 mm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Seat dimensions:
10 | Height in range: 45 to 50 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Width not less than: 45 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Depth not less than: 40 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Maximum load not less than: 110 Kg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Components, Spare parts, consumables:
14 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
15 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
D | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
16 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
17 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
18 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
E | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
F | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
G | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
H | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
I | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
J | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
K | Documentation and Licences
19 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
20 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
M | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L18 Patient Monitor
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 18: Patient Monitor
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Patient vital signs monitor, 6 physiological parameters: ECG, heart rate, respiration rate, NIBP, SPO2 and temperature | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Suitable for adults, paediatric and neonates | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | AC and battery powered | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Automatic self-test and continuous system monitoring. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Display:
5 | Russian language support in the user interface: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Colour flat display ≥ 10.4" | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Display of minimum 7 wave forms simultaneously. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | ELECTROCARDIOGRAM and HR (Heart Rate):
8 | Minimum 5-lead connection | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Simultaneously display of a minimum of 2 ECG traces | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | The attenuation coefficient of common-mode signals is not less than: 90 dB | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Heart rate in the range not less than: 15 - 300 bpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | The error of measuring the heart rate is not more than: ± 1 bpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | ST Segment analysis | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Analysis of arrhythmias (indicate the types of arrhythmia) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Defibrillator sync and protection. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Non-invasive pressure:
16 | Measurement method – oscillometric: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Measured parameters – systolic pressure, diastolic pressure, average pressure, pulse rate: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Manual and automatic modes | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Minimum diastolic and maximum systolic values should be equal or better than: 10 – 250 mmHg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Resolution: min 2mm Hg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | Maximum average error not more than: 5 mm Hg | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Pulse oximetry:
22 | Measuring range SpO2: 1- 100 % | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
23 | Accuracy in the range of 70 % - 100 % not more than: ± 3 % | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | Heart rate detection range derived from the SpO₂ is equal to or better than: 30 – 250 bpm. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | Error not more than: ± 3 bpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Respiratory rate:
26 | Measurement method – impedance: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Measuring range (for adults) not less than: 6 - 150 bpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
28 | Error: ± 3 bpm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Temperature: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Measuring range not less than: 15 - 45 degree | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Accuracy (including sensor) not more than: ± 0,3 degrees | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
H | Alarms:
31 | Acoustic and visual alarms: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
32 | Setting alarms for all parameters: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | Selection of volume level of the alarms: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
34 | Apnea Alarm: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
I | Data storage:
35 | Ability to save a minimum of 72 events: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
36 | Trend storage ≥ 24 hours | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
J | Power requirements:
37 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
38 | Built-in rechargeable battery, ensuring the operation of the device for at least 2 hours: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
39 | Data and network interface, LAN, USB or equivalent. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
K | Accessories
40 | ECG patient cable terminals (1 x neonatal/paediatric, 1 x adult). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
41 | ECG reusable electrodes (1 x neonatal/paediatric, 1 x adult). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
42 | ECG electrode, disposable, adhesive, for adults: 100 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
43 | ECG electrode, disposable, adhesive, for neonatal/paediatric,: 50 pcs. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
44 | SpO2 sensor, adult, reusable, with cable (Qty 2). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
45 | SpO2 sensor, for paediatric, reusable, with cable (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
46 | SpO2 sensor, for neonates, reusable, with cable (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
47 | NIBP hoses (1 x neonatal, 1 x paediatric, 1 x adult). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
48 | Blood pressure cuffs (1 x neonatal, 1 x paediatric, 1 x adult). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
49 | Skin temperature probes including cable. (Qty 2) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
50 | Mobile Stand: 1 pc. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
L | Components, spare parts, consumables:
51 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
52 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
M | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
53 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
54 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
55 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
O | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
P | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
Q | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
R | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
S | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
T | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
U | Documentation and Licences
56 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
57 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
58 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
V | Installation and Support
59 | The equipment must be installed and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
W | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L19 Functional Bed, mechanical
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 19: Functional Bed, mechanical
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Emergency Unit Bed, with mattress | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Manually-operated | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Bed Structure:
3 | Four sections adjustable bed | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
4 | Front and back panel | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | High-strength perforated steel bed, coated with impact-resistant epoxy or equivalent: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | With adjustable and removable folding side rails. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Mounted on 4 swivel castors, heavy duty, 2 with brake. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
8 | Protective bumpers at corner | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Functions:
9 | Manually operated foldable crank allows adjusting backrest (0-70°), | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Manually operated foldable crank that allows leg section and foot section (0-40°). | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Materials:
11 | High resistance to corrosion | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Epoxy painted tubular steel frame | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Water-repellent fabric mattress | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
14 | Caster wheel: mold-on type, non-hooded | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Ball-bearing swivel wheel | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Dimensions:
16 | Bed dimensions not less than: 200 x 85 cm (LxW) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Height of bed frame, without mattress: 50 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Mattress thickness: minimum 10 cm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
19 | Wheel diameter: 125 mm | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
20 | Load capacity not less than: min. 180 kg. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Components:
21 | Fitting mattress with cover | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Spare parts, consumables:
22 | The supplier (if necessary) must complete the equipment (taking into account the specifics of the proposed model) with all the necessary parts, components, materials (the cost of which should be included in the tender offer) for assembly, installation and commissioning at the workplace. Completeness should ensure the full functioning of the declared equipment. | ☐ Yes ☐ No | Attach Supporting Statement
23 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
H | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
24 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
25 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
26 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
I | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
J | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
K | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
L | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
M | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
N | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
O | Documentation and Licences
27 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
28 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
P | Installation and Support
29 | The equipment must be assembled and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
R | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## L20 Electrocardiogram (ECG)
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Technical specifications for goods – Comparative Data Table
Item 20: Electrocardiogram (ECG)
Model (full name) |  |  | Insert Details
Manufacturer (full name) |  |  | Insert Details
Country of Origin |  |  | Insert Details
S/N | Item Specification | Is quotation compliant? Bidder to complete | Details of goods offered. Bidder to complete
A | General description:
1 | Multi-channel electrocardiography portable system, designed for recording and analysing the ECG at resting | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
2 | Start/Stop buttons on the unit and on the ECG module | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
3 | Operating modes – automatic and manual | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
B | Display:
4 | Touch screen: not less than 12" | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
5 | Number of channels: 12 | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
6 | Simultaneous registration of 10 ECG leads | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
7 | Russian language support in the user interface: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Measurement Features:
8 | Common mode rejection ratio to be higher than 90 dB | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
9 | Input impedance shall not be less than 10 M Ω | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
10 | Signal sampling rate shall not be less than 8000 sample/sec | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
11 | Recording speed 25 mm/sec or better: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
12 | Automatic and manual sensitivity adjustment: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
13 | Signal noise suppression filters and techniques such as low and high pass, AC noise, baseline wander, muscle movements | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Signal Processing:
14 | Waveform (interval) analysis for width and amplitude | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
15 | Pacemaker detection | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
16 | Arrhythmia Analysis | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
17 | Heart rate computation | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
18 | Internal memory is not less than: 100 ECG | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
E | Alarms:
19 | Alarms for loose electrodes, artefacts, baseline wander, segment distortions, irregularities in ECG | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Thermal printer:
20 | Multi-channel (3/6/12 channels) printing: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
21 | High resolution, automatic feeder | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
22 | 300 metres paper rolls | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
F | Connections:
23 | Wireless or wired of data to a personal computer: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
24 | USB Port: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
G | Power Requirements: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
25 | 220 V ± 10 %, 50/60 Hz: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
26 | Built-in rechargeable battery: | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
27 | Battery life (new, fully charged) not less than: 2 hours | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
C | Accessories:
28 | ECG cables and electrodes | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
29 | Extra Battery Pack | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
30 | Mobile cart with at least one shelf or drawer. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
31 | Electrode gel 250 ml (Qty 4) | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
D | Components, spare parts, consumables:
32 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
33 | The supplier (if necessary) must equip the equipment (taking into account the specifics of the proposed model) with all the necessary spare parts (the cost of which must be included in the tender offer) for its full operation during the warranty period. A list of such spare parts shall be submitted in the bid. | ☐ Yes ☐ No | Attach Supporting Statement
E | Quality and Safety Standards (As per Section II.2b - Manufacturer & Supplier Certification)
34 | Copy of Bidder’s QMS (preferably compliant to ISO 9001 or other applicable quality management system requirements) is provided. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
35 | ISO 9001 and/or any of applicable ISO standards (ISO 10936, ISO 10934, ISO 21073) for the Manufacturer of the offered goods | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
36 | ISO 14001 EMS certificate, or similar, for the Manufacturer of the offered goods. | ☐ Yes ☐ No | As per Section II.2b, attach the certifications
F | Product Certification (As per Section II.2c - Product Certification) | ☐ Yes ☐ No | As per Section II.2c, attach the certifications
G | The awarded bidder is required to provide, where applicable, UNOPS with the hazardous property sheet (MSDS or alike) and documents under which the quality assurance was processed in the Manufacturer’s country, as shown in Operation and Service Manuals, within 2 weeks after award. | ☐ Yes ☐ No | Insert details and submit supporting documents as necessary
H | Confirm that Assembly and Calibration shall be completed as per Section II.2d of the Schedule of Requirements | ☐ Yes ☐ No | Provide the details and the relevant specific information regarding the service.
I | Confirm that Training services shall be provided as per Section II.2e of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the names and addresses of all the parties/institutions that will provide the services.
J | Confirm that Warranty and After-sales services shall be provided as per Section II.2f of the Schedule of Requirements | ☐ Yes ☐ No | Insert details of how these services are offered, including the name and address of the local agent (if any) that will carry out the services.
K | Confirm that General Requirements as per Section II.2g of the Schedule of Requirements are met | ☐ Yes ☐ No | Insert details of how these requirements are met
L | Documentation and Licences
37 | Operating instructions in English and/or Russian and/or Uzbek language (if not available, please confirm that they shall be supplied together with the equipment): | ☐ Yes ☐ No | Insert details of the goods offered
38 | Maintenance manual in English and/or Russian and/or Uzbek (if not available, please confirm that they shall be supplied together with the equipment). | ☐ Yes ☐ No | Insert details of the goods offered
39 | All software shall support English and/or Russian language, shall be original and officially licensed to the beneficiary, without additional fees/subscriptions required for operating the equipment over its lifecycle. | ☐ Yes ☐ No | Insert details of the goods offered
M | Installation and Support
40 | The equipment must be powered and tested by the supplier at each workplace separately. | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals
N | Confirmation statement on the unloading of the goods within the final destination as per RFQ specified Locations | ☐ Yes ☐ No | Insert details of goods offered, including relevant specifications and references to manuals

## Service and Delivery Requiremen
F
Form D: Technical Quotation Form
RFQ reference no: RFQ/2023/45678
Name of Bidder: [insert name of bidder]
Date: [insert submission date]
Bidders are required to complete the Comparative Data Tables below to demonstrate compliance with UNOPS requirements and inserted below. Bidders are NOT allowed to make any change in the “UNOPS requirements” columns of the Comparative Data Tables. Such changes might disqualify your quotation.
Service requirements – Comparative Data Table
S/N | Description of Services | Is quotation compliant? Bidder to complete | Details Bidder to complete
1 | The Supplier shall provide all the necessary documentation in a timely manner for facilitating the Customs Clearance process by the Consignee. This includes the Certificate of Registration / Authorization / Waiver for the import and use of the devices by the health regulatory authority of the Republic of Uzbekistan that shall be obtained by the Supplier prior to the goods’ arrival to the country. Any demurrage caused due to incomplete documentation shall be borne by the Supplier. | ☐ Yes ☐ No | Insert details
2 | After clearance of the goods from Customs by the Consignee, the Supplier (or its nominated local agent, partner or subcontractor) shall take the goods to its storage facility in Tashkent (or its vicinity), where the goods shall be segregated for further transportation to destination hospitals per the allocations provided in Table 2 of the Schedule of Requirements. | ☐ Yes ☐ No | Insert details
3 | Storage in Tashkent, delivery to hospitals, storage at hospitals together with associated safety and security arrangements for the goods until their handover to the authorities at each hospital shall remain the responsibility of the Supplier. The Supplier shall arrange for and maintain proper insurances for all the in-country storage and transportation activities. | ☐ Yes ☐ No | Insert details
4 | Storage of goods at Supplier’s warehouse in Tashkent (or its vicinity) shall be at no cost to UNOPS, unless deliveries to final destinations are delayed at UNOPS request. | ☐ Yes ☐ No | Insert details
5 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Republican Specialized Scientific and Practical Medical Center of Oncology and Radiology. Address: Shifokorlar Street, Almazar District, Tashkent city. | ☐ Yes ☐ No | Insert details
6 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Syrdarya regional branch. Address: Gulistan city, Dustlik MFY, Khondamir street, 1 | ☐ Yes ☐ No | Insert details
7 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Djizzakh regional branch. Address: Djizzakh region, Sharof Rashidov district, Samadova street, 6 | ☐ Yes ☐ No | Insert details
8 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Kokand city regional branch. Address: Kokand city, Kukon Sadosi Ruznomasi street, 6 | ☐ Yes ☐ No | Insert details
9 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Namangan regional branch. Address: Namangan city, Afrosiab street, 12 | ☐ Yes ☐ No | Insert details
10 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Samarkand regional branch. Address: Samarkand city, Amir Shoh Murod street, 86 | ☐ Yes ☐ No | Insert details
11 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Ferghana regional branch. Address: Fergana city, Gulistan street, 11 | ☐ Yes ☐ No | Insert details
12 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Andijan regional branch. Address: Andijan city, Y. Atabekov street, 5 | ☐ Yes ☐ No | Insert details
13 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Kashkadarya regional branch. Address: Karshi city, Chap sohil street, 28 | ☐ Yes ☐ No | Insert details
14 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Navoi regional branch. Address: Navoi city, F. Mardonov street, 2 | ☐ Yes ☐ No | Insert details
15 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Bukhara regional branch. Address: Bukhara city, Gijduvoniy street, 71 | ☐ Yes ☐ No | Insert details
16 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Surkhandarya regional branch. Address: Termez city, Ibn Sino street, 39 | ☐ Yes ☐ No | Insert details
17 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Khorezm regional branch. Address: Urgench city, A. Bahodirhon street, 176 | ☐ Yes ☐ No | Insert details
18 | Delivery, installation or assembly, calibration, testing, commissioning, users’ training services as well as maintenance during warranty period at Karakalpakstan regional branch. Address: Nukus city, Jumanazar street, 2 | ☐ Yes ☐ No | Insert details
Delivery requirements –– Comparative Data Table
UNOPS Requirements |  | Is quotation compliant? Bidder to complete | Details Bidder to complete
Delivery time for the goods | Bidder shall make the goods available at FCA (Incoterms2020) point within 60 days after Contract signature and/or deliver them to Tashkent, Uzbekistan within 90 days after Contract signature. | ☐ Yes ☐ No | Insert details
Delivery time for the services | Bidder shall deliver the goods from Tashkent to all the final destinations across the country and perform the related services within 60 days after goods’ arrival to their warehouse in Tashkent (or its vicinity). | ☐ Yes ☐ No | Insert details
Delivery place and Incoterms rules | FCA (Incoterms 2020) Bidder’s quoted point DPU (Incoterms 2020) Tashkent, Uzbekistan Door-to-door to all 14 locations across Uzbekistan per Table 2 of the Schedule of Requirements. | ☐ Yes ☐ No | Insert details
Consignee details | Consignee details for the preparation of shipping documents shall be provided at the time of contract signature. | ☐ Yes ☐ No | Insert details
UNOPS Right to vary requirements | At the time the Contract is awarded, UNOPS reserves the right to vary the quantity of the goods and associated services specified above, provided this does not exceed +/- 20%, without any change in the unit prices or other terms and conditions of the RFQ. | ☐ Yes ☐ No | Insert details