Examination Table:

Product description: Table, examination

Specifications:

* Examination table, in 2 sections Mounted on 4 sturdy supports, all finished with height adjustable feet.
* Both sections fitted with non-removable padded upholstery.
* Backrest angle adjustable via secured pawl and gear ratchet, safe for patient and operator.
* When fully extended, both sections align to perfectly flat surface.
* Transfer bars connect all lower distal portions of the 4 supports, providing maximal structural strength.
* Materials: High resistance to corrosion (tropical environment) Frame: epoxy coated tubular steel Adjustable feet: rubber or nylon Padded upholstery: high-density polyurethane foam, density 27-33 kg/m3Cover: plastic, flexible highly tear resistant, anti-static, flame retardant, disinfectant- and liquid proof, washable
* Dimensions: Examination table, two sections extended, including upholstery: 160-200x45-55x72-88cm (l x w x h) Frame: 2.7-3.3cm (outside, across), 1.35-1.65mm (thickness) Upholstery: 4.5-5.5cm (h) Carrying capacity: 135-165kg
* Knockdown construction: yes. Supplied with 1 x complete set of tools required for assembly List of accessories and parts
* Detailed step-by-step instructions for assembly and safe use, text-free pictorial based (i.e. line-drawings only)
* Packaging, labelling, instructions One (1) unit per box Identify Packaging Standards and provide Packaging Test Reports Dimensions: Unit Weight in Kg (including its packaging)Unit Volume in M3 (including its packaging) Dimensions of box, length x width x height in labelling: Compliance with EAN 128 bar code requirements

Special Notes:: All dimensions are to be interpreted with 10 % manufacturing tolerance.