## Instructions for Forms
Section II Schedule of Requirements
e-sourcing reference: ITB/2023/48792
Equipment for the Treatment of Infectious and Sharp Waste from Healthcare Facilities
Summary of Requirements UNOPS requirements are comprised of 12 items listed below. Quantities are undefined. UNOPS will not be committed to purchase any minimum quantity of goods and purchases will be made only if and when there is an actual demand. • Bidders are NOT allowed to make any change in the specifications listed. Such changes might disqualify your bid. Please be guided accordingly.
Item N | Item Description | Is bid Compliant? Bidder to complete with Yes or No
1 | Waste autoclave, 100 l. | Yes/No
2 | Waste autoclave, 300 l. | Yes/No
3 | Waste autoclave, 500 l. | Yes/No
4 | Waste autoclave, 1000 l. | Yes/No
5 | Hybrid autoclave, 20 kg/h | Yes/No
6 | Hybrid autoclave, 50 kg/h | Yes/No
7 | Hybrid autoclave, 100 kg/h | Yes/No
8 | Hybrid autoclave, 250 kg/h | Yes/No
9 | Microwave with integrated shredder, 20kg/h | Yes/No
10 | Microwave with integrated shredder, 50 kg/h | Yes/No
11 | Microwave with integrated shredder, 100 kg/h | Yes/No
12 | Microwave with integrated shredder, 250 kg/h | Yes/No
• Placeholders have been created in the eSourcing tender where the Excel Technical Bid Forms must be uploaded for each item. Failure to submit this form duly completed as per the ITB instructions may result in disqualification of your bid • Bidders must insert the specifications/details of their offered vequipment under Column D to demonstrate compliance with the technical criteria. • DO NOT SIMPLY MARK ‘YES/NO/COMPLIANT’ OR COPY AND PASTE THE TECHNICAL SPECIFICATIONS INTO COLUMN D OF EACH SHEET. • Not providing specifications under this column may render your bid non-compliant.
General Disclaimer
Equipment must be supplied with all electrical connections according to the destination country standards
In the event of termination of production of spare parts: advance notification to the purchaser of the pending termination, in sufficient time to permit the purchase.
Name and title of the person filling out the form
Contact details:
E-mail
Phone number
Address
Signature
Date

## Items 1-4. Waste Autoclave
1: Waste autoclave (treatment of solid infectious and sharp waste) |  | Compliant/ Non- Compliant | Insert your Specifications. Bidders are required to provide information on the specifications of their bid to demonstrate compliance to set specifications. Not providing specifications under this column may render your bid non-compliant. | Provide supporting documentation.
Minimum Requirements: |  | INSERT SUPPLIER NAME HERE |  | Technical data sheet
Supplier Name
Supplier Address
Manufacturer Name
Manufacturer Address
Mnufacturing Site Address
Model Number
Waste autoclave | Vacuum autoclave, medium
General requirements | Decontamination of solid infectious and sharp waste by steam in a pure, saturated steam atmosphere by a pre-vacuum or fractionated vacuum-steam process. | YES/NO
| Size of treatment chamber: 100, 300, 500, or 1000 liter ±20%. Bidder to specify | YES/NO
| 1 door version | YES/NO
| The autoclaving units are designed to remain operational in a tropical environment or in proximity to saltwater for a minimum period of 10 years, provided that the specified maintenance schedule is followed. All metal surfaces are resistant to atmospheric and process-induced corrosion, using non-corrosive stainless steel (e.g., UNS: S31603 / S30400). Supplier to indicate used metal in accordance with AISI, UNS or equivalent norms. | YES/NO |  | Stainless steel raw material technical data sheet (optional)
| Inactivation of vegetative bacteria, fungi, lipophilic/hydrophilic viruses, parasites and mycobacteria at a 6 Log10 reduction or greater; and inactivation of G. stearothermophilus spores and B. atrophaeus spores at a 4 Log10 reduction or greater (high level disinfection / STAATT Level III criteria). | YES/No |  | Validation result or testing results of the same model autoclave in former project, or testing results from an external accredited entity like Robert Koch Institute or Pasteur Institute or an accredited third party laboratory listed on https://ilac.org/about-ilac/
| Decontamination of waste with ≥121°C for ≥20 minutes or ≥15 minutes with a treatment temperature of ≥ 134°C. Supplier to indicate temperature and time. | YES/NO
| Built-in or attached automatic electric steam generator - capacity adapted to needs. | YES/NO
| Build in or attached vaccuum system (Providing of the maximum vacuum of the autoclave in both units: Bar and Pa) | YES/NO
| Air compressor - capacity adapted to needs (if applicalble). | YES/NO
| Equipped with condensate treatment system. | YES/NO
| Equipped with a filter or treatment system for potentially contaminated air from the autoclave. | YES/NO
| Fully automatic steering of the treatment process. | YES/NO
| Including of a visual monitoring system of the whole treatment cycle (time, temperature, pressure) - clearly readable analogue or digital system operation panel (english or local language). | YES/NO
| The chamber and the steam supply are equipped with pressure gauges. | YES/NO
| Indicate the number of all available pre-programmed cycles with details and minimum cycle time (in minutes). | YES/NO
| Bidder to provide information on average water and electricty consumption at the rated capacity: liter / cycle and electricity: kW / cycle | YES/NO
| Equipment expected lifespan in years | YES/NO
Safety Features | Equipped with a readily accessible emergency shut-off switchesfor immediate cycle shut-off. | YES/NO
| Included safety mechanism which prevent the cycle being started until the door is locked and sealed and to prevent the door being opened until chamber pressure is at atmospheric. | YES/NO
| Each pressure vessel (chamber etc.) equipped with pressure relief safety valves or equivalent pressure limiting device to prevent excess pressure. | YES/NO
| Air decontaminatin by a HEPA filter (0,2 µm, heat resistant, steam resistant, hydrophobic and pressure resistant for the pressure difference) or stream treatment or other equipvalent methods to prevent the release of pathogenic aerosols. Suppliers to describe the method/decontamination process. | YES/NO
| Information about the noise level of the autoclave during operation (dB). | YES/NO
| Steam and condensate discharged from the chamber is cooled to reduce discharge temperature to less than 60 degrees C to avoid risk of damage to the building drainage system. | YES/NO
| Furnished with external surface insulation to mitigate the risk of elevated temperatures, ensuring that surfaces do not exceed 50°C. This precautionary measure is implemented to protect personnel. | YES/NO
Loading | Manual loading of waste / batch loading | YES/NO
| 2 stainless steel loading baskets for processing waste bags designed to use the maximum chamber capacity and equipped with a liquid collection tray for each unit. | YES/NO
| 1 trolley made of stainless steel to load and unload the loading basket for each unit. | YES/NO
Electrical supply, generation and wiring | Electricity supply: The waste autoclave is designed for 3 phase, 400 V a.c. ±10% or for country specific conditions (e.g. 3 phase, 230 V a.c. ±10%) Information on electricity supply options available to be provided. | YES/NO
| All electrical connections attached with the autoclave and other machinery are provided with suitable and necessary safety devices and be protected by steel conduits, cable trays, and supports. | YES/NO
Certificates and markings | Marking on the autoclaves indicating: (i) manufacturer’s name or trademark (ii) style, type, model or, date of manufacture of the autoclave (iii) autoclave capacity (chamber volume) | YES/NO |  | Mock Label
| ISO 9001 Certificate - Quality management of the manufacturer or ISO 13485 for Medical Devices - Quality management | YES/NO |  | ISO 9001 Certificate or ISO 13485 Certificate
| Compliant with EN 285:2006+A2:2009 for large sterilizer or an equivalent standard (provided by a certified independent third party). | YES/NO |  | Validation result or testing results of the same autoclave type in former project, or testing results from an external accredited entity for example Robert Koch Institute or Pasteur Institute, an accredited third party laboratory listed on https://ilac.org/about-ilac/
| Markings /certificate for pressure vessel safety in compliance with EN 13445 or ASME Boiler and Pressure Code or equivalent. | YES/NO |  | Pressure vessel Certificate
| Calibration certificates of the pressure and temperature measurement devices. | YES/NO |  | Calibration certificates
| IEC 60601 Electrical equipment for medical use or IEC 61010-2-040, UL 61010A-2-041, or equivalent electrical safety standard. | YES/NO |  | Declaration of Conformity (DoC)
Documentation | Operating and maintainance / service manual including a list of proecedures required for local routine maintenance in English to be translated into the local language. | YES/NO |  | Manual in English language
| Short user manual for the autoclave operator (Safety Operation Procedure) in English to be translated into the local language. | YES/NO |  | SOP in english language
| Set of engineering drawings: Lay out drawing with all main equipment (dimentions), main piping, PID flow diagram, installation requirements. | YES/NO |  | Engineering drawings and installation requirements.
Warranty | Minimum 2 years warranty included in the price. Starting after the installation, testing and comissioning is finalized. | YES/NO
Packaging, labelling | Packaging, labelling and shipping must be compliant with national and international standards and regulations and in accordance with UNOPS packing and labelling instruction as indicated in the UNOPS General Conditions of Contract (GCC). | YES/NO
Please fill in information:
Name of manufacturer:
Origin country:
Bidder must be able to supply all the items listed below.
16
| Factory acceptance test (FAT): the system, prior to shipment, will be tested for conformity of the system with the manufacturer’s performance specifications and the minimum requirements specified. The supplier / manufacturer is responsible to provide the FAT report. An examplary FAT report to be provided. | YES/NO
Installation, Testing and Commissioning requirements | All associated housing, fittings and connections to enable full functionality of the autoclave shall be completed by the client (local entity where the autoclave will be installed). These shall be in line with the manufacturer’s instructions; including but not limited to housing, foundation, piping, electrical and water connections, drains, ventilation, etc., as applicable. | YES/NO |  | Manufacturer installation ´s guideline
| The supplier of the equipment is responsible for the installation and commissioning of the equipment. | YES/NO
| On-site acceptance test (SAT): the system, after delivery, will be tested by the contractor together with the user to demonstrate that the performance meets the manufacturer’s performance specifications. The waste decontamination efficiency of the autoclaves will be proofed after installation by adequate tests (series of biological, chemical, thermoelectric tests) under supervision of the client / an independent third party. The results of the testing of the system will be documented by the contractor in an acceptance protocol that is signed and dated by the user. | YES/NO
Operator training | Providing of onsite training for users and technicians: The training focus on providing hands on training to the client’s staff (operator, technicians) regarding the operation, testing, maintenance, trouble shooting and repair of the autoclave supplied under this Contract. This training session will be conducted at the client’s location, where the autoclave is installed. Followed by the training, the trainer supervise the operation of the autoclave by the client for at least 1 additional day. | YES/NO |  | Examplary training agenda / materials.
| Amount of days for necessary training . The trainer will be provided on a day-rate reimbursable basis. The bidder must specify the minimum scope of work in terms of operator working days . Rates for these to be provided on the financial offer form and MUST NOT include expenses of travel ticket and Daily Subsistence Alowance (DSA), such as hotel, food, transportation. These will be calculated according to UNOPS' standard travel and DSA fees. | YES/NO
Spare parts and post-installation service | Spare part package for 1 years to be provided. | YES/NO |  | List of spare parts with unit costs to be delivered.
| Declaration of spare parts availability post-warranty of minimum 5 years. | YES/NO
| List of countries where local service can be made available and contact details of local / regional service agents . | YES/NO |  | List of countries/regions and contact details
OPTIONAL | Description of types and cost of service contract of 5 years | YES/NO |  | filled out tables 6 and 7 in price form C
OPTIONAL: Accessories and Consumables
Assessories and consumables (optional - costs to be indicated separately and not included in the overall price of the quipment and above components) | Fresh water treatment system: sediment backwash filter and water softener system (water tank, water softening salt for startup, resin column) and all necessary mechanical and electrical elements needed for the activation and working of the station. | YES/NO
| Voltage limiter for electricity supply of the control system of the autoclave (if applicalble). | YES/NO
| Water meter (m³) | YES/NO
| Energy meter (kwh) | YES/NO
| 500 autoclavable bags: Strong plastic bags, single use, autoclavable at 121 /134 °C (heat resistant) based on the indicated waste treatment temperature, marked with biohazardous symbol, color: yellow | YES/NO
Providing of autoclave process testing equipment (in accordance with EN ISO 11140-1-4):
| Process Challenges Device for chemical indicators (PCD test): - Teflon tubing which is connected to a capsule inside where a chemical indicator strip is put. - 1,5 m long – 2mm diameter. - Disposable after 250 uses | YES/NO
| supplier to indicate shelf life of the above PCD tests |  |  | Indication of shelve life of product
| Chemical Indicators strips for the PCD: - Number 250 | YES/NO |  | Indication of shelve life of product
supplier to indicate shelf life of the above PCD test strips
| Chemical indicators test stripes: - Class 5 Integrating Indicators - Number: 250 | YES/NO
| supplier to indicate shelf life of class 5 indicators |  |  | Indication of shelve life of product
| Bowie Dick test 121 °C / 134 °C: - Number: 50 - Lead-free and no toxic heavy metals | YES/NO
| supplier to indicate shelf life of Bowie Dick tests |  |  | Indication of shelve life of product
| 100 Biological spore tests vials: Endospores of B. stearothermophilus tests (strains NCTC 10007, NCIB 8157, ATCC 7953). | YES/NO
| Supplier to indicate shelf life of vials |  |  | Indication of shelve life of product and storage requirements .
| Incubator for biological spore test vials: 55-60 °C | YES/NO

## Items 5-8. Hybrid Autoclave
2: Hybrid autoclave with integrated shredder (treatment of solid infectious and sharp waste) |  | Compliant/ Non- Compliant | Insert your Specifications. Bidders are required to provide information on the specifications of their bid to demonstrate compliance to set specifications. Not providing specifications under this column may render your bid non-compliant. | Provide supporting documentation.
Minimum Requirements: |  |  |  | Technical data sheet
Supplier Name
Supplier Address
Manufacturer Name
Manufacturer Address
Mnufacturing Site Address
Model Number
General requirements | Autoclave with integrated shredder. Decontamination of shredded solid infectious and sharp waste by steam in a pure, saturated steam atmosphere. | YES/No
| Size of treatment chamber: 20, 50, 100 or 250 kg/h ±20%. Bidder to specify | YES/No
| Build in shredder in stainless steel with minimum 2 shafts with knoves / blades with automatic reverse in case of jam to render the into small pieces before it is treated. | YES/No
| The hybrid autoclaving units are designed to remain operational in a tropical environment or in proximity to saltwater for a minimum period of 10 years, provided that the specified maintenance schedule is followed. All metal surfaces are resistant to atmospheric and process-induced corrosion, using non-corrosive stainless steel (e.g., UNS: S31603 / S30400). Supplier to indicate used metal in accordance with AISI, UNS or equivalent norms. | YES/No |  | Stainless steel raw material technical data sheet (optional)
| Inactivation of vegetative bacteria, fungi, lipophilic/hydrophilic viruses, parasites and mycobacteria at a 6 Log10 reduction or greater; and inactivation of G. stearothermophilus spores and B. atrophaeus spores at a 4 Log10 reduction or greater (high level disinfection / STAATT Level III criteria). | YES/No |  | Validation result or testing results of the same model autoclave in former project, or testing results from an external accredited entity like Robert Koch Institute or Pasteur Institute or an accredited third party laboratory listed on https://ilac.org/about-ilac/
| Decontamination of waste with ≥121°C for ≥20 minutes or higher temperatures with reduced treatment time (e.g. 138 °C for 10 minutes). Supplier to indicate temperature and time. | YES/NO
| Automatic electric steam generator - capacity adapted to needs. | YES/NO
| Air compressor - capacity adapted to needs if applicable). | YES/NO
| Fully automatic steering of the treatment process. | YES/NO
| Including of a visual monitoring system of the whole treatment cycle (time, temperature, pressure) - clearly readable analogue or digital system operation panel (english or local language). | YES/NO
| The chamber and the steam supply are equipped with pressure gauges. | YES/NO
| Indicate the number of all available pre-programmed cycles with details and minimum cycle time (in minutes). | YES/NO
| Bidder to provide information on average water and electricty consumption at the rated capacity: liter / cycle and electricity: kW / cycle | YES/NO
| Equipment expected lifespan in years. | YES/NO
Safety Features | Equipped with a readily accessible emergency shut-off switchesfor immediate cycle shut-off. | YES/NO
| Included safety mechanism which prevent the cycle being started until the door is locked and sealed and to prevent the door being opened until chamber pressure is at atmospheric. | YES/NO
| Each pressure vessel (chamber etc.) equipped with pressure relief safety valves or equivalent pressure limiting device to prevent excess pressure. | YES/NO
| Automatic decontamination / sterilization of the shredder as well as the upper and lower chambers during the treatment cycle. | YES/NO
| Noise level of the autoclave during operation (dB). Bidder to indicate. | YES/NO
| Steam and condensate discharged from the chamber is cooled to reduce discharge temperature to less than 60 degrees C to avoid risk of damage to the building drainage system. | YES/NO
| Furnished with external surface insulation to mitigate the risk of elevated temperatures, ensuring that surfaces do not exceed 50°C. This precautionary measure is implemented to protect personnel. | YES/NO
Loading and unloading | Manual loading of waste / batch loading . | YES/NO
| Optional: Automatic waste loading system and 2 waste loading bins suitable to the automatic loading system. | YES/NO
| Gravity based / manual unloading in to a bin / cart, with adequate volume of the complete batch of waste of one cylce (indicate the volume of the unloading bin / cart) | YES/NO
Electrical supply, generation and wiring | Electricity supply: The waste autoclave is designed for 3 phase, 400 V a.c. ±10% or for country specific conditions. Information on electricity supply options available to be provided. | YES/NO
| All electrical connections attached with the hybrid autoclave and other machinery to be provided with suitable and necessary safety devices and be protected by steel conduits, cable trays, and supports. | YES/NO
Certificates and markings | Marking on the device indicating: (i) manufacturer’s name or trademark (ii) style, type, model or, date of manufacture of the device (iii) autoclave capacity (chamber volume / treated waste per hour) | YES/NO |  | Mock Label
| ISO 9001 Certificate - Quality management of the manufacturer or ISO 13485 for Medical Devices - Quality management | YES/NO |  | ISO 9001 or ISO13485 Certificate
| Compliant with EN 285:2006+A2:2009 for large sterilizer or an equivalent standard (provided by a certified independent third party). | YES/NO |  | Validation result or testing results of the same autoclave type in former project, or testing results from an external accredited entity for example Robert Koch Institute or Pasteur Institute, an accredited third party laboratory listed on https://ilac.org/about-ilac/
| Markings /certificate for pressure vessel safety in compliance with EN 13445 or ASME Boiler and Pressure Code or equivalent. | YES/NO |  | Pressure vessel Certificate
| Calibration certificates of the pressure and temperature measurement devices. | YES/NO |  | Calibration certificates
| IEC 60601 Electrical equipment for medical use or IEC 61010-2-040, UL 61010A-2-041, or equivalent electrical safety standard. | YES/NO |  | Declaration of Conformity (DoC)
Documentation | Operating and maintainance / service manual including a list of proecedures required for local routine maintenance in English to be translated into the local language. | YES/NO |  | Manual in English language
| Short user manual for the autoclave operator (Safety Operation Procedure) in English to be translated into the local language. | YES/NO |  | SOP in english language
| Set of engineering drawings: Lay out drawing with all main equipment (dimentions), main piping, PID flow diagram, installation requirements. | YES/NO |  | Engeneering drawings and installation requirements.
Warranty | Minimum 2 years warranty included in the price. Starting after the installation, testing and comissioning is finalized. | YES/NO
Packaging, labelling | Packaging, labelling and shipping must be compliant with national and international standards and regulations and in accordance with UNOPS packing and labelling instruction as indicated in the UNOPS GCC. | YES/NO
Bidder must be able to supply all the items listed below.
MANDATORY: Installation, testing and commissioning, training, spare parts package
| Factory acceptance test (FAT): the system, prior to shipment, will be tested for conformity of the system with the manufacturer’s performance specifications and the minimum requirements specified. The supplier / manufacturer is responsible to provide the FAT report. An examplary FAT report to be provided. | YES/NO
Installation, Testing and Commissioning requirements | All associated housing, fittings and connections to enable full functionality of the autoclave shall be completed by the client (local entity where the autoclave will be installed). These shall be in line with the manufacturer’s instructions; including but not limited to housing, foundation, piping, electrical and water connections, drains, ventilation, etc., as applicable. | YES/NO
| The supplier of the equipment is responsible for the installation and commissioning of the equipment. | YES/NO
| On-site acceptance test (SAT): the system, after delivery, will be tested by the contractor together with the user to demonstrate that the performance meets the manufacturer’s performance specifications. The waste decontamination efficiency of the autoclaves will be proofed after installation by adequate tests (series of biological, chemical, thermoelectric tests) under supervision of the client / an independent third party. The results of the testing of the system will be documented by the contractor in an acceptance protocol that is signed and dated by the user. | YES/NO
Operator training | Providing of onsite training for users and technicians: The training focus on providing hands on training to the client’s staff (operator, technicians) regarding the operation, testing, maintenance, trouble shooting and repair of the autoclave supplied under this Contract. This training session will be conducted at the client’s location, where the autoclave is installed. Followed by the training, the trainer supervise the operation of the autoclave by the client for at least 1 additional day. | YES/NO |  | Examplary training agenda / materials.
| Amount of days for necessary training . The trainer will be provided on a day-rate reimbursable basis. The bidder must specify the minimum scope of work in terms of operator working days . Rates for these to be provided on the financial offer form and MUST NOT include expenses of travel ticket and DSA - Daily Subsistence Alowance, such as hotel, food, transportation. These will be calculated according to UNOPS' standard travel and DSA fees. | YES/NO
Spare parts and post-installation service | Spare part package for 1 years to be provided. | YES/NO |  | List of spare parts with unit costs to be delivered.
| Declaration of spare parts availability post-warranty of minimum 5 years. | YES/NO
| Description of types and cost of service contract of 5 years. | YES/NO
| List of countries where local service can be made available and contact details of local / regional service agents . | YES/NO |  | List of countries/regions and contact details
OPTIONAL | Description of types and cost of service contract of 5 years | YES/NO
OPTIONAL: Accessories and Consumables
Assessories and consumables (optional - costs to be indicated separately and not included in the overall price of the quipment and above components) | Water softener and adequate consumables if applicable. | YES/NO
| Voltage limiter for electricity supply of the control system of the autoclave. | YES/NO
| Water meter (m³) | YES/NO
| Energy meter (kwh) | YES/NO
Providing of autoclave process testing equipment:
| 250 Chemical indicators test strips (Class 5 integrated Indicators): EN ISO 11140-1-4 | YES/NO
supplier to indicate shelf life of the above test strips
| 100 Biological spore test vials: Endospores of B. stearothermophilus tests (strains NCTC 10007, NCIB 8157, ATCC 7953) and an incubator (55-60 °C). | YES/NO
supplier to indicate shelf life of the above vials

## Items 9-12. Microwave with shre
3: Microwave with integrated shredder (treatment of solid infectious and sharp waste) |  | Compliant/ Non- Compliant | Insert your Specifications. Bidders are required to provide information on the specifications of their bid to demonstrate compliance to set specifications. Not providing specifications under this column may render your bid non-compliant. | Provide supporting documentation.
Minimum Requirements: |  |  |  | Technical data sheet
Supplier Name
Supplier Address
Manufacturer Name
Manufacturer Address
Mnufacturing Site Address
Model Number
| Sterilization of shredded solid infectious and sharp waste by microwaving (frequencies 2450±50 MHz or 915±25 MHz) - the shredder is integrated into the system and decontaminated during the process. | YES/NO
| Size of treatment chamber: 20 kg/h ±20% | YES/NO
| Build in shredder in stainless steel with minimum 2 shafts with rotative knoves / blades. | YES/NO
| Size of treatment chamber: 20, 50, 100 or 250 kg/h ±20%. Bidder to specify | YES/NO |  | Stainless steel raw material technical data sheet (optional)
| Inactivation of vegetative bacteria, fungi, lipophilic/hydrophilic viruses, parasites and mycobacteria at a 6 Log10 reduction or greater; and inactivation of G. stearothermophilus spores and B. atrophaeus spores at a 4 Log10 reduction or greater (high level disinfection / STAATT Level III criteria or greater). | YES/NO |  | Validation result or testing results of the same model microwave in former project, or testing results from an external accredited entity like Robert Koch Institute or Pasteur Institute or an accredited third party laboratory listed on https://ilac.org/about-ilac/
General requirements | Decontamination of waste with ≥100°C for ≥20 minutes or adequate | YES/NO
| Equipped with a filter or treatment system for potentially contaminated air from the microwave. | YES/NO
| Fully automatic steering of the treatment process. | YES/NO
| Including of a visual monitoring system of the whole treatment cycle - clearly readable analogue or digital system operation panel (time, temperature, pressure). | YES/NO
| Indicate the number of all available pre-programmed cycles with details and minimum cycle time (in minutes). | YES/NO
| Bidder to provide information on average water and electricty consumption at the rated capacity: liter / cycle and electricity: kW / cycle | YES/NO
| Equipment expected lifespan in years | YES/NO
Safety Features | Equipped with a readily accessible emergency shut-off switchesfor immediate cycle shut-off. | YES/NO
| A safety mechanism preventing: 1. the cycle from starting until the doors are locked and sealed. 2. the door being opened until the process is finalised. | YES/NO
| Air decontamination by a HEPA filter (0,2 µm, heat resistant, steam resistant, hydrophobic and pressure resistant for the pressure difference) or steam treatment or other equivalent method to prevent the release of pathogenic aerosols. | YES/NO
| Automatic decontamination / sterilization of the shredder as well as the upper and lower chambers during the treatment cycle. | YES/NO
| Noise level of the device during operation (dB). Bidder to indicate | YES/NO
| Steam and condensate discharged from the chamber is cooled to reduce discharge temperature to less than 60 degrees C. to avoid risk of damage to the building drainage system. | YES/NO
| Furnished with external surface insulation to mitigate the risk of elevated temperatures, ensuring that surfaces do not exceed 50°C. This precautionary measure is implemented to protect personnel. | YES/NO
Loading and unloading | Manual loading of waste / batch loading | YES/NO
| Optional: Automatic waste loading system and 2 waste loading bins suitable to the automatic loading system. | YES/NO
| Gravity based / manual unloading in to a bin /cart, with adequate volume of the complete batch of waste of one cylce (indicate the volume of the unloading bin / cart) | YES/NO
Electrical supply, generation and wiring | Electricity supply: The microwave is designed for 3 phase, 400 V a.c. ±10% or for country specific conditions. Information on electricity supply options available to be provided. | YES/NO
| All electrical connections attached with the device and other machinery provided with suitable and necessary safety devices and protected by steel conduits, cable trays, and supports. | YES/NO
Certificates and markings | Marking on the device indicating: (i) manufacturer’s name or trademark (ii) style, type, model or, date of manufacture of the device (iii) microwave capacity (chamber volume / treated waste per hour) | YES/NO |  | Mock Label
| ISO 9001 Certificate - Quality management of the manufacturer or ISO 13485 for Medical Devices - Quality management | YES/NO |  | ISO 9001 or ISO13485 Certificate
| Calibration certificates of the pressure and temperature measurement devices. | YES/NO |  | Calibration certificates
| IEC 60601 Electrical equipment for medical use or IEC 61010-2-040, UL 61010A-2-041, or equivalent electrical safety standard. | YES/NO |  | Declaration of Conformity (DoC)
Documentation | Operating and maintainance / service manual including a list of proecedures required for local routine maintenance in English to be translated into the local language. | YES/NO |  | Manual in English language
| Short user manual for the microwave operator (Safety Operation Procedure) in English to be translated into the local language. | YES/NO |  | SOP in english language
| Set of engineering drawings: Lay out drawing with all main equipment (dimentions), main piping, PID flow diagram, installation requirements. | YES/NO |  | Engeneering drawings and installation requirements.
Warrant | Minimum 2 years warranty included in the price. Starting after the installation, testing and comissioning is finalized. | YES/NO
Packaging, labelling | Packaging, labelling and shipping must be compliant with national and international standards and regulations and in accordance with UNOPS packing and labelling instruction as indicated in the UNOPS GCC. | YES/NO
Bidder must be able to supply all the items listed below.
MANDATORY: Installation, testing and commissioning, training, spare parts package
| Factory acceptance test (FAT): the system, prior to shipment, will be tested for conformity of the system with the manufacturer’s performance specifications and the minimum requirements specified. The supplier / manufacturer is responsible to provide the FAT report. An examplary FAT report to be provided. | YES/NO
Installation, Testing and Commissioning requirements | All associated housing, fittings and connections to enable full functionality of the microwave with shredder shall be completed by the client (local entity where the device will be installed). These shall be in line with the manufacturer’s instructions; including but not limited to housing, foundation, piping, electrical and water connections, drains, ventilation, etc., as applicable. | YES/NO
| The supplier of the equipment is responsible for the installation and commissioning of the equipment. | YES/NO
| On-site acceptance test (SAT): the system, after delivery, will be tested by the contractor together with the user to demonstrate that the performance meets the manufacturer’s performance specifications. The waste decontamination efficiency of the device will be proofed after installation by adequate tests (series of biological, chemical, thermoelectric tests) under supervision of the client / an independent third party. The results of the testing of the system will be documented by the contractor in an acceptance protocol that is signed and dated by the user. | YES/NO
Operator training | Providing of onsite training for users and technicians: The training focus on providing hands on training to the client’s staff (operator, technicians) regarding the operation, testing, maintenance, trouble shooting and repair of the equipment supplied under this Contract. This training session will be conducted at the client’s location, where the microwave is installed. Followed by the training, the trainer supervise the operation of the microwave by the client for at least 1 additional day. | YES/NO |  | Examplary training agenda / materials.
| Amount of days for necessary training . The trainer will be provided on a day-rate reimbursable basis. The bidder must specify the minimum scope of work in terms of operator working days . Rates for these to be provided on the financial offer form and MUST NOT include expenses of travel ticket and DSA - Daily Subsistence Alowance, such as hotel, food, transportation. These will be calculated according to UNOPS' standard travel and DSA fees. | YES/NO
Spare parts and post-installation service | Spare part package for 1 years to be provided. | YES/NO |  | List of spare parts with unit costs to be delivered.
| Declaration of spare parts availability post-warranty of minimum 5 years. | YES/NO
| Description of types and cost of service contract of 5 years. | YES/NO
| List of countries where local service can be made available and contact details of local / regional service agents . | YES/NO |  | List of countries/regions and contact details
OPTIONAL | Description of types and cost of service contract of 5 years | YES/NO
OPTIONAL: Accessories and Consumables
Assessories and consumables (optional - costs to be indicated separately and not included in the overall price of the quipment and above components) | Voltage limiter for electricity supply of the control system of the device. | YES/NO
| Water meter (m³) | YES/NO
| Energy meter (kwh) | YES/NO
Providing of process testing equipment:
| 250 Chemical indicators test stripes (Class 5 integrated Indicators): EN ISO 11140-1-4 | YES/NO
Supplier to indicate shelf life of the above test strips
| 100 Biological spore tests vials: Endospores of B. stearothermophilus tests (strains NCTC 10007, NCIB 8157, ATCC 7953) and an incubator (55-60 °C) | YES/NO
Shelf life and storage requirements of the bove tests