## Sheet1
SN | Item | Spec | Qun | UOM | Picture detail
1 | Midray 240 Fully automated Clinal Chemistry Analyzer | Reagent volume: 100~250¬ul, step by 0.5¬ ul Sample volume: 2~45¬ul, step by 0.1¬ul Reagent/Sample probe: Liquid level detection Probe cleaning: Automatic washing for interior and exterior Carry over < 0.05% Automatic sample dilution: Pre-dilution and post-dilution Internal bar code reader option Capable to communicate with LIS in bi-directional mode Operating temperature: 37℃±0.1℃ Calibration modes: Linear (one point, two points and multi-points) Control Rules: X-R, L-J, Westgard multi-rule, Cumulative sum Operation system: Windows 8 Power Supply: 200~240V, 50/60Hz, _1000VA Complete Computer System With Multi Control and Multi Sera Calibrator (Optional) With Printer | 1 | PC
2 | Benchtop centrifuge: Nuve Laboratory centrifuge NF 1200 | Maximum Speed: Swing-out Rotor: 4.100 rpm Angle Rotor: 14.000 rpm Maximum RCF: Swing-out Rotor: 3.082xg Angle Rotor: 18.405 xg Tube Capacity: Swing-out Rotor: 4x280 ml Angle Rotor: 6x100 ml Control System: N-PrimeTM Programmable Microprocessor Control Speed Set Range: 500-14.000 rpm Speed Set Step: 10 rpm Speed Control Accuracy: ± 20 rpm Timer Set Range 1-99 minutes and hold position Timer Set Step 1 minute Acceleration/Braking Rates 10/10 Motor Induction Motor Body and lid Epoxy-polyester powder coated steel Chamber Stainless Steel Power Supply 230 V / 50-60 Hz Description Swing-out Rotor: Set of 4 inserts 12x5 ml size tube (Capacity) | 1 | PC
3 | Heavy-duty printer, | Medium-duty printer: ImageRUNNER ADVANCE DX 719iFZ without CASSETTE FEEDING UNIT-AR1;Monochrome Laser Multifunction Core Functions Print, Copy, Scan, Send, Fax Processor 1.8 GHz Dual Core Processor Control Panel 10.1" TFT LCD WSVGA Color Touch-panel Memory Main CPU Side: 2 GB Image Processing CPU Side: 1 GB + 512 MB Solid State Drive 256 GB Interface Connection Network: 1000Base-T/100Base-TX/10Base-T,Optional: Wireless LAN Board-D1 Others, Standard: USB 2.0 x2 (Host), USB 3.0 x1 (Host), USB 2.0 x1 (Device),Optional: Copy Control Interface,Paper Capacity (LTR, 20 lb. Bond),Standard: 650 Sheets Maximum: 3,200 Sheets,Paper Sources (LTR, 20 lb. Bond) Standard: 550-sheet Paper Cassette, 100-sheet Stack Bypass,Print Specifications Print Resolution (dpi),600 x 600 Standard Page Description Languages, UFR II, PCL®6, Adobe®, PS® 3;Copy Specifications First-Copy-Out Time (LTR) As fast as 5 Seconds Copy Resolution (dpi) 600 x 600 Multiple Copies Up to 9,999,Scan Specifications Type,Single-pass Duplexing Automatic Document Feeder13 Document Feeder Paper Capacity Up to 100 Sheets (20 lb. Bond) Document Feeder Supported Media Sizes Legal, Letter, Statement, Send Specifications Destination, Standard: Email/Internet Fax (SMTP), SMB, FTP, WebDAV, Mail Box Optional IP Fax, Super G3 Fax Address Book LDAP (2,000)/Local (1,600)/Speed Dial (200) Scan Resolution (dpi) Push: Up to 600 x 600,Pull: Up to 600 x 600 | 1 | PC
4 | Barcode Printer and scanner:Intermec PC43d thermal printer | Max. label width: 118 mm Max. label length: 203 dpi-1227 mm Print Direction: Prints text, barcodes and graphics in all four directions Print Resolution 8 dot/mm (203 dpi) or Print width: 203 dpi – max 104 mm (4.1 in); 300 dpi – max 106 mm (4.2 in) Print Speed: Selectable up to 8 ips ( 203.2 mm/sec) for 203 dpi; 6 ips (152.4 mm/sec) for 300 dpi Core Diameter: Min 25.4 mm (1 in) Max 38.1 mm (1.5 in) Electrical: Operating Frequency -50 - 60 Hz; Operating Voltage: 100 - 240 V AC | 2 | PC
5 | Small office printers: HP Color Laserjet Pro M283fdw Wireless Multi-function Laser Printer | Print, Scan, Copy, Fax Wireless Printing &Scanning Print Speed: 22 ppm Print Resolution: 600 x 600 dpi 50-Page ADF for Scanning and Copiny 2.7" Color Touchscreen USB 2.0 & Ethernet Connectivity | 1 | PC
6 | Variable Single channel Micropipette | Finnpipett Variable Volume Pipettes, range 5-50 µl, Increments 1μL | 4 | PCS
7 | Variable Single channel Micropipette | Finnpipett Variable Volume Pipettes, range 20-200 µl, Increments 1μL | 3 | PCS
8 | Variable Single channel Micropipette | Finnpipett Variable Volume Pipettes, range 100-1000 µl, Increments 1μL | 3 | PCS