## MAIN
UNICEF SUPPLY DIVISION TENDER
TENDER DESCRIPTION
2.2 billion people around the globe experience vision impairment or blindness, this is an area which UNICEF has committed to do as part of the AT2030 programme. In collaboration with WHO, UNICEF will add different assistive technology product ranges into our global supply catalogue, in this case for vision products. The goal of this project is to increase availability and affordability to optimal vision products and services in low and mid resource countries. The aim of the tender is to establish Long Term Arrangements (LTAs) with relevant, qualified suppliers for delivery of vision products and accessories for UNICEF and WHO programmes and partners over the coming years.
GUIDELINES FOR COMPLETING THE SUBMISSION
1 | Bidders are not allowed to copy and paste UNICEF's generic specifications as answers. In addition general phrases such as: "Offered as specification", "Compliant", As requested", "Compliant, but slightly larger, 20 cm", etc. Such statements might result in a rejected bid.
2 | All specifications for which the bidder has confirmed compliance, needs to be verified by UNICEF in the official publicised documents such as brochures, datasheets, etc. There suppliers need to ensure that this submission has documents attached which will allow for this verification. Links to webpages are not acceptable for verification purposes. Where specifications cannot be verified in official publicised documents, the bidder is to provide a documents in which the bidder inequivalently confirms compliance with those specifications. This document should be signed by an authorized person from the bidders' company and attached to this submission. Where verification cannot be carried out, bidders risk that the concerned product will be declared uncompliant.
3 | Bidders are reminded that not answering mandatory requirements truthfully risks their bid being disqualified.
4 | Where compliance with ranges, or minimum of maximum values are requested, indicate compliance with Yes or No and, in the free text field, indicate the actual range, or minimum or maximum value for the specific parameter.

## Schedules and Scoring
Evaluation Framework
Section | Headers | Scoring | Weight | Points
A | Manufacter Identification | Pass/Fail |  | 0
B | Device Classification | Pass/Fail |  | 0
C | Device Classification and Market Clearance | Pass/Fail |  | 0
D | If the Device Contains Batteries of any Kind | Pass/Fail |  | 0
E | Compliance to Technical Standards | Pass/Fail & Weighted Score | 0.25 | 50
F | Operating Enviroment | Pass/Fail |  | 0
G | Technical Specifications/Characteristics | Pass/Fail & Weighted Score | 0.5 | 100
H | Required Accessories, Supplies and/or Consumables | Pass/Fail | 0.05 | 10
I | Warranty and Maintenance | Pass/Fail & Weighted Score | 0.05 | 10
J | Packing | Pass/Fail |  | 0
K | Labelling / Documentation / Characterization | Pass/Fail |  | 0
L | Training, Instalation, Utilisation and Decommissioning | Pass/Fail & Weighted Score | 0.05 | 10 | *Some products do not require scored technical specification in this section. Please check the Maximum Total Points in the table below.
M | Sustainability Considerations | Pass/Fail & Weighted Score | 0.1 | 20
| Maximum Total points (*See below the Maximum Total Points for each product) | 200
This tender has four schedules, one for each category of health intervention. It encompasses all essential products necessary to provided or achieve an eye care intervention. The schedule one is for eye and vision screening and it has eye charts for distance and near, occluder and basic ophthalmoscope-loupe-otoscope. Its purpose is to support eye care interventions community and primary health workers. The schedule two includes products for the provision of comprehensive eye examination and refractive services. It includes products for the screening or determination of objective and subjective refractive error and ruling our presence of eye disease. Its main purpose is to support eye care practitioners in refractive error interventions, and eye disease detection. The schedule three includes spectacles and low vision aids. It aims to support the provision of ready to made/use refractive solutions in low resource contexts, where there is significant access barriers to refractive and low vision products. These products can be dispensed by secondary and tertiary eye health facilities, and also in the community. The schedule four is for addressing blindness. It includes a single product, white cane, folding and non-folding. Proposals that includes all products in one schedule will be preferred. The schedules are arranged as follows, with the total point that can be achieve and minimal threshold points to pass in the technical evaluation.
ITEM NO. | Material Number | MATERIAL | SHORT DESCRIPTION | SCHEDULE | Total Points | Minimal Threshold Points
10 | S0004256 | Eyechart,farvision,HOTV,nilluminated | Printed eye chart for measuring distance visual acuity with random ordered HOTV optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
20 | S0004257 | Eyechart,nearvision,HOTV,nilluminated | Printed eye chart for measuring near visual acuity with random ordered HOTV optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
30 | S0004258 | Eyechart,farvision,LogMAR,nonillumi | Printed eye chart for measuring distance visual acuity with random ordered letters optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
40 | S0004259 | Eyechart,nearvision,LogMAR,nonillumi | Printed eye chart for measuring near visual acuity with random ordered letters optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
50 | S0004260 | Eyechart,farvision,LEA,nilluminated | Printed eye chart for measuring distance visual acuity with random ordered LEA optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
60 | S0004261 | Eyechart,nearvision,LEA,nonilluminated | Printed eye chart for measuring near visual acuity with random ordered LEA optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
70 | S0004262 | Eyechart,farvision,tumblingE,nonillumi | Printed eye chart for measuring distance visual acuity with random ordered Tumbling E optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
80 | S0004263 | Eyechart,nearvision,tumblingE,nonillumi | Printed eye chart for measuring near visual acuity with random ordered Tumbling E optotypes according to ISO 8596:2017. | 1: Eye and vision screening | 190 | 0
90 | S0004230 | Readingchart,nearvision,nonillumi​ | Printed eye chart for measuring near reading acuity as described in ISO/DIS 7921 Ophthalmic optics and instruments — Near vision charts. | 1: Eye and vision screening | 190 | 0
100 | S0004238 | Occluder and pinhole | There are two common configuration. The first is in a shape of a ruler with two circles, one in each end, with about 25 mm diameter of opaque material. One of the circles as a pinhole in its centre. The second is the shape of rectangular mask with a hole of about 25 mm in the place of one of the eyes with a slide covert with a pinhole in is centre, and with a side holding handle. The pinhole must have a diameter between 0.94mm to 1.75mm. Desirable it should be of 1.2mm. | 1: Eye and vision screening | 190 | 25
110 | S0004264 | Ophthalmoscope-Loupe-Otoscope​ | ​​The portable ophthalmoscope-loupe-otoscope is a small sized, solar-charged battery powered device with LED lights and lenses for the observation of the eye fundus, eye and hear adnexa, ear canal and outer ear drum. ​ | 1: Eye and vision screening | 200 | 0
120 | S0004253 | Autorefractor, portable and low cost | Automated refraction measurement device, desktop set-up, with aided measurement, average of the measures with quality assessment. Support for the head stability and fast operation. | 2: Comprehensive eye examination – Refractive services | 200 | 60
130 | S0004254 | Projector & visual acuity system | Projector or illuminated display with a set of eye tests designed for measurement, assessment and determination of visual acuity, spherical, cylinder power and axis of the refractive error, bi-ocular and binocular balance, simultaneous perception. | 2: Comprehensive eye examination – Refractive services | 200 | 0
140 | S0004255 | Cross cylinder | Cross cylinder of ±0.25DC, with a handle and power and axis alignment markings in different colours. Red for negative and blue/green for positive power. | 2: Comprehensive eye examination – Refractive services | 190 | 0
150 | S0004240 | Ophthalmic torchlight | A ophthalmic torchlight, also known as an ophthalmic flashlight, is a portable hand-held electric lamp used in the medical field. It’s similar to a regular flashlight or torch, but it’s often designed to be more compact and lightweight for ease of use in medical examinations. Healthcare professionals use it to examine areas of the body that are not easily visible, such as eyes. The light source can be a miniature incandescent light bulb or light-emitting diodes (LEDs). It can also be used to stimulated pupil afferent and efferent responses or as a target to pursuit with your eyes. | 2: Comprehensive eye examination – Refractive services | 190 | 0
160 | S0004241 | Photo screener, infant and child | Device that takes a picture of the child's eyes and analyses their alignment, focus, and reflexes. The photo screener can identify signs of amblyopia and other vision problems, such as strabismus, astigmatism, or cataracts. The photo screener is easy to use, fast, and non-invasive. It can be used by anyone, even without specialized training. These images are analysed by software incorporated into the equipment to evaluate the alignment of the eye and estimate refractive error. | 2: Comprehensive eye examination – Refractive services | 200 | 0
170 | S0004242 | Fundus camera, non-mydriatic | It consists in a portable non-mydriatic high quality photographic camera with a head mount for the patient. It also has a software and display for analysis. It is able to analyse the image according a set of indicators grouped in a report. It includes all features for standalone usage including but not limited to digital camera body, computer hardware, firmware and software, head mount, monitor, controllers for image capturing and viewing, casing and support structure. It also should have connectivity with external computer and/or data transfer. | 2: Comprehensive eye examination – Refractive services | 200 | 0
180 | S0004243 | Retinoscope and/or ophthalmoscope | The retinoscope and the ophthalmoscope are battery powered devices for the objective measurement of the refractive error and the observation of the eye fundus in a reliable, accurate and non-invasive technique, respectively. | 2: Comprehensive eye examination – Refractive services | 200 | 0
190 | S0004244 | Retinoscopy lens, rack | Ruler with a set of positive and negative lens ranging from 0.50 to 10 D in 0.50DS steps, colour green for positive and red for negative. | 2: Comprehensive eye examination – Refractive services | 190 | 0
200 | S0004248 | Trial frame, universal, adult | Eye frame with two interconnected lens holders and means for holding the different full aperture trial lens in the required position relative to the subject's eyes. It must have adjustable sides and bridge for adult range sizes, with soft padding. It can hold full or reduce aperture trial lenses. The lens holders should be adjustable to a range of adult interpupillary distance. | 2: Comprehensive eye examination – Refractive services | 200 | 0
210 | S0004249 | Trial frame, universal, paediatric | Paediatric, 2 to 12 years of age, eye frame with two interconnected lens holders and means for holding the different full aperture trial lens in the required position relative to the subject's eyes. It must have adjustable sides and bridge for infant/child range sizes, with soft padding. It can hold full or reduce aperture trial lenses. The lens holders should be adjustable to a range of infant/child interpupillary distance. | 2: Comprehensive eye examination – Refractive services | 190 | 0
220 | S0004250 | Trial lens set | Consists in a full or reduced aperture trial lens set in protective and sorted, case resistant to impact, with lock. It should contain sufficient lenses for quick and easy determination of the subjective refraction of any refractive error or binocular and accommodation examination. They are used addictively, by insertion and overlapping each other on a compatible trial lens frame. | 2: Comprehensive eye examination – Refractive services | 190 | 0
230 | S0004231 | Magnifier, handheld, illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | 3: Spectacles and optical low vision aids | 190 | 0
240 | S0004232 | Magnifier, handheld, non-illuminated | Magnifier, handheld, non-illuminated An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | 3: Spectacles and optical low vision aids | 190 | 0
250 | S0004233 | Magnifier, pocket | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). The pocket magnifier is sized to fit in a pocket, and as a folding case for storage, transport and protection. | 3: Spectacles and optical low vision aids | 190 | 0
260 | S0004234 | Magnifier, spectacle | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Scratch- and impact-resistant plus powered lenses mounted in a spectacle frame. Once the power is more than +4.00 D, it is generally referred to as spectacle magnifiers. Base-in prisms are available from +5.00 D to +12.00 D in spectacle magnifiers. The availability of the spectacle magnifiers (without base-in prisms) is up to +44.00 D. However, considering the extremely short working distance that a person will be working at, usually spectacle magnifiers up to +24.00D can be prescribed for comfortable near work. | 3: Spectacles and optical low vision aids | 190 | 0
270 | S0004235 | Magnifiers, bar, ruler | An bar-shaped optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens shaped has a half-cane bar with the purpose to be used on in contact with the surface of a text or print, aligned with the line to be read. It may have a handle for easy manipulation and placement. Also, it may a colour line across the width of the magnifier, to facilitate the reading of the text. | 3: Spectacles and optical low vision aids | 190 | 0
280 | S0004236 | Magnifiers, neck | Lightweight, lens type plastic, high grade plastic frame, adjustable cord. | 3: Spectacles and optical low vision aids | 190 | 0
290 | S0004237 | Magnifiers, stand, illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | 3: Spectacles and optical low vision aids | 190 | 0
300 | S0004239 | Magnifiers, stand, non-illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | 3: Spectacles and optical low vision aids | 190 | 0
310 | S0004245 | Spectacles, pre-made | Pre-made glasses comprise a set of lenses with a power in the range -/+1.00D to -/+3.50D, in 0.25D steps, pre-cut to the shape and side of a specific eye frame. Left and right lenses can be of different power and are single-vision lenses (distribute focus evenly across the lens surface). The power range can be extended to included cylinder powers in different axis. At minimum, pupillary distance >18 year of age adult male 63 mm, adult women 61 mm. Male 14-18y: 61 mm; Female 14-18y: 59 mm. 9-13y: 57 mm; 4-8y: 52 mm. | 3: Spectacles and optical low vision aids | 200 | 0
320 | S0004246 | Spectacles, ready-made, distance vision | Ready-made glasses for distance comprise lenses with a power in the range -/+1.00D to -/+3.50D, in 0.25D steps. Left and right lenses have the same power and are single-vision lenses (distribute focus evenly across the lens surface). Ready-made glasses for distance are available in different frame sizes and styles, with no adjustability or customizing. Pupillary distance >18 year of age adult male 63 mm, adult women 61 mm. Male 14-18y: 61 mm; Female 14-18y: 59 mm. 9-13y: 57 mm; 4-8y: 52 mm. | 3: Spectacles and optical low vision aids | 190 | 0
330 | S0004247 | Spectacles, ready-made, near vision | Reading glasses comprise lenses with a power in the range +1.00D to +3.50D. Left and right lenses have the same plus power and are single-vision lenses (distribute focus evenly across the lens surface). Reading glasses are available in different frame sizes and styles, with no adjustability or customizing. Pupillary distance adult male 59 mm, adult women 57 mm. | 3: Spectacles and optical low vision aids | 190 | 0
340 | S0004251 | White cane | The white cane is made of a long tube of aluminium tubing, fiberglass or carbon fibber with a handle on one end and a tip on the other. The handle may have a wrist loop on the end or a small crook, which can be used to store the white cane when not in use. The body of the white cane is covered with a white/red paint or reflective tape to provide visibility. Some white canes have a few inches of red paint or reflective tape at the tip indicating the user to have both vision and hearing loss. | 4: Blindness | 190 | 0
350 | S0004252 | White cane, folding | The folding white cane looks similar to a straight white cane except that it is broken up into several sections which are held together by an elastic cord running through the middle of the tubing, which allows it to be held tightly together when unfolded or folded away when not in use. | 4: Blindness | 190 | 0

## List of Materials
LIST OF EQUIPMENT INCLUDED IN THE TENDER
ITEM NO. | Material Number | MATERIAL | SHORT DESCRIPTION | MATERIAL DESCRIPTION | SCHEDULE | EU CLASSIFICATION | GMDN | Intended use
10 | S0004256 | Eyechart,farvision,HOTV,nilluminated | Printed eye chart for measuring distance visual acuity with random ordered HOTV optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines. Essential for central visual acuity measurement in all ages. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of HOTV optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement in all ages and all people, including illiterate people and pre-school children. It can used in comprehensive eye examination, but it is designed for use in eye screening by community and primary healthcare workers.
20 | S0004257 | Eyechart,nearvision,HOTV,nilluminated | Printed eye chart for measuring near visual acuity with random ordered HOTV optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for near visual acuity measurement, arrange in a progression of reduced sizes by lines. Essential for central visual acuity measurement in all ages. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of HOTV optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement in all ages and all people, including illiterate people and pre-school children. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
30 | S0004258 | Eyechart,farvision,LogMAR,nonillumi | Printed eye chart for measuring distance visual acuity with random ordered letters optotypes according to ISO 8596:2017. | ​Charts with a set of optotypes designed for distance visual acuity measurement. Essential for central visual acuity measurement in all ages with Log Mar VA progression from the +1.3 to 0.0 Log Mar in Log Mar notation. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of letter optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes with Log MAR progression to be viewed at a standardize distance. It is used for central visual acuity measurement in literate people. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
40 | S0004259 | Eyechart,nearvision,LogMAR,nonillumi | Printed eye chart for measuring near visual acuity with random ordered letters optotypes according to ISO 8596:2017. | ​Charts with a set of optotypes designed for near visual acuity measurement. Essential for central visual acuity measurement in all ages with Log Mar VA progression from the +1.3 to 0.0 Log Mar in Log Mar notation. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of letter optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes with Log MAR progression to be viewed at a standardize distance. It is used for central visual acuity measurement in literate people. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
50 | S0004260 | Eyechart,farvision,LEA,nilluminated | Printed eye chart for measuring distance visual acuity with random ordered LEA optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines.. Essential for central visual acuity measurement in children and non-literate people. This eye chart uses the set of LEA symbols. It requires a card with a set of LEA symbols sized for near VA 0.05, for patient identification of viewed symbol. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of LEA optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement specially designed for pre-school children. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
60 | S0004261 | Eyechart,nearvision,LEA,nonilluminated | Printed eye chart for measuring near visual acuity with random ordered LEA optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for near visual acuity measurement, arrange in a progression of reduced sizes by lines.. Essential for central visual acuity measurement in children and non-literate people. This eye chart uses the set of LEA symbols. It requires a card with a set of LEA symbols sized for near VA 0.05, for patient identification of viewed symbol. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of LEA optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement specially designed for pre-school children. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
70 | S0004262 | Eyechart,farvision,tumblingE,nonillumi | Printed eye chart for measuring distance visual acuity with random ordered Tumbling E optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for distance visual acuity measurement. Essential for central visual acuity measurement in children and non literate people. This eye chart uses the set of the E arranged randomly in up, down, left, right direction, from the VA of 0.1 to 1.0, in decimal notation. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of Tumbling E optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement specially designed for pre-school children and illiterate people. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
80 | S0004263 | Eyechart,nearvision,tumblingE,nonillumi | Printed eye chart for measuring near visual acuity with random ordered Tumbling E optotypes according to ISO 8596:2017. | Chart with a set of optotypes designed for near visual acuity measurement. Essential for central visual acuity measurement in children and non-literate people. This eye chart uses the set of the E arranged randomly in up, down, left, right direction, from the VA of 0.1 to 1.0, in decimal notation. | 1: Eye and vision screening | Medical Device Class I | 16800 | Non illuminated chart with a set of Tumbling E optotypes designed for distance visual acuity measurement, arrange in lines of descending sizes to be viewed at a standardize distance. It is used for central visual acuity measurement specially designed for pre-school children and illiterate people. It can used in comprehensive eye examination, but it is designed for eye screening by community and primary healthcare workers.
90 | S0004230 | Readingchart,nearvision,nonillumi​ | Printed eye chart for measuring near reading acuity as described in ISO/DIS 7921 Ophthalmic optics and instruments — Near vision charts. | Standardized charts for the assessment of near reading acuity, demonstrating whether or not a patient can read sentences or paragraphs of text in a particular font size. Essential for near reading acuity measurement in literate people. It uses a set of paragraphs of an easy to understandable text, arranged in a progressive reduction in the font size with adequate notation. The notation used can be log RAD, M size, N size, decimal reading acuity, reduced Snellen fraction. | 1: Eye and vision screening | Medical Device Class I | 16800 | Standardized chart for the assessment of near reading acuity, demonstrating whether or not a patient can read sentences or paragraphs of text in a particular font size, at a standardize distance. Essential for near reading acuity measurement in literate people. It uses a set of paragraphs of easy to understand text, arranged in a progressive reduction in the font size with adequate notation. The notation used can be log RAD, M size, N size, decimal reading acuity, reduced Snellen fraction. It can be used by community and primary health workers in eye screening or as part of an comprehensive eye examination.
100 | S0004238 | Occluder and pinhole | There are two common configuration. The first is in a shape of a ruler with two circles, one in each end, with about 25 mm diameter of opaque material. One of the circles as a pinhole in its centre. The second is the shape of rectangular mask with a hole of about 25 mm in the place of one of the eyes with a slide covert with a pinhole in is centre, and with a side holding handle. The pinhole must have a diameter between 0.94mm to 1.75mm. Desirable it should be of 1.2mm. | Handheld occluder on one end with a pinhole on the other end. Essential in the vision screening and subjective refraction. It can be used to occlude one eye and determine the visual acuity of the other. The pinhole can be used to determine if the visual acuity can improve when the eye chart optotypes are viewed through it. This allows the ruling out refractive error as the cause of visual impairment. | 1: Eye and vision screening | Medical Device Class I | 37337 | An occluder and a pinhole are devices used to test the visual acuity and alignment of a person's eyes. An occluder is a flat, opaque object that covers one eye. The occluder can also be used to test and diagnose the eyes' misalignment by cover and uncover them. The pinhole is a small hole in the centre of an occluder, that allows only a narrow beam of light to enter the eye. By using a pinhole, the examiner can eliminate the defocus of refractive errors and determine how much can the visual acuity improve in case of refractive error correction. This method can help screen for the amblyopia, refractive error and diagnose for binocular conditions. It can be used in eye screenings with the eye charts, by a community or primary healthcare personnel. Also can be used in comprehensive eye care examination by eye care practitioner, optometrist or ophthalmologist.
110 | S0004264 | Ophthalmoscope-Loupe-Otoscope​ | ​​The portable ophthalmoscope-loupe-otoscope is a small sized, solar-charged battery powered device with LED lights and lenses for the observation of the eye fundus, eye and hear adnexa, ear canal and outer ear drum. ​ | Designed for the direct observation of the eye fundus, ear canal and outer drum. Also allows for magnified and illuminated observation of eye and ear adnexa, with white and blue light. It consists in a portable, pocket-sized single unit with a different built-in LED light sources of white and blue colour directed through the pupil to illuminate the interior of the eyeball, dial containing lenses of different powers. It is chargeable through its own solar panel or USB-C. If has a speculum for ear canal and outer eardrum observation, a video clip for phone attachment and imaging recording, a securing lanyard, cloth, visual acuity and reading chart, and storage and transportation pouche. | 1: Eye and vision screening | Medical Device Class I | 46786
120 | S0004253 | Autorefractor, portable and low cost | Automated refraction measurement device, desktop set-up, with aided measurement, average of the measures with quality assessment. Support for the head stability and fast operation. | Portable device for automatic objective measurement of the refractive error in children and adults in a fast, simples and non-invasive technique, as alternative to retinoscopy. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 36386 | Autorefractor are devices for the automatic measuring of refractive error as part of a comprehensive eye examination, in all ages, by eye care practitioners as ophthalmologist and optometrists. Can also be used by allied ophthalmic personnel. It consists in an image projecting and focusing system into the eye fundus, using non visible wavelength. It also has a visible image with variable focus for patient viewing, to relax accommodation. The amount of focus required to obtain a clear image projected in the eye fundus, in the different meridians is the refractive error value.
130 | S0004254 | Projector & visual acuity system | Projector or illuminated display with a set of eye tests designed for measurement, assessment and determination of visual acuity, spherical, cylinder power and axis of the refractive error, bi-ocular and binocular balance, simultaneous perception. | Projector of different eye tests with the purpose to perform a comprehensive assessment and determination of the subjective refraction. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 35922 | Projector and visual acuity systems are devices for the evaluation and determination of subjective refractive error, biocular and binocular vision, among other evaluations of the visual function, as part of a comprehensive eye examination, in all ages, by eye care practitioners as ophthalmologist and optometrists. Can also be used by allied ophthalmic personnel under supervision of an eye care practitioner. It has different eye charts for visual acuity for all ages and literacy levels, cylinder power and axis determination, simultaneous perception, fusion, stereopsis, phoria presence and its prismatic dioptric power. It has a support for installation near and above the patient's head, a projecting plate to be hang in the front wall and a remote controller for changing the different tests.
140 | S0004255 | Cross cylinder | Cross cylinder of ±0.25DC, with a handle and power and axis alignment markings in different colours. Red for negative and blue/green for positive power. | Useful for cylinder power and axis determination in the subjective refraction. Also known as Jackson's cross cylinder. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 0 | Used to determine the cylinder power and axis in cooperant patients, as part of and comprehensive eye examination, performed by an eye care practitioner. It is a lens with a crossed cylinder of ±0.25DC with a handle for easy manipulation in front of patient's eye. It has three markings for identification of each power axis and its 45 degree axis that is aligned with the handle.
150 | S0004240 | Ophthalmic torchlight | A ophthalmic torchlight, also known as an ophthalmic flashlight, is a portable hand-held electric lamp used in the medical field. It’s similar to a regular flashlight or torch, but it’s often designed to be more compact and lightweight for ease of use in medical examinations. Healthcare professionals use it to examine areas of the body that are not easily visible, such as eyes. The light source can be a miniature incandescent light bulb or light-emitting diodes (LEDs). It can also be used to stimulated pupil afferent and efferent responses or as a target to pursuit with your eyes. | Torchlight in shape of a pen, with a click on-off switch. Useful for external eye examination, eye adnexa, ocular motility target and pupillary function. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 38832 | A torchlight is a small sized, handheld, battery powered device that produces a bright beam of light when switched on. It can be used for various purposes, such evaluating the exterior of the eye and adnexa, and the pupillary function. A torchlight works by converting electrical energy from batteries into light energy from a bulb or an LED. The light is then focused by a lens to form a narrow beam. It is used by all eye care workforce in all levels.
160 | S0004241 | Photo screener, infant and child | Device that takes a picture of the child's eyes and analyses their alignment, focus, and reflexes. The photo screener can identify signs of amblyopia and other vision problems, such as strabismus, astigmatism, or cataracts. The photo screener is easy to use, fast, and non-invasive. It can be used by anyone, even without specialized training. These images are analysed by software incorporated into the equipment to evaluate the alignment of the eye and estimate refractive error. | A photo screener is a handheld, portable device used for vision screening, particularly in children. It uses a camera to capture images of a child’s non-dilated eyes. Allows for screening of the presence of refractive error, ocular axis alignment and amblyopia in a fast, simples and non-invasive technique. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 0 | A photo screener is a device for fast and simple eye screening in young children. It takes a picture of the child's eyes and analyses the reflection of light from the retina. The photo screener can detect common eye disorders such as amblyopia, strabismus, refractive errors, and cataracts. A photo screener is intended to be used by trained personnel as part of a routine vision screening program.
170 | S0004242 | Fundus camera, non-mydriatic | It consists in a portable non-mydriatic high quality photographic camera with a head mount for the patient. It also has a software and display for analysis. It is able to analyse the image according a set of indicators grouped in a report. It includes all features for standalone usage including but not limited to digital camera body, computer hardware, firmware and software, head mount, monitor, controllers for image capturing and viewing, casing and support structure. It also should have connectivity with external computer and/or data transfer. | A fundus camera is a specialized imaging device for the observation of the eye fundus. This includes the retina, retinal vasculature, optic disc, macula, and posterior pole. The main structures that can be visualized on a fundus photo are the central and peripheral retina, optic disc, and macula. Fundus photography can be performed with coloured filters, or with specialized dyes including fluorescein and indocyanine green. Visualization of the anterior segment is a desirable feature. | 2: Comprehensive eye examination – Refractive services | Medical Device Class IIa | 10551 | A fundus camera is a device that captures images of the retina. It uses a low-power microscope and a flash to illuminate the retina and create a digital or film record of its condition. The images can help diagnose and monitor various eye diseases, such as glaucoma, diabetic retinopathy, macular degeneration, and more. It can be used for diabetic retinopathy screening in primary eye care, by trained personnel, and in secondary and tertiary health care, supporting comprehensive eye care.
180 | S0004243 | Retinoscope and/or ophthalmoscope | The retinoscope and the ophthalmoscope are battery powered devices for the objective measurement of the refractive error and the observation of the eye fundus in a reliable, accurate and non-invasive technique, respectively. | Allows for objective measurement of the refractive error and the observation of the eye fundus in a reliable, accurate and non-invasive technique. Hand-held streak retinoscopes are electrically-powered, hand-held, ophthalmic instrument designed to be held close to the examiner's eye, to project a streak of light into the patient's eye and observe the shape and motion of the streak reflex from the interior of the eye. It consists of a main unit (commonly referred to as the head) with a built-in light source directed through the pupil to illuminate the interior of the eyeball, a plane mirror with a single hole for viewing, and a moving lever for changing the focus of the streak; a dedicated power supply handpiece may be included with the head. Hand-held direct ophthalmoscopes are electrically-powered, hand-held, ophthalmic instrument designed to be held close to the patient’s eye to examine the interior of the eye and related structures [e.g., fundus (retina), cornea, aqueous, lens, and vitreous] by producing an upright (unreversed) image of approximately 5 times magnification. It consists of a main unit (commonly referred to as the head) with a built-in light source directed through the pupil to illuminate the interior of the eyeball, a mirror with a single hole for viewing, and a dial containing lenses of different powers; a dedicated power supply handpiece may be included with the head. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 46786 | Allows for objective measurement of the refractive error and the observation of the eye fundus in a reliable, accurate and non-invasive technique. Hand-held streak retinoscopes are electrically-powered, hand-held, ophthalmic instrument designed to be held close to the examiner's eye, to project a streak of light into the patient's eye and observe the shape and motion of the streak reflex from the interior of the eye. It consists of a main unit (commonly referred to as the head) with a built-in light source directed through the pupil to illuminate the interior of the eyeball, a plane mirror with a single hole for viewing, and a moving lever for changing the focus of the streak; a dedicated power supply handpiece may be included with the head. Hand-held direct ophthalmoscopes are electrically-powered, hand-held, ophthalmic instrument designed to be held close to the patient’s eye to examine the interior of the eye and related structures [e.g., fundus (retina), cornea, aqueous, lens, and vitreous] by producing an upright (unreversed) image of approximately 5 times magnification. It consists of a main unit (commonly referred to as the head) with a built-in light source directed through the pupil to illuminate the interior of the eyeball, a mirror with a single hole for viewing, and a dial containing lenses of different powers; a dedicated power supply handpiece may be included with the head.
190 | S0004244 | Retinoscopy lens, rack | Ruler with a set of positive and negative lens ranging from 0.50 to 10 D in 0.50DS steps, colour green for positive and red for negative. | Ruler with a set of positive and negative lens ranging from 0.50 to 10DS in 0.50DS steps, colour green for positive and red for negative. Allows for objective measurement of the refractive error in a reliable, accurate and non-invasive technique. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 0 | A crossed cylinder is a device used in comprehensive eye care examination to determine the subjective astigmatism of the eye. It consists of a lens with two cylindrical powers but opposite sign, placed at right angles to each other, with an handheld for easy manipulation. The crossed cylinder is rotated in front of the eye, in specific procedures, until the patient sees the clearest image. It is used by eye care practitioners, optometrist or ophthalmologist.
200 | S0004248 | Trial frame, universal, adult | Eye frame with two interconnected lens holders and means for holding the different full aperture trial lens in the required position relative to the subject's eyes. It must have adjustable sides and bridge for adult range sizes, with soft padding. It can hold full or reduce aperture trial lenses. The lens holders should be adjustable to a range of adult interpupillary distance. | Trial frame with adjustment to the face, for insertion and holding of trial set lenses in from of adult's eyes. Essential for the determination of the subjective refraction. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 0 | A universal trial frame is an adjustable eye frame used to hold different trial lenses in front of a patient's eyes. It consists of a metal or plastic eye frame that can hold up several trial lens and filters on each side, and can be adjusted for the pupil distance, nose bridge height, and temple length of the patient. The lenses are inserted into slots on the frame, and can be rotated or flipped to change their position. The universal trial frame allows the optometrist or ophthalmologist to find the best combination of lenses for the patient's refractive correction or to test visual system function.
210 | S0004249 | Trial frame, universal, paediatric | Paediatric, 2 to 12 years of age, eye frame with two interconnected lens holders and means for holding the different full aperture trial lens in the required position relative to the subject's eyes. It must have adjustable sides and bridge for infant/child range sizes, with soft padding. It can hold full or reduce aperture trial lenses. The lens holders should be adjustable to a range of infant/child interpupillary distance. | Trial frame with adjustment to the face, for insertion and holding of trial set lenses in from of child's eyes. Essential for the determination of the subjective refraction. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 0 | A paediatric universal trial frame is an adjustable eye frame used to hold different trial lenses in front of a children's eyes. It consists of a metal or plastic eye frame that can hold up several trial lens and filters on each side, and can be adjusted for the children's pupil distance, nose bridge height, and temple length of the patient. The lenses are inserted into slots on the frame, and can be rotated or flipped to change their position. The paediatric universal trial frame allows the optometrist or ophthalmologist to find the best combination of lenses for the patient's refractive correction or to test visual system function.
220 | S0004250 | Trial lens set | Consists in a full or reduced aperture trial lens set in protective and sorted, case resistant to impact, with lock. It should contain sufficient lenses for quick and easy determination of the subjective refraction of any refractive error or binocular and accommodation examination. They are used addictively, by insertion and overlapping each other on a compatible trial lens frame. | A set of trial lens to be placed in front of the subject's eye for the determination of the subjective refraction. | 2: Comprehensive eye examination – Refractive services | Medical Device Class I | 36068 | A trial lens set is a collection of lenses of different powers and filters used to determine the subjective refraction and evaluate the visual function of one person, as part of a comprehensive eye examination. A trial lens set is used by placing the lenses and filters in a trial frame or holder and adjusting them in front of the person's eyes. The lenses and filters can be changed quickly and easily. It is used by eye care practitioners, optometrist or ophthalmologist.
230 | S0004231 | Magnifier, handheld, illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 30045 | A handheld illuminated magnifier is an optical device that enlarges the appearance of objects for people who experience vision impairment or low vision. It consists of a lens that bends the light rays coming from the object, making it look bigger and clearer. A handheld magnifier can be used for various purposes, such as reading, writing, sewing, or inspecting small details. To use a handheld magnifier, hold it close to your eye and move it over the object until you find the best focus and magnification level.
240 | S0004232 | Magnifier, handheld, non-illuminated | Magnifier, handheld, non-illuminated An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 30045 | A handheld non illuminated magnifier is an optical device that enlarges the appearance of objects for people who experience vision impairment or low vision. It consists of a lens that bends the light rays coming from the object, making it look bigger and clearer. A handheld magnifier can be used for various purposes, such as reading, writing, sewing, or inspecting small details. To use a handheld magnifier, hold it close to your eye and move it over the object until you find the best focus and magnification level.
250 | S0004233 | Magnifier, pocket | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). The pocket magnifier is sized to fit in a pocket, and as a folding case for storage, transport and protection. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 30045 | A pocket magnifier is a small optical device that can be used to enlarge the appearance of objects or texts, designed to fit um a pocket. It works by using a convex lens that bends the light rays and creates a magnified image on the retina of the eye. Pocket magnifiers are useful for people who have vision impairment or low vision.
260 | S0004234 | Magnifier, spectacle | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Scratch- and impact-resistant plus powered lenses mounted in a spectacle frame. Once the power is more than +4.00 D, it is generally referred to as spectacle magnifiers. Base-in prisms are available from +5.00 D to +12.00 D in spectacle magnifiers. The availability of the spectacle magnifiers (without base-in prisms) is up to +44.00 D. However, considering the extremely short working distance that a person will be working at, usually spectacle magnifiers up to +24.00D can be prescribed for comfortable near work. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 42832 | A spectacle magnifier is an optical device that attaches to eyeglasses and provides magnification for near vision tasks. It works by increasing the size of the image on the retina, allowing the user to see more details. Spectacle magnifiers are intended for people with low vision who need to perform activities such as reading, writing, sewing, or crafts. They are to be used by someone that experiences vision impairment or low vision.
270 | S0004235 | Magnifiers, bar, ruler | An bar-shaped optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens shaped has a half-cane bar with the purpose to be used on in contact with the surface of a text or print, aligned with the line to be read. It may have a handle for easy manipulation and placement. Also, it may a colour line across the width of the magnifier, to facilitate the reading of the text. | Bar-shaped optical magnifiers can produce magnified (enlarged) height of the images of text and print, without changing the width. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 15652 | A bar or ruler magnifier is an optical device shaped as an half cane magnifying lens to help read lines of small or fine print. It works by placing the bar magnifier over the text or object and sliding it along the surface. A bar ruler magnifier is useful for reading maps, contracts, newspapers, labels, and other documents with small or fine print. It is used by people experiencing vision impairment or low vision.
280 | S0004236 | Magnifiers, neck | Lightweight, lens type plastic, high grade plastic frame, adjustable cord. | Non-illuminated, magnifier mounted in a frame, suspended by a cord around a person's neck and usually stabilized against their chest. Used for fine motor tasks. | 3: Spectacles and optical low vision aids | Medical Device Class I | 15652 | A non-illuminated neck magnifier is an handless optical device that can be worn around the neck to magnify small objects or texts. It is useful for people who have difficulty seeing or reading fine details, such as hobbyists, crafters, or seniors. The device consists of a large lens attached to a cord that can be adjusted to fit the user's neck. The lens can be moved up and down to change the focal distance and the magnification level.
290 | S0004237 | Magnifiers, stand, illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 15652 | A stand magnifier is an handless optical device that helps people with vision impairment or low vision to read printed materials or see details on small objects. It consists of a lens mounted on a base that rests on the page. The lens magnifies the text and the base provides a fixed focal distance. A stand magnifier can also have a built-in light source to improve contrast and clarity.
300 | S0004239 | Magnifiers, stand, non-illuminated | An optical magnifier uses positive power lens to produce a magnifying effect. It comprises a single lens or more than one lens (lens system). Most lenses or lens systems are mounted into a frame, with the exception of a dome magnifier. Different frame types suit different tasks. The range of magnifying power (magnification) is measured in dioptres (D). Magnifiers most commonly have a lens power range from +4D to +76D. Some magnifiers include a built-in light source and are described as illuminated magnifiers. | Optical magnifiers can produce magnified (enlarged) images of close objects and print. They are intended for use by children and adults with low vision that cannot be fully corrected with spectacles or other treatment such as surgery. | 3: Spectacles and optical low vision aids | Medical Device Class I | 15652 | A stand magnifier is an handless optical device that helps people with vision impairment or low vision to read printed materials or see details on small objects. It consists of a lens mounted on a base that rests on the page. The lens magnifies the text and the base provides a fixed focal distance.
310 | S0004245 | Spectacles, pre-made | Pre-made glasses comprise a set of lenses with a power in the range -/+1.00D to -/+3.50D, in 0.25D steps, pre-cut to the shape and side of a specific eye frame. Left and right lenses can be of different power and are single-vision lenses (distribute focus evenly across the lens surface). The power range can be extended to included cylinder powers in different axis. At minimum, pupillary distance >18 year of age adult male 63 mm, adult women 61 mm. Male 14-18y: 61 mm; Female 14-18y: 59 mm. 9-13y: 57 mm; 4-8y: 52 mm. | Pre-made spectacles for distance vision are plus or minus-powered lenses assemble into a spectacle frame. They consist in a set of different power lenses precut to the shape of a specific eye frame, for assemblage accordingly the refractive error for each eye in a specific person. They are intended for use by children and adults with myopia, hyperopia, presbyopia with or without astigmatism. | 3: Spectacles and optical low vision aids | Medical Device Class I | 35065 | Pre-made spectacles are eye frames and pre-shaped lenses that are pre-manufactured and can be assembled with a range of standard powers and dispensed over the counter. They are intended for people who have hypermetropia or myopia, a condition that causes difficulty in focusing on distant objects. Pre-made spectacles work by diverging, negative powers for myopia, or converging, positive powers for hypermetropia, and focusing on the retina the beam of light which helps to improve the clarity of vision. They have the ability of assembling a different power for each eye. However, they are not customized to the individual's refractive error, so they may not correct other vision conditions such as astigmatism.
320 | S0004246 | Spectacles, ready-made, distance vision | Ready-made glasses for distance comprise lenses with a power in the range -/+1.00D to -/+3.50D, in 0.25D steps. Left and right lenses have the same power and are single-vision lenses (distribute focus evenly across the lens surface). Ready-made glasses for distance are available in different frame sizes and styles, with no adjustability or customizing. Pupillary distance >18 year of age adult male 63 mm, adult women 61 mm. Male 14-18y: 61 mm; Female 14-18y: 59 mm. 9-13y: 57 mm; 4-8y: 52 mm. | Ready-to-wear spectacles for distance vision are plus or minus-powered lenses mounted into a spectacle frame. They are commonly known as reading glasses. They are intended for use by children and adults with myopia and hyperopia. | 3: Spectacles and optical low vision aids | Medical Device Class I | 35065 | Spectacles ready-made for distance vision are eye glasses that are pre-manufactured or assembled with a range of standard powers and dispensed over the counter. They are intended for people who have hypermetropia or myopia, a condition that causes difficulty in focusing on distant objects. Spectacles ready-made for distance vision work by diverging, negative powers for myopia, or converging, positive powers for hypermetropia, and focusing on the retina the beam of light which helps to improve the clarity of vision. However, they are not customized to the individual's refractive error, so they may not correct other vision conditions such as astigmatism.
330 | S0004247 | Spectacles, ready-made, near vision | Reading glasses comprise lenses with a power in the range +1.00D to +3.50D. Left and right lenses have the same plus power and are single-vision lenses (distribute focus evenly across the lens surface). Reading glasses are available in different frame sizes and styles, with no adjustability or customizing. Pupillary distance adult male 59 mm, adult women 57 mm. | Ready-to-wear spectacles for near vision are plus-powered lenses mounted into a spectacle frame. They are commonly known as reading glasses. They are intended for use by adults with presbyopia, helping them to visually focus on near tasks such as reading, writing, sewing, or cooking. | 3: Spectacles and optical low vision aids | Medical Device Class I | 35065 | Spectacles ready-made for near vision are eye glasses that are pre-manufacter or assembled with a range of standard powers and dispensed over the counter. They are intended for people who have presbyopia, a condition that causes difficulty in focusing on near objects, after forty years of age. Spectacles ready-made for near vision work by converging and focusing on the retina, the beam of light which helps to improve the clarity of vision. However, they are not customized to the individual's refractive error, so they may not correct other vision conditions such as astigmatism or myopia.
340 | S0004251 | White cane | The white cane is made of a long tube of aluminium tubing, fiberglass or carbon fibber with a handle on one end and a tip on the other. The handle may have a wrist loop on the end or a small crook, which can be used to store the white cane when not in use. The body of the white cane is covered with a white/red paint or reflective tape to provide visibility. Some white canes have a few inches of red paint or reflective tape at the tip indicating the user to have both vision and hearing loss. | A white cane is a long rod-like device used by blind or visually impaired travellers to give them Information about the environment they are traveling through. Using a cane can warn them of obstacles in their path, tell them of stairs they are coming to, warn them that they are coming up to a curb, and tell them of many other things in the environment that they should deal with. The cane will also alert others to identify them as blind and visually impaired thus needful help can be obtained. | 4: Blindness | Medical Device Class I | 0 | A white cane is a device that helps people who are blind or visually impaired to navigate their surroundings. It is usually held in front of the body and swept from side to side to detect obstacles, changes in elevation, or other features of the environment.
350 | S0004252 | White cane, folding | The folding white cane looks similar to a straight white cane except that it is broken up into several sections which are held together by an elastic cord running through the middle of the tubing, which allows it to be held tightly together when unfolded or folded away when not in use. | A white cane is a long rod-like device used by blind or visually impaired travellers to give them Information about the environment they are traveling through. Using a cane can warn them of obstacles in their path, tell them of stairs they are coming to, warn them that they are coming up to a curb, and tell them of many other things in the environment that they should deal with. The cane will also alert others to identify them as blind and visually impaired thus needful help can be obtained. | 4: Blindness | Medical Device Class I | 0 | A folding white cane is a device that can be folded to be easily transported or stored and helps people who are blind or visually impaired to navigate their surroundings. It is usually held in front of the body and swept from side to side to detect obstacles, changes in elevation, or other features of the environment.

## 10
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,farvision,HOTV,nilluminated
MATERIAL: |  | S0004256
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses HOTV optotypes. | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 20
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,nearvision,HOTV,nilluminated
MATERIAL: |  | S0004257
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses HOTV optotypes | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 30
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,farvision,LogMAR,nonillumi
MATERIAL: |  | S0004258
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses Log MAR notation and progression | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 40
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,nearvision,LogMAR,nonillumi
MATERIAL: |  | S0004259
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses Log MAR notation and progression | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 50
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,farvision,LEA,nilluminated
MATERIAL: |  | S0004260
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses LEA symbols optotypes | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 60
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,nearvision,LEA,nonilluminated
MATERIAL: |  | S0004261
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses LEA symbols optotypes | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 70
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,farvision,tumblingE,nonillumi
MATERIAL: |  | S0004262
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses tumbling E optotypes | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 80
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Eyechart,nearvision,tumblingE,nonillumi
MATERIAL: |  | S0004263
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
26 |  | Uses tumbling E optotypes | Pass/Fail | Yes |  | No
27 |  | Non-illuminated | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 90
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Readingchart,nearvision,nonillumi​
MATERIAL: |  | S0004230
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
19 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
20 |  | ISO/DIS 7921 Ophthalmic optics and instruments — Near vision charts | Pass/Fail | Yes |  | No
21 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
22 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
23 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
24 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
25 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
26 |  | Chart with a set of optotypes designed for distance visual acuity measurement, arrange in a progression of reduced sizes by lines with a minimum range of 0.1 to 1.25 decimal visual acuity. | Pass/Fail | Yes |  | No
27 |  | Uses a set of paragraphs with decreasing font size as reading target | Pass/Fail | Yes |  | No
28 |  | Non-illuminated | Pass/Fail | Yes |  | No
29 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
30 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
31 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
32 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
33 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
34 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
35 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
36 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
37 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
38 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
39 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
40 |  | Primary packaging: Unit of use: | Info | No |  | Yes
41 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
42 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
43 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
44 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
45 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
46 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
47 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
48 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
52 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
53 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
54 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
61 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
62 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 100
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Occluder and pinhole
MATERIAL: |  | S0004238
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
19 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
20 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
21 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
22 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
23 |  | Handheld occluder on one end with a pinhole on the other end. | Pass/Fail | Yes |  | No
24 |  | Configuration: Ruler with about 150mm with one circle of opaque material coloured black on each end. One of the circles has a pinhole in its centre. | Score | Yes |  | No
25 |  | Alternative configuration: Rectangular mask with a hole of about 25mm in the place of one of the eyes with a slide covert with a pinhole in is centre, and with a side holding handle. The pinhole must have a diameter between 0.94mm to 1.75mm. | Score | Yes |  | No
26 |  | The pinhole must have a diameter between 0.94mm to 1.75mm | Pass/Fail | Yes |  | No
27 |  | Pinhole as 1.2mm of diameter | Pass/Fail | Yes |  | No
28 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
29 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
30 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
31 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
32 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
33 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
34 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
35 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
36 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
37 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
38 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
39 |  | Primary packaging: Unit of use: | Info | No |  | Yes
40 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
41 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
42 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
43 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
44 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
45 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
46 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
47 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
48 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
51 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
52 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
53 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
54 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
60 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 110
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Ophthalmoscope-Loupe-Otoscope​
MATERIAL: |  | S0004264
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 10942:2006 Ophthalmic instruments — Direct ophthalmoscopes | Score | Yes |  | No
23 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Score | Yes |  | No
24 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Score | Yes |  | No
25 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Score | Yes |  | No
26 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Score | Yes |  | No
27 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
28 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Score | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
29 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
30 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
31 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
32 |  | Portable, pocket sized ophthalmoscope-loupe-otoscope for eye fundus and adnexa, eardrum, ear canal and adnexa observation | Pass/Fail | Yes |  | No
33 |  | Incorporated solar panel for battery recharging | Pass/Fail | Yes |  | No
34 |  | Battery powered | Pass/Fail | Yes |  | No
35 |  | Ear canal observation speculum | Pass/Fail | Yes |  | No
36 |  | Ability to magnify viewed image | Pass/Fail | Yes |  | No
37 |  | Dial with different viewing lens powers | Pass/Fail | Yes |  | No
38 |  | White and blue light | Pass/Fail | Yes |  | No
39 |  | Provided with an LED light source | Pass/Fail | Yes |  | No
40 |  | Distance visual acuity chart | Pass/Fail | Yes |  | No
41 |  | Reading chart | Pass/Fail | Yes |  | No
42 |  | Video clip for phone attachment | Pass/Fail | Yes |  | No
43 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
44 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
45 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
46 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
47 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
48 |  | Storage and transport case | Pass/Fail | Yes |  | No
49 |  | Cleaning cloth | Pass/Fail | Yes |  | No
50 |  | Securing lanyard | Pass/Fail | Yes |  | No
51 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
52 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
53 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
54 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
55 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
56 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
57 |  | Primary packaging: Unit of use: | Info | No |  | Yes
58 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
59 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
60 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
61 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
62 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
63 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
64 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
65 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
66 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
67 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
68 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
69 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
70 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 points
71 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
72 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
73 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
74 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
75 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
76 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
77 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
78 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
79 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
80 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 120
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Autorefractor, portable and low cost
MATERIAL: |  | S0004253
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments. | Pass/Fail | Yes |  | No
23 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
24 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. | Pass/Fail | Yes |  | No
25 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements. | Pass/Fail | Yes |  | No
26 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
27 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
28 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
29 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
30 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
31 |  | Automated refraction measurement device | Pass/Fail | Yes |  | No
32 |  | Desktop set-up | Pass/Fail | Yes |  | No
33 |  | Aided measurement | Pass/Fail | Yes |  | No
34 |  | Average of the measures with quality assessment | Pass/Fail | Yes |  | No
35 |  | Support for the head stability and fast operation. | Pass/Fail | Yes |  | No
36 |  | Refraction measurement: Spherical power -25.00DS to +22.00DS (0.12D/0.25D), cylinder power 0.00DC to ±10.00DC (0.12D/0.25D), axis 0⁰ to 180⁰ (1⁰) | Pass/Fail | Yes |  | No
37 |  | Vertex distance 0.0-12.0-13.5-15.0mm | Pass/Fail | Yes |  | No
38 |  | Minimum working pupil size 3mm | Pass/Fail | Yes |  | No
39 |  | Interpupillary distance 10mm to 88mm (steps 1mm) | Pass/Fail | Yes |  | No
40 |  | Screen for easy focus and measures taking | Pass/Fail | Yes |  | No
41 |  | Access to the last 10 measures for each eye, with average indication and quality of the measure assessment. | Pass/Fail | Yes |  | No
42 |  | Portable | Pass/Fail | Yes |  | No
43 |  | Battery operation autonomy in hours (includes all batteries supplied with the device, and that can be charged without interrupting device operation) | Score | Yes |  | Yes
44 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
45 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
46 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
47 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
48 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
49 |  | Portable autorefractor with forehead support. | Pass/Fail | Yes |  | No
50 |  | Integrated printer with RS 232C and USB connection. | Pass/Fail | Yes |  | No
51 |  | Power 100-240AC - 50-60HZ with power plug adequate for the specific country power network | Pass/Fail | Yes |  | No
52 |  | Leads, cables, and attachments can be securely attached to the unit, providing tidy storage when the product in use | Pass/Fail | Yes |  | No
53 |  | Other accessories no mentioned elsewhere | Info | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
54 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
55 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
56 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
57 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
58 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
59 |  | Primary packaging: Unit of use: | Info | No |  | Yes
60 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
61 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
62 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
63 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
64 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
65 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
66 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
67 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
68 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
69 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
70 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
71 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
72 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 point
73 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
74 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
75 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
76 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
77 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
78 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
79 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
80 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
81 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
82 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 130
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Projector & visual acuity system
MATERIAL: |  | S0004254
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points
18 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments. | Pass/Fail | Yes |  | No
19 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. | Pass/Fail | Yes |  | No
20 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. | Pass/Fail | Yes |  | No
21 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements. | Pass/Fail | Yes |  | No
22 |  | ISO 10938:2016 Ophthalmic optics — Chart displays for visual acuity measurement — Printed, projected and electronic | Pass/Fail | Yes |  | No
23 |  | ISO 8596:2017 Ophthalmic optics — Visual acuity testing — Standard and clinical optotypes and their presentation | Pass/Fail | Yes |  | No
24 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
25 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT ***
26 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
27 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
28 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
29 |  | Eye charts with a minimum range of 0.1 to 1.25 decimal visual acuity for all ages and literacy levels | Pass/Fail | Yes |  | No
30 |  | Tests for astigmatism power and axis determination | Pass/Fail | Yes |  | No
31 |  | Red green bi-ocular and binocular balance | Pass/Fail | Yes |  | No
32 |  | Red-green or equivalent binocular assessment of simultaneous perception | Pass/Fail | Yes |  | No
33 |  | Remote controller | Pass/Fail | Yes |  | No
34 |  | Support and fixation platform | Pass/Fail | Yes |  | No
35 |  | Projector screen, in case is it a projector system | Pass/Fail | Yes |  | No
36 |  | Easy access to the lamp for replacement | Pass/Fail | Yes |  | No
37 |  | Luminance equal or higher than 200 cd/m2 | Score | Yes |  | No
38 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
39 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
40 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
41 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
42 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
43 |  | Remote controller | Pass/Fail | Yes |  | No
44 |  | Support and fixation platform | Pass/Fail | Yes |  | No
45 |  | Power 100-240AC - 50-60HZ with power plug adequate for the specific country power network | Pass/Fail | Yes |  | No
46 |  | Power cord with minimum length of 2m | Pass/Fail | Yes |  | No
47 |  | Leads, cables, and attachments can be securely attached to the unit, providing tidy storage | Pass/Fail | Yes |  | No
48 |  | Other accessories no mentioned elsewhere | Info | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
49 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
50 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
51 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
52 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
53 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
54 |  | Primary packaging: Unit of use: | Info | No |  | Yes
55 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
56 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
57 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
58 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
59 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
60 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
61 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
62 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
63 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
65 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
66 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
67 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 points
68 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
69 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
70 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
71 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
72 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
73 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
74 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
75 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
76 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
77 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 140
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Cross cylinder
MATERIAL: |  | S0004255
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 21987:2017 Ophthalmic optics — Mounted spectacle lenses | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | Yes |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Lens with crossed cylinder of ±0.25DC power with markings for power and axis cylinder determination | Pass/Fail | Yes |  | No
25 |  | Red for negative and blue/green for positive power | Pass/Fail | Yes |  | No
26 |  | Handle for hand holding and fast rotation | Pass/Fail | Yes |  | No
27 |  | Material lens in plastic or glass | Pass/Fail | Yes |  | No
28 |  | Handle with a minimum of 68mm | Pass/Fail | Yes |  | No
29 |  | Lens diameter a minimum of 25mm | Pass/Fail | Yes |  | No
30 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | Yes |  | Yes
31 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | Yes |  | Yes
| x2 | PHYSICAL DIMENSIONS
32 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
33 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
34 |  | Storage case | Score | Yes |  | No
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
35 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
36 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
37 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
38 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
39 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
40 |  | Primary packaging: Unit of use: | Info | No |  | Yes
41 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
42 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
43 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
44 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
45 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
46 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
47 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
48 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
49 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
52 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
53 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
54 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
55 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
61 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
62 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 150
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Ophthalmic torchlight
MATERIAL: |  | S0004240
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments. | Pass/Fail | Yes |  | No
23 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
24 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests. | Pass/Fail | Yes |  | No
25 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
26 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
27 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
28 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
29 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
30 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
31 |  | Torchlight in shape of a pen, with a click on-off switch | Pass/Fail | Yes |  | No
32 |  | Portable and powered by battery | Pass/Fail | Yes |  | No
33 |  | Compact and lightweight for ease of use in eye care examinations | Pass/Fail | Yes |  | No
34 |  | The light source can be a miniature incandescent light bulb or light-emitting diodes (LEDs) | Pass/Fail | Yes |  | No
35 |  | The battery can be replaceable or integrated and sealed | Pass/Fail | Yes |  | No
36 |  | The light emitting source emits sufficient light for eliciting pupillary response in a normal eye, and for easy observation of the anterior eye surface and ocular adnexa. Indicate luminance in cd/m2. | Score | Yes |  | Yes
37 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
38 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
39 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
40 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
41 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
42 |  | Anti-scracth storage case | Score | Yes |  | No
43 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
44 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
45 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
46 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
47 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
48 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
49 |  | Primary packaging: Unit of use: | Info | No |  | Yes
50 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
51 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
52 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
53 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
54 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
55 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
56 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
57 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
58 |  | Upload a copy of the Free Sales Certifciate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
59 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
60 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
61 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
62 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
63 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
66 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
67 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
69 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
70 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
71 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 160
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Photo screener, infant and child
MATERIAL: |  | S0004241
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class IIa medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
23 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
24 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Pass/Fail | Yes |  | No
25 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
26 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
27 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
28 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
29 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
30 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
31 |  | Handheld, portable device used for vision screening, particularly in children. It uses a camera to capture images of a child’s non-dilated eyes. Allows for screening of the presence of refractive error, ocular axis alignment and amblyopia in a fast, simples and non-invasive technique. | Pass/Fail | Yes |  | No
32 |  | Sensitivity and specificity in refractive error and amblyopia above 70%, for each | Pass/Fail | Yes |  | No
33 |  | Pupil size 3.0mm to 8.0mm in 0.1mm increments | Pass/Fail | Yes |  | No
34 |  | Power measurement ranging from -7.00DS to +5.00DS in 0.25D increments | Pass/Fail | Yes |  | No
35 |  | Ability to detect misalignment greater than 3 prismatic dioptres | Pass/Fail | Yes |  | No
36 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
37 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
38 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
39 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
40 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
41 |  | Operation time with full charge, including backup batteries, is great than 8 hours | Pass/Fail | Yes |  | No
42 |  | All involved batteries have their recharge stand and are able to be recharged in less than 8 hours | Pass/Fail | Yes |  | No
43 |  | Recharge stand and backup batteries | Pass/Fail | Yes |  | No
44 |  | Regular power outlet 100-240V 5-15A with power plug adequate for the specific country power network | Pass/Fail | Yes |  | No
45 |  | Storage and transport case | Pass/Fail | Yes |  | No
46 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
47 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
48 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
49 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
50 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
51 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
52 |  | Primary packaging: Unit of use: | Info | No |  | Yes
53 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
54 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
55 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
56 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
57 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
58 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed rquirements after 1 October 2023. | Pass/Fail | Yes |  | No
59 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
60 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
62 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
63 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
64 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
65 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 points
66 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
67 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
69 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
70 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
71 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
72 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
73 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
74 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
75 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 170
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Fundus camera, non-mydriatic
MATERIAL: |  | S0004242
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class IIa medical device under EU regulations. Then indicate the applicable directive (MDD 93/42/EEC or MDR 2017/745/EEC ) and applicable annex. Upload a valid EC certificate under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 10940:2009 Ophthalmic instruments — Fundus cameras | Pass/Fail | Yes |  | No
23 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
24 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
25 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Pass/Fail | Yes |  | No
26 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
27 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
28 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
29 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
30 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
31 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
32 |  | Specialized imaging device for the observation and recording of the eye fundus | Pass/Fail | Yes |  | No
33 |  | Non-mydriatic | Pass/Fail | Yes |  | No
34 |  | Minimum working pupillary diameter equal to or smaller than 4mm | Pass/Fail | Yes |  | No
35 |  | Field of view >30° | Pass/Fail | Yes |  | No
36 |  | Desirable field of view with mosaic =>90° | Pass/Fail | Yes |  | No
37 |  | Camera resolution >15 megapixels | Pass/Fail | Yes |  | No
38 |  | Image resolution should be specified | Pass/Fail | Yes |  | No
39 |  | Autofluorescence | Score | Yes |  | No
40 |  | Light source: LED, halogen, or laser diode | Pass/Fail | Yes |  | No
41 |  | Auto-focus and manual focus | Pass/Fail | Yes |  | No
42 |  | Auto-capture and manual capture | Pass/Fail | Yes |  | No
43 |  | Software integration | Pass/Fail | Yes |  | No
44 |  | Portability | Pass/Fail | Yes |  | No
45 |  | Three-dimensional imaging | Pass/Fail | Yes |  | No
46 |  | Data export connectivity | Pass/Fail | Yes |  | No
47 |  | Ability to export generated reports in PDF format | Pass/Fail | Yes |  | No
48 |  | Database software organized by name and/or date | Pass/Fail | Yes |  | No
49 |  | Monitor resolution or external monitor connectivity | Pass/Fail | Yes |  | No
50 |  | Fixation targets and field of view should make it relatively straightforward to capture images centred on the foveal area, centred on the optic disc, following the local screening protocol, in the macular and nasal fields (EURODIAB Protocol), temporal to the macula, superior temporal, inferior temporal, superior nasal, inferior nasal (fields 3-7 of ETDRS) | Pass/Fail | Yes |  | No
51 |  | It should facilitate image capture by a unit that is TWAIN or customized API integrated, allowing the use of software not provided by the supplier for image capture | Pass/Fail | Yes |  | No
52 |  | It should automatically identify the eye from which the image was captured (left or right) | Pass/Fail | Yes |  | No
53 |  | It must achieve a resolution of at least 30 pixels per degree at all field points and orientations, excluding borders, where true colour is defined as the colour that would be seen with a direct view | Pass/Fail | Yes |  | No
54 |  | The unit must have the capability to manage and manipulate images without losing any clinically significant data, regardless of any down-sampling or compression applied | Pass/Fail | Yes |  | No
55 |  | When the raw image is compressed with a 'lossy' compression algorithm, such as JPEG, the unit must not apply further 'lossy' compression, for example, applying a software image mask | Pass/Fail | Yes |  | No
56 |  | It should be able to derive any image parameters necessary for displaying, transferring, or manipulating the image, such as height, width, resolution, colour depth, and file size | Pass/Fail | Yes |  | No
57 |  | These parameters should be readily available in a format that can be easily transferred to text or other files | Pass/Fail | Yes |  | No
58 |  | The unit must have the capability to save and transfer images with a file size that allows fast transfer from the unit to an external PC, ensuring fast transfer for subsequent grading and quality assurance, and affordable mass storage (each image is expected to fall within the 500KB - 2MB range) | Pass/Fail | Yes |  | No
59 |  | Images should be saved to a specified folder with a specified filename on the same or connected PC/peripheral device | Pass/Fail | Yes |  | No
60 |  | Supported by direct and immediate image transfer from the image capture device in the unit to the external computer without the need for any manual export process | Pass/Fail | Yes |  | No
61 |  | This transfer should allow the identification of multiple images acquired from any patient via a standard USB, linked to Hospital Information Systems for storage and via the internet for viewing images at other locations | Pass/Fail | Yes |  | No
62 |  | Indoor use | Pass/Fail | Yes |  | No
63 |  | Working distance inferior to 50cm, measured from patient’s eye to the front lens | Pass/Fail | Yes |  | No
64 |  | Wi-Fi and/or ethernet connection | Pass/Fail | Yes |  | No
65 |  | USB with personal computer and/or external monitor | Pass/Fail | Yes |  | No
66 |  | Software with database manager and images viewing | Pass/Fail | Yes |  | No
67 |  | Visualization of the anterior segment | Score | Yes |  | No
68 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
69 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
70 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
71 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
72 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
73 |  | Power 100-240AC - 50-60HZ with power plug adequate for the specific country power network | Pass/Fail | Yes |  | No
74 |  | Storage and transport case | Pass/Fail | Yes |  | No
75 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
76 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
77 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
78 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
79 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
80 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
81 |  | Primary packaging: Unit of use: | Info | No |  | Yes
82 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
83 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
84 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
85 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
86 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
87 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
88 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
89 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
90 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
91 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
92 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
93 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
94 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 point
95 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
96 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
97 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
98 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
99 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
100 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
101 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
102 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
103 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
104 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 180
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Retinoscope and/or ophthalmoscope
MATERIAL: |  | S0004243
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| D.0 | SECTION D. *** IF THE DEVICE CONTAINS BATTERIES OF ANY KIND ***
18 |  | Does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Info | Yes |  | No
19 |  | Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Info | Yes |  | No
20 |  | Indicate the UN hazardous classification code here, if applicable: | Info | No |  | Yes
21 |  | Confirm here that the Material Safety Datasheet for the offered product has been uploaded under the following placeholder for upload: ”MSDS Doc”. Indicate the filename of the uploaded document here. | Info | No |  | Yes
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
22 |  | ISO 12865:2006 Ophthalmic instruments — Retinoscopes | Pass/Fail | Yes |  | No
23 |  | ISO 10942:2006 Ophthalmic instruments — Direct ophthalmoscopes | Pass/Fail | Yes |  | No
24 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
25 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
26 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Pass/Fail | Yes |  | No
27 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
28 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
29 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
30 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
31 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
32 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
33 |  | Retinoscope: Streak retinoscope for refractive error determination. The head contains a white light source, a beam-split mirror, a condensing lens, a slit stop and a sleeve that controls both meridian (or streak rotation) and vergence (or streak focus) | Pass/Fail | Yes |  | No
34 |  | Ophthalmoscope for eye fundus observation: group A of ISO 10942:2022 | Pass/Fail | Yes |  | No
35 |  | Recharge stand with two slots for simultaneous recharging and cord for power outlet | Pass/Fail | Yes |  | No
36 |  | Battery powered | Pass/Fail | Yes |  | No
37 |  | Range of lenses not smaller than -35D to +20D, adjustable in steps | Pass/Fail | Yes |  | No
38 |  | Anti-reflection lens | Pass/Fail | Yes |  | No
39 |  | Ability to magnify at least five times | Pass/Fail | Yes |  | No
40 |  | Available apertures at a minimum: small, large and semi-circle, fixation star | Pass/Fail | Yes |  | No
41 |  | Colour temperature: Cool white in the range 3,500-4,000 K | Pass/Fail | Yes |  | No
42 |  | Red-free, blue, and green filters | Pass/Fail | Yes |  | No
43 |  | Handle with on/off switch | Pass/Fail | Yes |  | No
44 |  | Provided with an LED light source | Pass/Fail | Yes |  | No
45 |  | Is able to withstand frequent cleaning and disinfecting with hospital grade products | Pass/Fail | Yes |  | No
46 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
47 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
48 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
49 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
50 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
51 |  | Storage and transport case resistant to impact | Score | Yes |  | No
52 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
53 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
54 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
55 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
56 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
57 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
58 |  | Primary packaging: Unit of use: | Info | No |  | Yes
59 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
60 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
61 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
62 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
63 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
64 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
65 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
66 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
67 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
68 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
69 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
70 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
71 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 point
72 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
73 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
74 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
75 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
76 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
77 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
78 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
79 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
80 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
81 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 190
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Retinoscopy lens, rack
MATERIAL: |  | S0004244
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
19 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Retinoscopy lens rack that consists of two linear frames on a single rack —one containing positive lenses (increasing in optical power) and the other containing negative lenses to be used with a retinoscopy for easy and quick objective refraction | Pass/Fail | Yes |  | No
26 |  | Ruler with a set of positive and negative lens ranging from 0.50DS to 10DS in 0.50DS steps | Pass/Fail | Yes |  | No
27 |  | Colour green for positive and red for negative | Pass/Fail | Yes |  | No
28 |  | With a handle for grip and holding | Pass/Fail | Yes |  | No
29 |  | Lens diameter no less than 10mm, nor greater than 25mm | Pass/Fail | Yes |  | No
30 |  | The lenses are recessed to prevent scratching and increase durability | Score | No |  | Yes
31 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
32 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
33 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
34 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
35 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
36 |  | Storage and transport case | Score | Yes |  | No
37 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
38 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
39 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
40 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
41 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
42 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
43 |  | Primary packaging: Unit of use: | Info | No |  | Yes
44 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
45 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
46 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
47 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
48 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
49 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
50 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
51 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
55 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
56 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
57 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
64 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
65 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 200
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Trial frame, universal, adult
MATERIAL: |  | S0004248
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 12867:2010 Ophthalmic instruments — Trial frames | Pass/Fail | Yes |  | No
19 |  | ISO 9801:1997 Ophthalmic instruments — Trial case lenses | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Adjustable eye frame used to hold different trial lenses in front of an adult patient's eyes | Pass/Fail | Yes |  | No
26 |  | PD range from 45-82mm | Pass/Fail | Yes |  | No
27 |  | Lens holder for 10 trial lenses | Pass/Fail | Yes |  | No
28 |  | Adjustable temple angle and length for adult head sizes, with self-lock | Pass/Fail | Yes |  | No
29 |  | Adjustable bridge height with self-lock and smooth axis adjustment | Pass/Fail | Yes |  | No
30 |  | Easy removal and holding of the trial lens | Pass/Fail | Yes |  | No
31 |  | Non-allergenic materials | Pass/Fail | Yes |  | No
32 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
33 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
34 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
35 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
36 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
37 |  | Storage and transport case | Score | Yes |  | No
38 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
39 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
40 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
41 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
42 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
43 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
44 |  | Primary packaging: Unit of use: | Info | No |  | Yes
45 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
46 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
47 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
48 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
49 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
50 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
51 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
52 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
55 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
56 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
57 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
58 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
65 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
66 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 210
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Trial frame, universal, paediatric
MATERIAL: |  | S0004249
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 12867:2010 Ophthalmic instruments — Trial frames | Pass/Fail | Yes |  | No
19 |  | ISO 9801:1997 Ophthalmic instruments — Trial case lenses | Pass/Fail | Yes |  | No
20 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
21 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
22 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
23 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
24 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
25 |  | Adjustable eye frame used to hold different trial lenses in front of a paediatric patient's eyes | Pass/Fail | Yes |  | No
26 |  | Pupillary distance range within 40-60mm | Pass/Fail | Yes |  | No
27 |  | Lens holder for 10 trial lenses | Pass/Fail | Yes |  | No
28 |  | Adjustable angle and length of temples for children head sizes | Pass/Fail | Yes |  | No
29 |  | Adjustable bridge height with self-lock and smooth axis adjustment | Pass/Fail | Yes |  | No
30 |  | Easy removal and holding of the trial lens | Pass/Fail | Yes |  | No
31 |  | Non-allergenic materials | Pass/Fail | Yes |  | No
32 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
33 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
34 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
35 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
36 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
37 |  | Storage and transport case | Score | Yes |  | No
38 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
39 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
40 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
41 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
42 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
43 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
44 |  | Primary packaging: Unit of use: | Info | No |  | Yes
45 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
46 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
47 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
48 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
49 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
50 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
51 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
52 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
55 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
56 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
57 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
58 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
65 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
66 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 220
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Trial lens set
MATERIAL: |  | S0004250
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 9801:2009 Ophthalmic instruments Trial case lenses | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | The trial lens set includes a minimum set of plastic or glass lenses with plastic or metal rim and holder | Pass/Fail | Yes |  | No
25 |  | Spherical lenses, both positive and negative, from 0.25D to 4.00D in 0.25D steps, from 4.00D to 8.00D in 0.50D steps, from 8.00D to 20.00D in 1.00D steps | Pass/Fail | Yes |  | No
26 |  | Cylindrical lenses in a range of powers, often both positive and negative, from 0.25D to 4.00D in 0.25D steps, from 4.00D to 6.00D in 0.50D steps | Pass/Fail | Yes |  | No
27 |  | Accessory lenses should include prisms from 0.50DP, 1DP to 12 DP in 1DP steps | Pass/Fail | Yes |  | No
28 |  | Range of lenses not smaller than -35D to +40D, adjustable in steps | Pass/Fail | Yes |  | No
29 |  | Green and red filters | Pass/Fail | Yes |  | No
30 |  | Occluder | Pass/Fail | Yes |  | No
31 |  | Pinhole | Pass/Fail | Yes |  | No
32 |  | Maddox rod in red | Pass/Fail | Yes |  | No
33 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
34 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
35 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
36 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
37 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
38 |  | Anti-scratch and impact resistance storage and transport case with easy lens presentation and handling in consultation | Pass/Fail | Yes |  | No
39 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
40 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
41 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
42 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
43 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
44 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
45 |  | Primary packaging: Unit of use: | Info | No |  | Yes
46 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
47 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
48 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
49 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
50 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
51 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
52 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
53 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
55 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
56 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
57 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
58 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
59 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
66 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
67 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 230
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifier, handheld, illuminated
MATERIAL: |  | S0004231
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
20 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
21 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Pass/Fail | Yes |  | No
22 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
23 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
24 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
25 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
26 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
27 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
28 |  | Handled optical magnifier that magnifies (enlarged) images of close objects and print. Available as lens held in round frame with handle. | Pass/Fail | Yes |  | No
29 |  | Lens power range: +4D to +56D | Pass/Fail | Yes |  | No
30 |  | Minimum lens diameter as follows: ≤ +8D: ≥ 90 mm, > +8D to +12D: ≥ 80 mm, > +12D to +16D: ≥ 60 mm, > +16D to +20D: ≥ 45 mm, > +20D to +26D: ≥ 35 mm, > +26D to +40D: ≥ 30 mm, ≥ +40D: ≥ 26 mm | Pass/Fail | Yes |  | No
31 |  | Lens material: plastic | Pass/Fail | Yes |  | No
32 |  | Frame (mount) and handle: high grade plastic | Pass/Fail | Yes |  | No
33 |  | Illumination options for illuminated magnifiers: LED, SMD LED or incandescent bulbs | Pass/Fail | Yes |  | No
34 |  | In-built illumination distributed evenly across area being viewed | Pass/Fail | Yes |  | No
35 |  | Bundle of 3 magnifiers distributed as follows for power, magnification and diameter: 8D 2x 90mm, 10D 3x 80mm, 18D 5.5x 45mm. | Score | Yes |  | No
36 |  | Other bundles with justification | Score | No |  | Yes
37 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
38 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
39 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
40 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
41 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
42 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
43 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
44 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
45 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
46 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
47 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
48 |  | Primary packaging: Unit of use: | Info | No |  | Yes
49 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
50 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
51 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
52 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
53 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
54 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
55 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
56 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
57 |  | Upload a copy of the Free Sales Certifciate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
58 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
59 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
60 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
61 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
62 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
66 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
67 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
69 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
70 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 240
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifier, handheld, non-illuminated
MATERIAL: |  | S0004232
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Handled optical magnifier that magnifies (enlarged) images of close objects and print. Available as lens held in round frame with handle. | Pass/Fail | Yes |  | No
25 |  | Lens power range: +4D to +56D | Pass/Fail | Yes |  | No
26 |  | Minimum lens diameter as follows: ≤ +8D: ≥ 90 mm, > +8D to +12D: ≥ 80 mm, > +12D to +16D: ≥ 60 mm, > +16D to +20D: ≥ 45 mm, > +20D to +26D: ≥ 35 mm, > +26D to +40D: ≥ 30 mm, ≥ +40D: ≥ 26 mm | Pass/Fail | Yes |  | No
27 |  | Lens material: plastic | Pass/Fail | Yes |  | No
28 |  | Frame (mount) and handle: high grade plastic | Pass/Fail | Yes |  | No
29 |  | Bundle of 3 magnifiers distributed as follows for power, magnification and diameter: 8D 2x 90mm, 10D 3x 80mm, 18D 5.5x 45mm. | Score | Yes |  | No
30 |  | Other bundles with justification | Score | No |  | Yes
31 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
32 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
33 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
34 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
35 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
36 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
37 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
38 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
39 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
40 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
41 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
42 |  | Primary packaging: Unit of use: | Info | No |  | Yes
43 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
44 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
45 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
46 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
47 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
48 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
49 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
50 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
54 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
55 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
56 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
63 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 250
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifier, pocket
MATERIAL: |  | S0004233
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Pocket sized optical magnifier that magnifies (enlarged) images of close objects and print. | Pass/Fail | Yes |  | No
25 |  | Lens power of 8D to 10D with lens diameter of 50mm | Pass/Fail | Yes |  | No
26 |  | Lens power of 12D with lens diameter of 55mm | Pass/Fail | Yes |  | No
27 |  | Lens power of 24D with lens diameter of 55mm | Pass/Fail | Yes |  | No
28 |  | Bundle of 3 magnifiers distributed as follows for power and diameter: 8D 50mm, 12D 85mm, 24D 45mm | Score | Yes |  | No
29 |  | Other bundles with justification | Score | No |  | Yes
30 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
31 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
32 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
33 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
34 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
35 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
36 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
37 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
38 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tetriary packaging. | Pass/Fail | Yes |  | No
39 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
40 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
41 |  | Primary packaging: Unit of use: | Info | No |  | Yes
42 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
43 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
44 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
45 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
46 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
47 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
48 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
49 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
50 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
53 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
54 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
55 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
56 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
62 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
63 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 260
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifier, spectacle
MATERIAL: |  | S0004234
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Lens or lens system mounted into spectacle frames for magnification imaging of closed objects or print. | Pass/Fail | Yes |  | No
25 |  | Enables binocularity when base-in prisms are incorporated, to offer better field of view | Pass/Fail | Yes |  | No
26 |  | Lens power: single-vision lens power range +4.00 D to +24.00 D, in steps of 1.00D | Pass/Fail | Yes |  | No
27 |  | Solar UV transmittance: for UVB transmittance <5%, for UVA transmittance calculated up to 400nm < 50% | Pass/Fail | Yes |  | No
28 |  | Bundle of 4 spectacles magnifiers with one for each powers: +4D, +8D, +12D and +20D. | Score | Yes |  | No
29 |  | Others bundles with justification | Score | No |  | Yes
30 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
31 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
32 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
33 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
34 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
35 |  | Anti-scracth storage case | Pass/Fail | Yes |  | No
36 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
37 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
38 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
39 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
40 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
41 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
42 |  | Primary packaging: Unit of use: | Info | No |  | Yes
43 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
44 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
45 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
46 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
47 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
48 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
49 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
50 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
54 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
55 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
56 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
63 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 270
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifiers, bar, ruler
MATERIAL: |  | S0004235
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Non-illuminated magnifier, shaped like a half-cane with no less than 210 mm of viewing usable length, and 25 mm width | Pass/Fail | Yes |  | No
25 |  | Power magnification range: 1.5x to 4x | Pass/Fail | Yes |  | No
26 |  | Lens material: Plastic | Pass/Fail | Yes |  | No
27 |  | Grip handle, guide line across its width | Score | Yes |  | No
28 |  | Bundle of 2 bar magnifiers distributed as one for each magnification: 2x and 4x | Score | Yes |  | No
29 |  | Others bundles with justification | Score | No |  | Yes
30 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
31 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
32 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
33 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
34 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
35 |  | Anti-scracth storage case | Score | Yes |  | No
36 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
37 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
38 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
39 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tetriary packaging. | Pass/Fail | Yes |  | No
40 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
41 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
42 |  | Primary packaging: Unit of use: | Info | No |  | Yes
43 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
44 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
45 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
46 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
47 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
48 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
49 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
50 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
54 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
55 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
56 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
63 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 280
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifiers, neck
MATERIAL: |  | S0004236
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Non-illuminated, magnifier mounted in a frame, suspended by a cord around a person's neck and usually stabilised against their chest | Pass/Fail | Yes |  | No
25 |  | Magnification range: 2x to 4x | Pass/Fail | Yes |  | No
26 |  | Lens material: Plastic | Pass/Fail | Yes |  | No
27 |  | Diameter range: 100 to 90 mm | Pass/Fail | Yes |  | No
28 |  | Bundle of 2 neck magnifiers distributed as follows for powers: one for 2x and one for 4x | Score | Yes |  | No
29 |  | Other bundles with justification | Score | No |  | Yes
30 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
31 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
32 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
33 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
34 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
35 |  | Anti-scracth storage case | Score | Yes |  | No
36 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
37 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
38 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
39 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
40 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
41 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
42 |  | Primary packaging: Unit of use: | Info | No |  | Yes
43 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
44 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
45 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
46 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
47 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
48 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
49 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
50 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
51 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
54 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
55 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
56 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
57 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
63 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 290
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifiers, stand, illuminated
MATERIAL: |  | S0004237
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | ISO 15004-1:2020 Ophthalmic instruments — Fundamental requirements and test methods — Part 1: General requirements applicable to all ophthalmic instruments | Pass/Fail | Yes |  | No
20 |  | IEC 60601-1:2005 + A1:2012(E) Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | Pass/Fail | Yes |  | No
21 |  | IEC 60601-1-2:2014 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests | Pass/Fail | Yes |  | No
22 |  | EN ISO 15223-1:2021 (EN 980) Medical devices – Symbols to be used with medical device labels, labelling and Information to be supplied – Part 1: General requirements | Pass/Fail | Yes |  | No
23 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
24 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
25 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
26 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
27 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
28 |  | Stand optical magnifier that magnifies (enlarged) images of close objects and print. Consists in a lens mounted on a base that rests on the page. The lens magnifies the text and the base provides a fixed focal distance. | Pass/Fail | Yes |  | No
29 |  | Lens power range +6D to +76D | Pass/Fail | Yes |  | No
30 |  | Minimum lens diameter as follows: ≤ +8D: ≥ 90 mm, > +8D to +12D: ≥ 80 mm, > +12D to +16D: ≥ 60 mm, > +16D to +20D: ≥ 45 mm, > +20D to +26D: ≥ 35 mm, > +26D to +40D: ≥ 30 mm, ≥ +40D: ≥ 26 mm | Pass/Fail | Yes |  | No
31 |  | Lens material: plastic | Pass/Fail | Yes |  | No
32 |  | Frame (mount) and handle: high grade plastic | Pass/Fail | Yes |  | No
33 |  | Illumination options for illuminated magnifiers: LED, SMD LED or incandescent bulbs | Pass/Fail | Yes |  | No
34 |  | In-built illumination distributed evenly across area being viewed | Pass/Fail | Yes |  | No
35 |  | Bundle of 3 magnifiers distributed as follows for power, magnification and diameter: 8D 2x 90mm, 10D 3x 80mm, 18D 5.5x 45mm. | Score | Yes |  | No
36 |  | Others bundles with justification | Score | No |  | Yes
37 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
38 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
39 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
40 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
41 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
42 |  | Anti-scracth storage case | Score | Yes |  | No
43 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
44 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
45 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
46 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
47 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
48 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
49 |  | Primary packaging: Unit of use: | Info | No |  | Yes
50 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
51 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
52 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
53 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
54 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
55 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
56 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
57 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
58 |  | Upload a copy of the Free Sales Certifciate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
59 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
60 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
61 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
62 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
63 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
66 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
67 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
69 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
70 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
71 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 300
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Magnifiers, stand, non-illuminated
MATERIAL: |  | S0004239
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 15253: 2021 Optical and electro-optical devices for enhancing low vision | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Stand optical magnifier that magnifies (enlarged) images of close objects and print. Consists in a lens mounted on a base that rests on the page. The lens magnifies the text and the base provides a fixed focal distance. | Pass/Fail | Yes |  | No
25 |  | Lens power range +6D to +76D | Pass/Fail | Yes |  | No
26 |  | Minimum lens diameter as follows: ≤ +8D: ≥ 90 mm > +8D to +12D: ≥ 80 mm > +12D to +16D: ≥ 60 mm > +16D to +20D: ≥ 45 mm > +20D to +26D: ≥ 35 mm > +26D to +40D: ≥ 30 mm ≥ +40D: ≥ 26 mm | Pass/Fail | Yes |  | No
27 |  | Lens material: plastic | Pass/Fail | Yes |  | No
28 |  | Frame (mount): high-grade plastic | Pass/Fail | Yes |  | No
29 |  | Bundle of 3 magnifiers distributed as follows for power, magnification and diameter: 8D 2x 90mm, 10D 3x 80mm, 18D 5.5x 45mm. | Score | Yes |  | No
30 |  | Supports around lens mount are small or transparent to avoid blocking light from area being viewed | Score | Yes |  | No
31 |  | Special designed smaller supporting stands for more area exposer for better spot-writing function | Score | Yes |  | No
32 |  | Has adjustable lens to compensate for user’s uncorrected refractive error | Score | Yes |  | No
33 |  | Other bundles with justification | Score | No |  | Yes
34 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
35 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
36 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
37 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
38 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
39 |  | Anti-scracth storage case | Score | Yes |  | No
40 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
41 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
42 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
43 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
44 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
45 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
46 |  | Primary packaging: Unit of use: | Info | No |  | Yes
47 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
48 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
49 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
50 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
51 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
52 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
53 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
54 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
55 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
56 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
57 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
58 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
59 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
60 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
66 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
67 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
68 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 310
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Spectacles, pre-made
MATERIAL: |  | S0004245
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 16034:2002 Ophthalmic Optics − Specifications for Single Vision Ready-to-wear Near-vision Spectacles where the power can be either positive or negative | Pass/Fail | Yes |  | No
19 |  | ISO 21987 Ophthalmic optics — Mounted spectacle lenses | Pass/Fail | Yes |  | No
20 |  | ISO 7998:2005 Ophthalmic optics — Spectacle frames — Lists of equivalent terms and vocabulary | Pass/Fail | Yes |  | No
21 |  | ISO 8624:2020 Ophthalmic optics — Spectacle frames — Measuring system and vocabulary | Pass/Fail | Yes |  | No
22 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
23 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
24 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
25 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
26 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
27 |  | Eye frames and pre-shaped lenses that are pre-manufactured and can be assembled with a range of standard powers and dispensed over the counter | Pass/Fail | Yes |  | No
28 |  | Lens power range +3.50DS to -3.50DS in 0.25D steps | Pass/Fail | Yes |  | No
29 |  | Lens material: plastic | Pass/Fail | Yes |  | No
30 |  | Solar UV transmittance: UV absorption should be 95% and above close to 400nm | Pass/Fail | Yes |  | No
31 |  | Full-field and unisex | Pass/Fail | Yes |  | No
32 |  | Pupillary distance >18 year of age adult male 63 mm | Pass/Fail | Yes |  | No
33 |  | Pupillary distance >18 year of age adult women 61 mm | Pass/Fail | Yes |  | No
34 |  | Pupillary distance adult women 57 mm | Pass/Fail | Yes |  | No
35 |  | Pupillary distance male 14-18y: 61 mm | Pass/Fail | Yes |  | No
36 |  | Pupillary distance female 14-18y: 59 mm | Pass/Fail | Yes |  | No
37 |  | Pupillary distance 9-13y: 57 mm | Pass/Fail | Yes |  | No
38 |  | Pupillary distance Child 4-8y: 52 mm | Pass/Fail | Yes |  | No
39 |  | Lens hard coating | Score | Yes |  | No
40 |  | Bundle of 44 spectacles/ready-made for distance vision, 11 spectacles for each of the following pupillary distances: Adult male 63 mm, adult female 61 mm , male 14-18y 59 mm, female 14-18y 57 mm and child 4-8y 52 mm, distributed as one for each power step in the following range: -3.50 to -1.00 and +1.00 to +3.50 in 0.25DS | Score | Yes |  | No
41 |  | Other bundles with justification | Score | No |  | Yes
42 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
43 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
44 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
45 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
46 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
47 |  | Anti-scracth storage case | Pass/Fail | Yes |  | No
48 |  | Cleaning cloth | Pass/Fail | Yes |  | No
49 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
50 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
51 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
52 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
53 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
54 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
55 |  | Primary packaging: Unit of use: | Info | No |  | Yes
56 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
57 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
58 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
59 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
60 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
61 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed rquirements after 1 October 2023. | Pass/Fail | Yes |  | No
62 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
63 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
64 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
65 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
66 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
67 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
68 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING *** | Max 10 point
69 |  | Indicate whether end-user training is available and its format (in-person or remote, live or asynchronous) | Score | Yes |  | Yes
70 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
71 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
72 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
73 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
74 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
75 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
76 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
77 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
78 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 320
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Spectacles, ready-made, distance vision
MATERIAL: |  | S0004246
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 16034:2002 Ophthalmic Optics − Specifications for Single Vision Ready-to-wear Near-vision Spectacles where the power can be either positive or negative | Pass/Fail | Yes |  | No
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Eye glasses that are pre-manufactured or assembled with a range of standard powers and dispensed over the counter | Pass/Fail | Yes |  | No
25 |  | Lens power range +3.50DS to -3.50DS in 0.25D steps | Pass/Fail | Yes |  | No
26 |  | Lens material: plastic | Pass/Fail | Yes |  | No
27 |  | Solar UV transmittance: UV absorption should be 95% and above close to 400nm | Pass/Fail | Yes |  | No
28 |  | Full-field and unisex | Pass/Fail | Yes |  | No
29 |  | Pupillary distance >18 year of age adult male 63 mm | Pass/Fail | Yes |  | No
30 |  | Pupillary distance >18 year of age adult women 61 mm | Pass/Fail | Yes |  | No
31 |  | Pupillary distance adult women 57 mm | Pass/Fail | Yes |  | No
32 |  | Pupillary distance male 14-18y: 61 mm | Pass/Fail | Yes |  | No
33 |  | Pupillary distance female 14-18y: 59 mm | Pass/Fail | Yes |  | No
34 |  | Pupillary distance 9-13y: 57 mm | Pass/Fail | Yes |  | No
35 |  | Pupillary distance child 4-8y: 52 mm | Pass/Fail | Yes |  | No
36 |  | Lens hard coating | Score | Yes |  | No
37 |  | Bundle of 44 spectacles/ready-made for distance vision, 11 spectacles for each of the following pupillary distances: Adult male 63 mm, adult female 61 mm , male 14-18y 59 mm, female 14-18y 57 mm and child 4-8y 52 mm, distributed as one for each power step in the following range: -3.50 to -1.00 and +1.00 to +3.50 in 0.25DS | Score | Yes |  | No
38 |  | Other bundles with justification | Score | No |  | Yes
39 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
40 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
41 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
42 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
43 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
44 |  | Anti-scracth storage case | Pass/Fail | Yes |  | No
45 |  | Cleaning cloth | Pass/Fail | Yes |  | No
46 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
47 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
48 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
49 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
50 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
51 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
52 |  | Primary packaging: Unit of use: | Info | No |  | Yes
53 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
54 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
55 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
56 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
57 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
58 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
59 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
60 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | Upload a copy of the Free Sales Certifciate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
62 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
63 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
64 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
65 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
66 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
67 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
69 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
70 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
71 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
72 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
73 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
74 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 330
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | Spectacles, ready-made, near vision
MATERIAL: |  | S0004247
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | ISO 16034:2002 Ophthalmic Optics − Specifications for Single Vision Ready-to-wear Near-vision Spectacles | Pass/Fail | Yes |  | No
19 |  | EN 14139:2002 Ophthalmic Optics − Specifications for Single Vision Ready-to-wear Near-vision Spectacles | Pass/Fail | Yes |  | No
20 |  | US ANSI standard: Z80.31–2017 Specifications for Ready-to-wear Near-vision Spectacles | Pass/Fail | Yes |  | No
21 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
22 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
23 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
24 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
25 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
26 |  | Eye glasses for near vision that are pre-manufactured or assembled with a range of standard powers and dispensed over the counter | Pass/Fail | Yes |  | No
27 |  | Lens power range +1.00DS to +3.50 DS in 0.25D steps | Pass/Fail | Yes |  | No
28 |  | Lens material: plastic | Pass/Fail | Yes |  | No
29 |  | Solar UV transmittance: UV absorption should be 95% and above close to 400nm | Pass/Fail | Yes |  | No
30 |  | Full-field and half-eye | Pass/Fail | Yes |  | No
31 |  | Length range: tall adults: 145cm ± 5cm adults: 120−140cm ± 5cm children: 70cm ±5 cm | Pass/Fail | Yes |  | No
32 |  | Pupillary distance adult male 59 mm | Pass/Fail | Yes |  | No
33 |  | Pupillary distance adult women 57 mm | Pass/Fail | Yes |  | No
34 |  | Lens hard coating | Score | Yes |  | Yes
35 |  | Bundle of 22 spectacles/ready-made for near vision, 11 with PD of 59 mm and 11 with PD of 57 mm, distributed as one for each power step in the following range: +1.00 to +3.50 in 0.25DS | Score | Yes |  | No
36 |  | Other bundles with justification | Score | No |  | Yes
37 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
38 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
39 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
40 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
41 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
42 |  | Anti-scracth storage case | Pass/Fail | Yes |  | No
43 |  | Cleaning cloth | Pass/Fail | Yes |  | No
44 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
45 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
46 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
47 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
48 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
49 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
50 |  | Primary packaging: Unit of use: | Info | No |  | Yes
51 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
52 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
53 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
54 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
55 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
56 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
57 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
58 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
59 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
60 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
61 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
62 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
63 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
64 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
65 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
66 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
67 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
68 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
69 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
70 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
71 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
72 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 340
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | White cane
MATERIAL: |  | S0004251
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | White canes should comply with and be tested according to relevant national or international standards. Documents that provide detailed reports of tests performed, should also be provided. If white canes do not comply with or are not tested according to relevant national or international standards, an explanation should be provided. All documentation should be in the official language or in English (other languages could be specified too). There are no specific product standards for white canes, however considerations could be taken to the general standards e.g. EN 12182 Assistive products for persons with disability - General requirements and test methods. | Info | No |  | Yes
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Physical device shaped like a cane that helps people who are blind or visually impaired to navigate their surroundings | Pass/Fail | Yes |  | No
25 |  | Straight cane: Cane with straight tubing, roller tip (e.g. nylon) | Score | No |  | Yes
26 |  | Angular white cane: Straight tubing with handgrip bent 150°, roller tip (e.g. nylon) | Score | No |  | Yes
27 |  | Aluminium, fibreglass or carbon fibre straight tubing with outer diameter 1.2cm or smaller | Pass/Fail | Yes |  | No
28 |  | 360° turn roller or marshmallow-style nylon tip 2−3cm thick | Pass/Fail | Yes |  | No
29 |  | Length range: tall adults: 145cm ± 5cm adults: 120−140cm ± 5cm children: 70cm ±5 cm | Pass/Fail | Yes |  | No
30 |  | Plastic or rubber handgrip of length 20cm and diameter 2.5cm and wrist loop | Pass/Fail | Yes |  | No
31 |  | Angular white cane: Same as above. In addition, the handgrip should be bent at 150 degrees from the main body of the cane | Pass/Fail | Yes |  | No
32 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
33 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
34 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
35 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
36 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
37 |  | Storage and transport case | Score | Yes |  | No
38 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
39 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
40 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
41 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
42 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
43 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
44 |  | Primary packaging: Unit of use: | Info | No |  | Yes
45 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
46 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
47 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
48 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
49 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
50 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
51 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
52 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
55 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
56 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
57 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
58 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
64 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
65 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
66 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes

## 350
Bidders are instructed to exclusively respond within the designated white shaded cells of the questionnaire. The form has been secured to guarantee adherence to this stipulation. Any attempt to unlock or make modifications to this questionnaire may result in the disqualification of the bid.
MATERIAL SPECIFICATIONS
DESCRIPTION: |  | White cane, folding
MATERIAL: |  | S0004252
No. | Header | DESCRIPTION | Score or Pass/Fail or Info | Show compliant question? | YES / NO | Allow free text to be entered? | FREE TEXT (Or indicate the actual specification of the offered unit)
| A.0 | SECTION A. *** MANUFACTURER IDENTIFICATION ***
1 |  | Is the supplier ISO 9001 and/or ISO 13485 certified? If yes, provide expiry date and upload the certificate(s) to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
2 |  | If the bidder is not the manufacturer provide manufacturers details: - Name manufacturer - Address manufacturer - Postal code - City - Country manufacturer | Pass/Fail | No |  | Yes
3 |  | If the manufacturer has a website, provide the weblink: | Pass/Fail | No |  | Yes
4 |  | If the manufacturer has an online catalogue, provide link or upload the catalogue to the placeholder named: "Brochures" Indicate the filename of the uploaded document(s) here. | Info | No |  | Yes
5 |  | Is the manufacturer ISO 9001 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
6 |  | Is the manufacturer ISO 13485 certified? If yes, provide expiry date and upload the certificate to the placeholder named: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
7 |  | If the manufacturing process(es) is/are subcontracted provide: - Subcontracted process - Name and address and country of subcontractor - Attach QMS certificate of subcontractor to this submission and Indicate the filename of the uploaded document(s) here. | Pass/Fail | No |  | Yes
| B.0 | SECTION B. *** DEVICE IDENTIFICATION ***
8 |  | Product name: | Pass/Fail | No |  | Yes
9 |  | Product reference/model number from supplier: | Pass/Fail | No |  | Yes
10 |  | Product reference/model number from manufacturer, if different from supplier: | Pass/Fail | No |  | Yes
11 |  | Confirm here that a general brochure of the offered item has been included in this bid submission. | Pass/Fail | Yes |  | No
12 |  | Confirm here that all Information required for UNICEF to verification all listed specifications has been included in this bid submission, e.g. publicised datasheets or brochures. Where specifications are not available in such documents, the bidder has developed a customised document detailing those specifications and confirming their compliance. This document is to be signed by an authorised representative of the bidder and should be attached to this submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
13 |  | Confirm here that an image of the product with a white background has been attached to this submission (Min: 2000 x 3000 pixels without compression. preferred: 3000 x 4000, PNG- no Word documents) By including an image in this submission the bidder automatically provides authorization to UNICEF to use the image in UNICEF's catalogue for the purpose of identifying the product. | Pass/Fail | Yes |  | No
| C.0 | SECTION C. *** DEVICE CLASSIFICATION AND MARKET CLEARANCE ***
14 |  | Confirm if the product has market clearance as a class I medical device under EU MDR 2017/745/EEC regulations. Upload the EC self-declaration under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
15 |  | Confirm if the product has market clearance from the USA FDA, indicate the classification here and Upload proof of USA market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
16 |  | Confirm if the product has market clearance from either Canada, Japan, or Australia, Upload proof of this market clearance under the placeholder: "Certificates". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
17 |  | Is the product compliant with the UNICEF's Technical Provisions for Medical Devices as published on the UNICEF website: https://www.unicef.org/supply/documents/technical-requirements-medical-devices | Pass/Fail | Yes |  | No
| E.0 | SECTION E. *** COMPLIANCE TO TECHNICAL STANDARDS *** If the declaration of compliance is based on report(s) issued by an independent testing laboratory, the reference of the test report must be indicated. (Mandatory for safety compliance of electro-medical devices) | Max 50 points | 0 |  | 0
18 |  | White canes should comply with and be tested according to relevant national or international standards. Documents that provide detailed reports of tests performed, should also be provided. If white canes do not comply with or are not tested according to relevant national or international standards, an explanation should be provided. All documentation should be in the official language or in English (other languages could be specified too). There are no specific product standards for white canes, however considerations could be taken to the general standards e.g. EN 12182 Assistive products for persons with disability - General requirements and test methods. | Info | Yes |  | Yes
19 |  | Indicate here any other series standards the products complies with (e.g. ISO, IEC, etc.) attached any available certifications to this submission. Indicate the file name of the uploaded document here. (one point for each standard) | Score | No |  | Yes
20 |  | Confirm a Declaration of Conformity to applicable standards signed by an authorized representative of the bidder, has been attached to this submission. Indicate the file name of the uploaded document here. | Pass/Fail | Yes |  | Yes
| F.0 | SECTION F. *** OPERATING ENVIRONMENT *** |  | 0 |  | 0
21 |  | Storage conditions: 0 - 50°C / 90% relative humidity. | Pass/Fail | Yes |  | Yes
22 |  | Operating conditions: 10 - 40°C / 90% relative humidity: | Pass/Fail | Yes |  | Yes
23 |  | Indicate the atmospheric pressure range the equipment is able to handle: | Info | No |  | Yes
| G.0 | SECTION G. *** TECHNICAL SPECIFICATIONS / CHARACTERISTICS *** | Max 100 points | 0 |  | 0
24 |  | Physical device shaped like a cane that helps people who are blind or visually impaired to navigate their surroundings | Pass/Fail | Yes |  | No
25 |  | Cane with foldable tubing, roller tip (e.g. nylon) | Pass/Fail | Yes |  | No
26 |  | Foldable aluminium, fibreglass or carbon fibre tubing with outer diameter 1.2cm or smaller; with four or five folds; and with elastic cord running through the middle of the tubing | Pass/Fail | Yes |  | No
27 |  | Aluminium, fibreglass or carbon fibre straight tubing with outer diameter 1.2cm or smaller | Pass/Fail | Yes |  | No
28 |  | 360° turn roller or marshmallow-style nylon tip 2−3cm thick | Pass/Fail | Yes |  | No
29 |  | Length range: tall adults: 145cm ± 5cm adults: 120−140cm ± 5cm children: 70cm ±5 cm | Pass/Fail | Yes |  | No
30 |  | Plastic or rubber handgrip of length 20cm and diameter 2.5cm and wrist loop | Pass/Fail | Yes |  | No
31 |  | Additional technical characteristics or features, not mentioned elsewhere | Score | No |  | Yes
32 |  | Additional capabilities (includes capabilities integrated in the device besides its primary main function) | Score | No |  | Yes
| x2 | PHYSICAL DIMENSIONS
33 |  | Indicate the size of the packaged product in dm³ or m³. | Pass/Fail | No |  | Yes
34 |  | Indicate the total weight of the packaged product in gr or kg. | Pass/Fail | No |  | Yes
| H.0 | SECTION H. *** REQUIRED ACCESSORIES, SUPPLIES AND/OR CONSUMABLES *** | Max 10 points | 0 |  | 0
35 |  | Instructions for assembly, use and maintenance in English, French and Spanish. | Pass/Fail | Yes |  | No
36 |  | Storage and transport case | Score | Yes |  | No
37 |  | Other accessories no mentioned elsewhere | Score | No |  | Yes
| I.0 | SECTION I. *** WARRANTY AND MAINTENANCE *** | Max 10 points
38 |  | Indicate the estimated life span of the product: | Info | No |  | Yes
39 |  | For products which have an expiry date, indicate here the maximum shelf-life: | Pass/Fail | Yes |  | Yes
40 |  | For those products which have a confirmed shelf-life in the above question, confirm here that the expiry date is indicated on the primary, secondary and tertiary packaging. | Pass/Fail | Yes |  | No
41 |  | Warranty period should be minimal two (2) years including supplied accessories | Pass/Fail | Yes |  | No
42 |  | Additional warranty extension (more than two years) | Score | No |  | Yes
| J.0 | SECTION J. *** PACKAGING ***
43 |  | Primary packaging: Unit of use: | Info | No |  | Yes
44 |  | Secondary packaging: Number of units on a palete on in a carton: | Info | No |  | Yes
45 |  | Product packaging and labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
| K.0 | SECTION K. *** LABELLING / DOCUMENTATION / CATEGORIZATION *** Indicate with which labelling requirements the product is in compliance: At a minimum at least one should be selected.
46 |  | Products MDR - CE certified follow labelling regulations as stipulated in EU MDR 2017/745 Annex I, Chapter III, point 23.2. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:32017R0745&qid=1622012817907&from=EN#d1e32-94-1 | Pass/Fail | Yes |  | No
47 |  | Products MDD - CE certified follow labelling requirements as stipulated in EU MDD 93/42/EEC, Annex I, point 13.3. Follow below hyperlink for details on this regulation. https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:31993L0042&qid=1638369660570&from=EN | Pass/Fail | Yes |  | No
48 |  | Products with market clearance from the USA FDA follow labelling requirements as stipulated in CFR code of regulations Title 21, part 801 - Labelling. Follow below hyperlink for details on this regulation. https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-801 | Pass/Fail | Yes |  | No
49 |  | Product labelling is compliant with UNICEF packaging and labelling requirements as published on UNICEF website: https://www.unicef.org/supply/technical-specifications-packing-packaging-and-labelling. Bidders are to take note of the changed requirements after 1 October 2023. | Pass/Fail | Yes |  | No
50 |  | Additionally, compliant with the below requirements for secondary and primary packing: - The name or trade name and address of the manufacturer. - The details strictly necessary to identify the device and the contents of the packaging especially for the users; - Where appropriate, the batch code, preceded by the word ‘LOT’, or the serial number; - Where appropriate, an indication of the date by which the device should be used, in safety, expressed as the year and month; - Year and month of manufacture; - Any special storage and/or handling conditions; - Any symbol used on the labelling should be in accordance with ISO 15223-1:2016. | Pass/Fail | Yes |  | No
51 |  | Samples of primary, secondary and tertiary labels are to be included in the submission. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
52 |  | Upload a copy of the Free Sales Certificate for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
53 |  | Upload a copy of the Certificate of Origin for this product under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
54 |  | For sterile products only, deliveries will include a Certificate of Sterilization indicating the batch number(s) of the products included in the delivery. | Pass/Fail | Yes |  | No
55 |  | For applicable product(s) only, deliveries will include a Certificate of Analyses indicating the batch number(s) of the product(s) included in the delivery. | Pass/Fail | Yes |  | No
56 |  | Indicate the Harmonization System code (HS Code) for this product here: | Pass/Fail | No |  | Yes
| L.0 | SECTION L. *** TRAINING, INSTALLATION, UTILISATION AND DECOMMISSIONING ***
57 |  | Either indicate here if, and what, any pre-installation requirements are, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
58 |  | Either indicate whether installation is required, or attach an explanatory document under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
59 |  | Indicate here the required consumables and/or accessories, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
60 |  | Indicate here the required spare parts, their source, reference numbers and indicative prices, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
61 |  | Indicate if there are specific maintenance requirements for users, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
62 |  | Indicate here the specific maintenance requirements for technicians, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Info | Yes |  | Yes
63 |  | The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span: | Info | Yes |  | No
64 |  | Indicate here the specific disposal or decommissioning requirements, or upload the Information under the placeholder: “Technical Docs”. Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
65 |  | Confirm that the Instruction manual for the offered product has been uploaded under the following placeholder for upload: ”Technical Docs". Indicate the filename of the uploaded document(s) here. | Pass/Fail | Yes |  | Yes
M.0 | M.0 | SUSTAINABILITY CONSIDERATIONS *** These questions should be completed for the manufacturer if the supplier is not the manufacturer. *** Any statements should have complimentary documents to support the statement, e.g. claims of a policy being in place, should be support by including the policy document in the submission, statements regarding certifications can only be accepted if the certificates are included in the submission. | Max 20 points
| X2 | SOCIAL CONSIDERATIONS
M.1 |  | Does your Company have a SA 8000 certification? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.2 |  | Does your company invest in community development activities in the markets you source from and/or operate in? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.3 |  | Does your company have in place a documented system to ensure fair workplace and labour practices? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.4 |  | Does your Company have any child labour policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.5 |  | Does your Company have the capacity to meet its client's demand with its own workforce or does it have to integrate with agencies' workers? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.6 |  | Does your company have a system in place to check its employees' identity? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.7 |  | Does your company have a system in place to check its employees' age? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.8 |  | Does your Company have unions/work committees? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.9 |  | Does your Company have any independent certifications on Safety Standards? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.10 |  | Have you signed up with the UN Global Compact initiative? https://www.unglobalcompact.org/. If so provide proof of membership and indicate the filename here. | Score | Yes |  | Yes
| X2 | ENVIRONMENTAL CONSIDERATIONS
M.11 |  | Does your company have a documented environmental management system? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.12 |  | If the manufacturer is either 14001 certified or in the process of obtaining ISO 14001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.13 |  | Does your company have a sustainability/environmental policy? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.14 |  | Does your company have a recycling programme in place? If so, upload a copy under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.15 |  | Is the packaging made of recycled or recyclable materials? | Score | Yes |  | Yes
M.16 |  | Does your company maintain records of potentially hazardous environmental substances in the raw materials used in the product? If so, has your company replaced these substances with non-hazardous ones? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.17 |  | Is the device compliant with RoHS (EU Directive 2002/95/EC) on the restriction of the use of certain hazardous substances in electrical and electronic equipment? If so, provide Information here or upload a document under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes
M.18 |  | Has your company started to monitor its carbon emissions in order to set reduction targets or objectives? Any claims made here require documents(s) to be uploaded under the placeholder "Sustainability" to be accepted. Indicate filenames of any uploaded documents here. | Score | Yes |  | Yes
M.19 |  | Is the manufacturer either 50001 certified or in the process of obtaining ISO 50001 certification. If so, provide expiry date and upload certificate under the placeholder: "Sustainability" and provide the filename here. | Score | Yes |  | Yes