Clarifications for ITB No.UNFPA/UKR/ITB/24/05

Question No.1: We would like to clarify when submitting offers for lots. Please let me know if it is necessary to have a product offer for the entire lot or is it possible to partially submit not all positions in the lot?

Answer: Thank you for your interest in participating in ITB №UNFPA/UKR/ITB/24/05. Please refer to the main document where it is specified that partial bids are allowed under this ITB per Lots only, which means all items within one Lot must be proposed by the bidder, but not all Lots are required in the offer. I hope this clarifies. Please also note that UNFPA will receive technical questions till July 15, 2024 at 11:00 AM (Kyiv time) and will provide the unified response to all bidders at the same time by July 17, 2024 at 3:00 PM (Kyiv time).

Question No.2: I am contacting you regarding the technical requirements for lot 3 (Mobile X-ray system).

1. The system configuration requires an X-ray tube with a stationary anode.

Equipment which we can offer is equipped with an X-ray tube with a rotating anode.

2. A collimator with its own filtering of at least 2 mm in aluminum equivalent is required.

Equipment which we can offer is equipped with a collimator with its own filtering of at least 1.6 mm in aluminum equivalent.

3. Required range of motion:

- rotation of the X-ray tube around the transverse axis ≥ 90° to -30° and around the longitudinal axis ≥ ±110°

- rod rotation around the vertical axis ≥ ±90° and around the transverse axis ≥ 120°.

Our equipment has a completely different principle of mechanical movements. The rotation of the X-ray tube around the longitudinal axis is performed in the range of ≥ ±180°, the rotation of the rod around the vertical axis in the range of ≥ ±270° and with collimator inclinations in the range from -10° to 90°.

Will our offer be rejected if we propose a mobile x-ray system equipped with an x-ray tube with a rotating anode, a collimator with at least 1.6 mm self-filtering aluminum equivalent, and the above-mentioned ranges of movement of the x-ray tube and rod? If our proposal is rejected due to non-compliance with technical requirements, please make appropriate changes to avoid discrimination against other potential bidders

Answer: Thank you for your interest in participating in ITB №UNFPA/UKR/ITB/24/05. UNFPA shall compare all substantially responsive bids per lot to determine the lowest priced substantially responsive bid. Your proposal will not be rejected, but will be evaluated in accordance with the requirements of ITB by the technical expert.

A substantially responsive bid is one that conforms to all the terms, conditions, and specifications of the bidding documents without material deviation, reservation, or omission. A material deviation, reservation, or omission is one that:

a.	affects in any substantial way the scope, quality, or performance of the goods and related services specified in the contract; or

b.	limits in any substantial way, inconsistent with the bidding documents, UNFPA’s rights or the bidder’s obligations under the contract; or

c.	if rectified would unfairly affect the competitive position of other bidders presenting substantially responsive bids.

Question No.3: Can you please kindly confirm that all goods delivered DAP Kyiv will be installed in one location in Kyiv?

Answer: No, after being delivered to Kyiv, the goods will be distributed to other cities or city by UNFPA.

Question No.4: Could you possibly share the exact address for the services? Further, I understand that if we deliver DAP, you will move the goods from a warehouse to the final destination at your own capacity? Or is the delivery to be made to the final destination?

Answer: Exact addresses will be shared with the contractors. UNFPA plans to arrange deliveries by own capacity, but final distribution will be planned on later stages.

Question No.5: I would like to ask if you will kindly accept an offer for products with local service, however no products registration in the country?

Answer: No, UNFPA doesn’t have the capacity to arrange product registration in the country. Given the requirements of Ukrainian legislation for medical devices & equipment it is important that the product which is offered is fully compliant with all the requirements including registration, certification and others.

Question No.6: In Lot No2. Anesthesia workstation with patient ventilator. The technical specification indicates the presence of AG (multigas) and CO2 modules, although the AG (multigas) module by default includes the measurement of the CO2 indicator. Does it mean that only this module AG will be enough, or is there still a need to have both one(AG) and the other(CO2)?

Answer: The Anesthesia workstation has to be equipped with a CO2 measurement module/indicator. It might be a separate CO2 measurement unit  or this unit might be integrated into an AG module.

Question No.7: All equipment requires the presence of a power cord with “F” type plug for medical devices. But this is not mentioned anywhere in any technical document. Does this point require confirmation from the official manufacturer? Or, for example, if our service center independently equips the equipment with plugs of this type, will a warranty letter from our service center be enough?

Answer: Medical equipment in Ukraine uses the power cable with type F plug. It means the type F plug should be provided by a manufacturer of equipment (not by service centre).

Question No.8: In the requirements for this procurement, it is stated that a list of available accessories/spare parts for the equipment must be specified. Does it mean a list of all existing accessories/parts or only those related to the model of the equipment we will be offering?

Answer: The spare parts list is related to the model of equipment that is offered.

Question No.9: Сan we confirm their (accessories) availability by listing them in the commercial offer together with the equipment? And not submit a separate list?

Answer: Yes, you can. You can offer equipment and a list of spare parts to it.

Question No.10: In Lot 8. Surgical table. Usually, in all tables, the leg section is manually adjusted, even if the table is electric. Isn't it?

Answer: Surgical table with manual patient's leg adjustment is acceptable (but just for legs).

Question No.11: All technical documentation must be completed in English, accordingly, references confirming technical compliance are also made to the English-language manual. But it is also required to provide its Ukrainian version. Although the content and pages may differ slightly from the English version. Is this acceptable?

Answer: All technical documentation must be completed in Ukrainian and in English. And each paragraph of the English version must correspond to the Ukrainian one.