PS-10 - Patient scale

**Description**

*Mechanical patient scale. Beam type.*

**Reference Picture**



| Brand / Marque | |  | |
| --- | --- | --- | --- |
| Model / Modèle | |  | |
| Model code as per manufacturer classification / Numéro de référence du produit fournisseur | |  | |
| Manufacturer name / Nom du Fabricant | |  | |
| Manufacturing site / Site du Fabricant | |  | |
| Country of manufacturing / Pays de fabrication | |  | |
| Country of origin / Pays d'origine | |  | |
|  |  |  |  |
|  | **UNOPS Minimum technical requirements // Exigences techniques minimales de l'UNOPS** | **Complies (Yes/No) // Conforme (Oui/Non)** | **Comments/ Technical data //**  **Commentaires/Données techniques** |
| **A** | **Functionality and Performance** |  |  |
| 1 | Mechanical scale, no electronics. |  |  |
| 2 | Steel structure with anticorrosive powder coating, resistant to corrosion and the use of hospital-grade disinfectants. |  |  |
| 3 | Adjustable feet |  |  |
| 4 | Scale at eye level |  |  |
| 5 | Measurement unit: at least kilograms |  |  |
| 6 | Two sliding weights: for gram setting and for kg setting. |  |  |
| 7 | Weighing capacity: at least 150 kg. |  |  |
| 8 | Graduation: at least 100 g. |  |  |
| 9 | Approximate platform dimensions: 30 cm x 40 cm |  |  |
| 10 | Non-slip platform surface |  |  |
| 11 | Integrated telescopic measuring rod. Measurement up to at least 200 cm, 1 mm graduation. |  |  |
| 12 | Preferably reset to zero function |  |  |
| **B** | **Certifications and post sales included in the offered price** |  |  |
| 1 | All the requirements included in the document “Certification, technical documentation and related services requirements: **type II for medical devices**” are applicable. |  |  |
| **C** | **Supplied with** |  |  |
| 1 | Instructions for assembly, use, reset to zero and cleaning in pictograms. |  |  |
| 2 | Tools necessary for assembly. |  |  |
| 3 | Accessories included in the price offered. | N/A |  |
| 4 | Consumables included in the price offered. | N/A |  |
| **D** | **ANY DEVIATIONS OR SUPERIOR CHARACTERISTICS MUST BE INDICATED BELOW // TOUTE DÉVIATION OU CARACTÉRISTIQUE SUPÉRIEURE DOIT ÊTRE INDIQUÉE CI-DESSOUS :** | | |
|  | | |