OTF-33 - Medical bed

**Description**

*Hydraulic dual-sided mattress platform with 3 articulated sections.*

**Reference Picture**

*Note: The images below are intended solely as a guiding support and should be considered as purely indicative and not restrictive of the expected item characteristics*



| Brand / Marque | |  | |
| --- | --- | --- | --- |
| Model / Modèle | |  | |
| Model code as per manufacturer classification / Numéro de référence du produit fournisseur | |  | |
| Manufacturer name / Nom du Fabricant | |  | |
| Manufacturing site / Site du Fabricant | |  | |
| Country of manufacturing / Pays de fabrication | |  | |
| Country of origin / Pays d'origine | |  | |

|  | **UNOPS Minimum technical requirements // Exigences techniques minimales de l'UNOPS** | **Complies (Yes/No) // Conforme (Oui/Non)** | **Comments/ Technical data //**  **Commentaires/Données techniques** |
| --- | --- | --- | --- |
| **A** | **Functionality and Performance / Fonctionnalité et performances** |  |  |
| 1 | Frame:   1. Made of 4 epoxy-coated steel at least 12 gauge, round tube legs including seats for wheels and locking gears; 2. Equipped with plastic bumpers and receptacles for patient-lifting and/or IV poles; 3. ~~All lifting mechanisms and hydraulic pumps support made of epoxy coated, cold-drawn steel.~~ |  |  |
| 2 | Mattress platform:   1. Base consisting of rigid non conductive hard plastic surfaces to allow save cardiopulmonary resuscitation; 2. Back section must be made of radiolucent material and shall include an x-ray cassette holder; 3. Mattress retainers. |  |  |
| 3 | Mattress and pillow:   1. Mattress with a steady (stable) pressure reduction surface at all body points. It shall avoid the accumulation of corporal heat and shall have a removable waterproof and seamless cover in order to keep dry the skin and to avoid skin sores. 2. Mattress thickness of at least 140 mm; 3. High density polyurethane foam interior: at least 28kg/m3; 4. Pillow made of antibacterial, anti-static, toxic-free, high-density foam core, provided with removable and washable hard cotton cover. |  |  |
| 4 | Movements/controls:   1. Back-rest and knee-brake controlled/adjusted by screw-gear based retractable crank handles at the foot end; 2. Calf section with drop-prevention device; 3. Backrest elevation of at least +60º; 4. Knee-brake elevation of at least 20º; 5. Easily removable head & foot ends with plastic-covered bends and grips. |  |  |
|  | **Accessories** |  |  |
| 1 | 3/4 bed size pair of stainless steel folding side-rails. |  |  |
| 2 | Telescopic IV-pole with 4 hooks. |  |  |
| 3 | Blanket support. |  |  |
| 4 | Chart holder. |  |  |
| 5 | Urine/Liquids sack holder. |  |  |
| **B** | **Certifications and post sales included in the offered price** |  |  |
| 1 | The Supplier, included in the offered price, will provide the post sales services described in the document “OT-PS Type II post sales requirements for Medical Devices” |  |  |
| C | **Supplied with** |  |  |
| 1 | Detailed step-by-step instructions for assembly and cleaning in pictograms. |  |  |
| 2 | Tools required for assembly. |  |  |
| **D** | **ANY DEVIATIONS OR SUPERIOR CHARACTERISTICS MUST BE INDICATED BELOW // TOUTE DÉVIATION OU CARACTÉRISTIQUE SUPÉRIEURE DOIT ÊTRE INDIQUÉE CI-DESSOUS\* :** | | |
|  | | |