OTF-23 - Delivery table

**Description**

Hospital delivery bed, adjustable positioning by hydraulic, mechanical or combined mechanism.

**Reference Picture**

*Note: The images below are intended solely as a guiding support and should be considered as purely indicative and not restrictive of the expected item characteristics*



| Brand / Marque | |  | |
| --- | --- | --- | --- |
| Model / Modèle | |  | |
| Model code as per manufacturer classification / Numéro de référence du produit fournisseur | |  | |
| Manufacturer name / Nom du Fabricant | |  | |
| Manufacturing site / Site du Fabricant | |  | |
| Country of manufacturing / Pays de fabrication | |  | |
| Country of origin / Pays d'origine | |  | |

|  | **UNOPS Minimum technical requirements // Exigences techniques minimales de l'UNOPS** | **Complies (Yes/No) // Conforme (Oui/Non)** | **Comments/ Technical data //**  **Commentaires/Données techniques** |
| --- | --- | --- | --- |
| **A** | **Functionality and Performance / Fonctionnalité et performances** |  |  |
| 1 | Manually-operated. |  |  |
| 2 | At least 3 articulated lying sections: backrest, seat and foot. |  |  |
| ~~3~~ | ~~Bumpers at least on two corners~~ |  |  |
| 4 | IV Poles on both sides |  |  |
| 5 | Seat section with oval shape (u-shape or c-shape) to perineum access. |  |  |
| 6 | Leg end section: removable, retractable or sliding under the main section. |  |  |
| 7 | Delivered with all necessary accessories for its good operation |  |  |
| 8 | Suitable for use, and with necessary accessories, at least for the following labor positions:   1. Lithotomy; 2. Side-lying; 3. Sitting or Semi-sitting; 4. On all fours; 5. Squatting; 6. Cesarean section. |  |  |
| 9 | Back section adjustment, minimum range: 0 - 60°. |  |  |
| 10 | Foot or leg supports adjustable horizontally and vertically. |  |  |
|  | **Mechanical Requirements** |  |  |
| 11 | Frame and bed platform construction material: Epoxy or polyester powder coating steel with highly resistant to aggressive corrosive agents such as active chlorine. Or stainless steel AISI 304/304L or other stainless steel with higher performance & corrosion resistance. |  |  |
| 12 | Preferably: detachable head and footboard, material: rigid plastic, ABS or similar. |  |  |
| 13 | Mounted on 4 supports, all finished with non-slip rubber or plastic. |  |  |
| 14 | Two dedicated rails for attaching and detaching the accessories. Made of steel, aluminum, stainless steel AISI 304/304L or better. |  |  |
| 15 | All parts of the bed (including accessories) must be designed for frequent clean and disinfection with hospital-grade. |  |  |
|  | **Mattress** |  |  |
| 16 | At least two separate and removable sections |  |  |
| 17 | Minimum thickness (mattress with cover): 125 mm in backrest and seat section. |  |  |
| 18 | Polyurethane foam or polyurethane high density |  |  |
| 19 | Minimum density: 28 kg/m3 |  |  |
| 20 | Designed to pressure injury prevention (ulcers) |  |  |
| 21 | Mattress cover (antistatic, impermeable and washable material) resistant to corrosive agents such as active chlorine. |  |  |
| 22 | Resistant to corrosion, water, detergent soap, ethyl alcohol solution with or without nitrite and to the hypochlorite of sodium. |  |  |
| 23 | Recommended by the bed manufacturer. |  |  |
|  | **Dimensions and weight capacity** |  |  |
| 24 | Mattress platform dimension (LxWxH): 190x65x80 cm, +/- 15%. |  |  |
| 25 | Maximum patient weight (MPW): 165 kg or higher. |  |  |
| 26 | Safe Working Load (SWL): 210 kg or higher. |  |  |
|  | **Accessories included in the offered price** |  |  |
| 27 | Two hand grips. |  |  |
| 28 | Removable stainless steel fluid basin. |  |  |
| 29 | Height adjustable IV Pole with at least two hooks. |  |  |
| 30 | Two (2) padded leg holders/Goepel holders:   * 1. Mounted on stainless steel rods;   2. Lockable and removable, with clamps for fixation if they are necessary;   3. Pad waterproof and cleanable resistant to hospital-grade disinfectant products;   4. Foam padded or similar;   5. Adjustable in height and width. |  |  |
| 31 | Birthing stool/seat u shape to support upright birth, see picture. That can be used as a stand-alone product. |  |  |
| 32 | Two bobath balls of 65 and 95 cm, inflatable, made of dense PVC. |  |  |
| **B** | **Certifications and post sales included in the offered price** |  |  |
| 1 | The Supplier, included in the offered price, will provide the post sales services described in the document “OT-PS Type II post sales requirements for Medical Devices”. |  |  |
| C | **Supplied with** |  |  |
| 1 | Detailed step-by-step instructions for assembly and cleaning in pictograms. |  |  |
| 2 | Tools required for assembly. |  |  |
| **D** | **ANY DEVIATIONS OR SUPERIOR CHARACTERISTICS MUST BE INDICATED BELOW // TOUTE DÉVIATION OU CARACTÉRISTIQUE SUPÉRIEURE DOIT ÊTRE INDIQUÉE CI-DESSOUS\* :** | | |
|  | | |