OTS-26 - Obstetrics suction cup

Description

Suction cup for obstetrics, to assist childbirth. With reusable cup

Reference Picture

Note: The images below are intended solely as a guiding support and should be considered as purely indicative and not restrictive of the expected item characteristics

| Brand / Marque |  |  |  |
| --- | --- | --- | --- |
| Model / Modèle |  |  |  |
| Model code as per manufacturer classification / Numéro de référence du produit fournisseur |  |  |  |
| Manufacturer name / Nom du Fabricant |  |  |  |
| Manufacturing site / Site du Fabricant |  |  |  |
| Country of manufacturing / Pays de fabrication |  |  |  |
| Country of origin / Pays d'origine |  |  |  |

|  | UNOPS Minimum technical requirements // Exigences techniques minimales de l'UNOPS | Complies (Yes/No) // Conforme (Oui/Non) | Comments/ Technical data // Commentaires/Données techniques |
| --- | --- | --- | --- |
| A | Functionality and Performance / Fonctionnalité et performances |  |  |
| 1 | Hand-operated vacuum pump. |  |  |
| 2 | Reusable. |  |  |
| 3 | Vacuum level at least up to 580 mmHg. |  |  |
| 4 | Manometer with green and red zones clearly identified. |  |  |
| 5 | Vacuum release security valve. |  |  |
| 6 | Vacuum gauge. |  |  |
| 7 | All parts must be removable. disinfectable with hospital grade products and autoclavable. |  |  |
|  | Accessories included |  |  |
| 1 | Four (4) Silicone reusable vacuum delivery cups, preferably two different sizes. |  |  |
| 2 | Four (4) reusable or single use tubings for the cups. |  |  |
| 3 | One (1) pump maintenance kit, if applicable. |  |  |
| 4 | If tubing is single use, at least two hundred (200) units must be provided. |  |  |
| 5 | Storage box. |  |  |
| B | Certifications and post sales included in the offered price |  |  |
| 1 | The Supplier, included in the offered price, will provide the post sales services described in the document “OT-PS Type II post sales requirements for Medical Devices”. |  |  |
| C | Supplied with |  |  |
| 1 | Detailed step-by-step instructions for assembly, use and cleaning in pictograms. |  |  |
| D | ANY DEVIATIONS OR SUPERIOR CHARACTERISTICS MUST BE INDICATED BELOW // TOUTE DÉVIATION OU CARACTÉRISTIQUE SUPÉRIEURE DOIT ÊTRE INDIQUÉE CI-DESSOUS* : |  |  |