OTS-25 - Sphygmomanometer

**Description**

*Aneroid blood pressure monitor (Sphygmomanometer) with cuff set*

**Reference Picture**

*Note: The images below are intended solely as a guiding support and should be considered as purely indicative and not restrictive of the expected item characteristics*

|  |  |
| --- | --- |

| Brand / Marque | |  | |
| --- | --- | --- | --- |
| Model / Modèle | |  | |
| Model code as per manufacturer classification / Numéro de référence du produit fournisseur | |  | |
| Manufacturer name / Nom du Fabricant | |  | |
| Manufacturing site / Site du Fabricant | |  | |
| Country of manufacturing / Pays de fabrication | |  | |
| Country of origin / Pays d'origine | |  | |

|  | **UNOPS Minimum technical requirements // Exigences techniques minimales de l'UNOPS** | **Complies (Yes/No) // Conforme (Oui/Non)** | **Comments/ Technical data //**  **Commentaires/Données techniques** |
| --- | --- | --- | --- |
| **A** | **Functionality and Performance / Fonctionnalité et performances** |  |  |
| 1 | For adult and pediatric patients. |  |  |
| 2 | Body material: stainless steel or metal with anticorrosive treatment, or ABS or similar plastic . |  |  |
| 3 | Measurement range of at least 0 to 300 mmHg. |  |  |
| 4 | Accuracy: +/- 3 mmHg or better |  |  |
| 5 | Gauge (measuring clock) graduated from 0 to 300 mmHg (minimum), resolution: at least 2 mmHg. Dimension of at least 10 x 10 cm or 10 cm in diameter. |  |  |
| 6 | Two (2) rubber bulbs, with manual pressure release valve. |  |  |
| 7 | Spiral rubber hose. |  |  |
| 8 | Two (2) sets of at least five (5) reusable, hypoallergenic cuffs with tubes each: Infant, Child, Small adult, Adult and Large adult. With Velcro fastening. Ten (10) units in total. |  |  |
| 9 | Latex and mercury free design. |  |  |
| 10 | Mobile cart mounted, with basket to house accessories. Structure material: epoxy coated tubular steel (or similar finish), resistant to hospital-grade disinfectant products; or stainless steel. |  |  |
| 11 | Rolling base with at least 5 swivel castors, made of polypropylene, nylon, rubber or similar material. Diameter: at least 50 mm. |  |  |
| 12 | Complete Stethoscope |  |  |
| 13 | All materials resistant to corrosion and the use of hospital-grade disinfectants. |  |  |
| **B** | **Certifications and post sales included in the offered price** |  |  |
| 1 | The Supplier, included in the offered price, will provide the post sales services described in the document “OT-PS Type II post sales requirements for Medical Devices”. |  |  |
| C | **Supplied with** |  |  |
| 1 | Detailed step-by-step instructions for assembly, use and cleaning in pictograms.  Tools required for assembly. |  |  |
| **D** | **ANY DEVIATIONS OR SUPERIOR CHARACTERISTICS MUST BE INDICATED BELOW // TOUTE DÉVIATION OU CARACTÉRISTIQUE SUPÉRIEURE DOIT ÊTRE INDIQUÉE CI-DESSOUS\* :** | | |
|  | | |