United Nations Population Fund

Website: www.unfpa.org

. M

14-08-2024

INVITATION TO BID

ITB No. UNFPA/DNK/ITB/24/014

MANUFACTURE AND/OR SUPPLY OF PRODUCTS AND RELATED SERVICES

INTRODUCTORY LETTER

Dear Sir/Madam,

- The United Nations Population Fund (UNFPA), an international development agency, invites sealed bids for the supply of:

| Item No & (internal number) | Item short description | Qty |
| --- | --- | --- |
| Item 1 (OR 1). | Activated Clotting Time (ACT) Measuring device | 2 |
| Item 2 (OR 2). | Portable Blood Gas Analyzer | 2 |
| Item 3 (OR 3). | Extra Corporeal Circulation (ECC) Machine | 1 |
| Item 4 (OR 4). | Hypo-hyperthermia Machine | 1 |
| Item 5 (OR 33). | Patient Auto-Transfusion System | 1 |
| Item 6 (OR 5). | Infusion Rack (Syringe Pump and Infusion Pump including racks and trolley with holder pole with hooks) | 4 |
| Item 7 (OR 10). | Monitor/defibrillator + specific spoons for children + specific options + accessories | 2 |
| Item 8 (OR 11). | Suction system, surgery | 4 |
| Item 9 (OR 12). | Cardiovascular ultrasound imaging system | 1 |
| Item 10 (OR 16). | Electrosurgical system | 1 |
| Item 11 (OR 18). | Hand-held ultrasound imaging system | 1 |
| Item 12 (OR 19). | Brain Monitoring System | 1 |
| Item 13 (OR 20). | Monitor for cerebral oximetry | 1 |
| Item 14 (OR 22). | Multiparametric TRANSPORT patient monitor for adult and pediatric patients | 4 |
| Item 15 (OR 23). | Mounting system for monitoring arm or lateral support for ventilation machine | 1 |
| Item 16 (OR 24). | Slave Monitor 28’’+ Wall fixation Arm | 1 |
| Item 17 (OR 26). | Lamp, operating room, ceiling | 1 |
| Item 18 (OR 27). | Video-laryngoscope | 1 |
| Item 19 (OR 28). | Patient heater | 2 |
| Item 20 (OR 29). | Blood / fluid warmer | 2 |
| Item 21 (OR 31). | Patient Ventilator/Respirator | 1 |
| Item 22 (OR 32). | Portable external pacemaker | 3 |
| Item 23 (OR 34). | Patient drainage system | 1 |
| Item 24 (OR 37). | Sternal, oscillating saw | 2 |
| Item 25 (OR 38). | Surgeon's headlamp | 1 |
| Item 26 (ICU 3). | Dialysis Equipment for Neonate - Peritoneal dialysis system | 1 |
| Item 27 (ICU 18). | Anesthetic Monitor & Accessories | 1 |
| Item 28 (ICU 26). | ICU Emergency & Transport ventilator with Acc & Options: “IPPV, SIMV, CPAP/ASB, Bibap, Mes. CO2” | 1 |
| Item 29 (ICU 27). | ICU Ventilators, ADU/PED, mode, «IPPV, SIMV, CPAP /ASB, BIBAP Ventilation » | 3 |
| Item 30 (ICU 28). | Lamp, examination | 3 |
| Item 31 (ICU 29). | Non-invasive ventilator | 1 |

for its programme in Mongolia.

- Bidding shall be conducted through ONE envelope. The technical bid containing the technical specifications and the financial bid containing price information shall be submitted together.

- The Bidder shall not be required to quote for all items. However, Bidders are encouraged to quote for as many items as possible.

- To enable you to submit a bid, please read the following attached documents carefully:

| Section I: | Instructions to Bidders |
| --- | --- |
| Section II: | Technical Specifications and Schedule of Requirements |
| Section III: | UNFPA General Conditions of Contract |
| Section IV: | UNFPA Special Conditions for Contracts |
| Section V: Section VI: | Bidding Forms Contract Forms |

- The bid shall reach UNFPA’s reception or the email inbox of bidtender@unfpa.org no later than 11th September 2024, at 10:00AM Copenhagen time .

- The bid shall be opened on 16th of September 2024, at 10:00AM Copenhagen time. Bidders or their authorized representatives may attend the bid opening.  Kindly confirm by email by 21-08-2024 whether your company shall be represented at the bid opening.

- Bids received after the stipulated date and time shall not be accepted under any circumstances. Bids delivered through courier and posted later than the due date shall not be registered and shall be returned unopened or shall be shredded. Bids submitted to any other email address than bidtender@unfpa.org shall be rejected.

- Bidders shall acknowledge receipt of this Invitation to Bid according to the Bid Confirmation Form,  Section V, 1 of this solicitation document by email to Tsetsenbaatar Batsuuri, email: batsuuri@unfpa.org  no later than 21-08-2024 and to indicate whether or not a bid shall be submitted. If you are declining to bid please state the reasons for UNFPA to improve its effectiveness in future invitations.

- Any questions relating to the attached documents shall be addressed in writing to the following UNFPA personnel:

- Tsetsenbaatar Batsuuri, email: batsuuri@unfpa.org

- Mirza Trozic, email: trozic@unfpa.org

Do not submit your bid to these contacts, or your bid will be disqualified.

- This letter is not to be construed in any way as an offer to contract with your firm.

- UNFPA strongly encourages all Bidders to register on the United Nations Global Marketplace (http://www.ungm.org). The UNGM is the procurement portal of the United Nations system. By registering on UNGM, vendors become part of the database that UN buyers use when searching for suppliers. Vendors can also access all UN tenders online and, by subscribing to the Bid Tender Service, vendors can be automatically notified via email of all UN business opportunities that match the products and services for which they have registered. Instructions on how to subscribe to the Tender Alert Service can be found in the UNGM Interactive Guide for Suppliers.

http://www.ungm.org/Publications/UserManuals/Suppliers/UserManual_Supplier.pdf .

Yours sincerely,

Tsetsenbaatar Batsuuri

UNFPA

UNITED NATIONS POPULATION FUND

INVITATION TO BID

ITB NO.: UNFPA/DNK/ITB/24/014

Bid document for the manufacture and/or supply of products and related services

14th of August 2024

Table of Contents

SECTION I: Instructions to Bidders	7

A.	Introduction	7

1.	Scope	7

2.	Eligible Bidders	8

3	Eligible Goods and Related Services	8

4	Cost of Bid	8

5	Fraud and Corruption	9

B.	Solicitation Documents	9

6	UNFPA Solicitation document	9

7	Clarifications of solicitation document	9

8	Amendments to UNFPA bid solicitation document	9

C.	Preparation of Bids	10

9	Language of the bid	10

10	Documents to be submitted with the bid	10

11	Bid Currency and Prices	11

12	Validity of Bid	12

D.	Submission of Bids and Bid Opening	12

13	Partial Bids	12

14	Alternative Bids	12

15	Bids	12

16	Electronic Submissions	13

17	Bid Submission Deadline/Late Bids	13

18	Withdrawal, Substitution and Modification of Bids	14

19	Storage of Bids	14

20	Bid Opening	14

E.	Evaluation and Comparison of Bids	15

21	Confidentiality	15

22	Clarification of Bids	15

23	Responsiveness of bids	15

24	Nonconformities, Errors, and Omissions	16

25	Preliminary examination of Bids	17

26	Examination of Terms and Conditions and Technical Evaluation	17

27	Conversion to Single Currency	17

28	Domestic Preference	17

29	Evaluation of Bids	17

30	Comparison of Price Bids	18

31	Post-qualification of the Bidder	18

32	UNFPA’s Right to Accept Any Bid and to Reject Any or All Bids	20

33	UNFPA’s Right to Annul a Bidding Process	20

F.	Award of Contract	20

34	Award Criteria	20

35	Right to Vary Requirements at Time of Award	21

36	Signing of the contract	21

37	Publication of Contract Award	21

SECTION II: Technical Specifications and Schedule of Requirements	22

2.1.	Technical Specifications	22

2.2.	Schedule of Requirements	86

SECTION III: UNFPA General Conditions of Contract	89

SECTION IV: UNFPA Special Conditions for Contracts	90

SECTION V: Bidding Forms	92

1. Bid Confirmation Form	98

2. Bid Submission Form	99

3. Bidders Identification Form	100

4. Performance Statement Form	102

5.	Product Item Overview Form	103

6. Price Schedule Form	104

7. Joint Venture Partner Information Form	106

SECTION VI: Contract Forms	107

1.	Bank Guarantee for Advance Payment	108

2.	Performance Security	109

3.	Contract Forms	110

# SECTION I: Instructions to Bidders

# Introduction

## Scope

  - The goods and related services to be procured are:

| Item No | Item short description | Qty |
| --- | --- | --- |
| Item 1 (OR 1). | Activated Clotting Time (ACT) Measuring device | 2 |
| Item 2 (OR 2). | Portable Blood Gas Analyzer | 2 |
| Item 3 (OR 3). | Extra Corporeal Circulation (ECC) Machine | 1 |
| Item 4 (OR 4). | Hypo-hyperthermia Machine | 1 |
| Item 5 (OR 33). | Patient Auto-Transfusion System | 1 |
| Item 6 (OR 5). | Infusion Rack (Syringe Pump and Infusion Pump including racks and trolley with holder pole with hooks) | 4 |
| Item 7 (OR 10). | Monitor/defibrillator + specific spoons for children + specific options + accessories | 2 |
| Item 8 (OR 11). | Suction system, surgery | 4 |
| Item 9 (OR 12). | Cardiovascular ultrasound imaging system | 1 |
| Item 10 (OR 16). | Electrosurgery generator | 1 |
| Item 11 (OR 18). | Hand-held ultrasound imaging system | 1 |
| Item 12 (OR 19). | Brain Monitoring System | 1 |
| Item 13 (OR 20). | Monitor for cerebral oximetry | 1 |
| Item 14 (OR 22). | Multiparametric TRANSPORT patient Monitor for adult and pediatric patients | 4 |
| Item 15 (OR 23). | Mounting system for monitoring arm or lateral support for ventilation machine | 1 |
| Item 16 (OR 24). | Slave Monitor 28’’, + Wall fixation Arm | 1 |
| Item 17 (OR 26). | Lamp, operating room, ceiling | 1 |
| Item 18 (OR 27). | Video-laryngoscope | 1 |
| Item 19 (OR 28). | Patient heater | 2 |
| Item 20 (OR 29). | Blood / fluid warmer | 2 |
| Item 21 (OR 31). | Patient Ventilator/Respirator | 1 |
| Item 22 (OR 32). | Portable external pacemaker | 3 |
| Item 23 (OR 34). | Patient drainage system | 1 |
| Item 24 (OR 37). | Sternal, oscillating saw | 2 |
| Item 25 (OR 38). | Surgeon's headlamp | 1 |
| Item 26 (ICU 3). | Dialysis Equipment for Neonate - Peritoneal dialysis system | 1 |
| Item 27 (ICU 18). | Anesthetic Monitor & Accessories | 1 |
| Item 28 (ICU 26). | ICU Emergency & Transport ventilator with Acc & Options: “IPPV, SIMV, CPAP/ASB, Bibap, Mes. CO2” | 1 |
| Item 29 (ICU 27). | ICU Ventilators, ADU/PED, mode, «IPPV,SIMV, CPAP /ASB, BIBAP Ventilation » | 3 |
| Item 30 (ICU 28). | Lamp, examination | 3 |
| Item 31 (ICU 29). | Non-invasive ventilator | 1 |

for UNFPA’s Programme located in Mongolia.

## Eligible Bidders

  - This bid is open to primary manufacturers, authorized agents and authorized resellers.

  - A Bidder and all parties constituting the Bidder may hold any nationality.

  - A Bidder shall not have a conflict of interest. All Bidders found to have a conflict of interest shall be disqualified. Bidders may be considered to have a conflict of interest with one or more parties in this bidding process, if they:

  - Are or have been associated in the past, with a firm or any of its affiliates that have been engaged by UNFPA to provide consulting services for the preparation of the design, specifications, and other documents to be used for the procurement of the goods to be purchased under these bidding documents; or

  - A Bidder that is under a declaration of ineligibility by UNFPA in accordance with Instructions to Bidders Clause 2 at the date of contract award shall be disqualified.  Bidders shall not be eligible to submit a bid if at the time of bid submission:

- The Bidder is listed as suspended on United Nations Global Marketplace (http://www.ungm.org) as a result of having committed fraudulent activities,

- The Bidder’s name is mentioned in the UN 1267 list issued by the Security Council resolution 1267 that establishes a sanctions regime to cover individuals and entities associated with Al-Qaida and/or the Taliban;

- The Bidder is debarred by the World Bank Group

  - Bids may be submitted by a Joint Venture (JV). In the case of a JV:

- The duly filled Joint Venture Partner Information Form, Section V, 7,, must be included with the bid; and

- All parties to the JV shall be jointly and severally liable; and

- The JV shall nominate a Representative who shall have the authority to conduct all businesses:

  - for and on behalf of any and all the parties of the JV during the bidding process; and

  - in the event the JV is awarded the contract, during contract execution.

## Eligible Goods and Related Services

- All the goods and related services to be supplied under the contract may have their origin in any country.

- For purposes of this Clause, the term “origin” means the country where the goods have been produced, manufactured or processed; or, through manufacture, processing, or assembly, another commercially recognized article results that differs substantially in its basic characteristics from its components.

## Cost of Bid

- The Bidder shall bear all costs associated with the preparation and submission of the bid, and the procuring UN entity shall in no case be responsible or liable for those costs, regardless of the conduct or outcome of the bid.

## Fraud and Corruption

- UNFPA’s policy regarding fraud and corruption is available at http://www.unfpa.org/about-procurement#FraudCorruption and applies fully to this Invitation to Bid. The submission of any offer implies that the Bidder is aware of this policy.

# Solicitation Documents

## UNFPA Solicitation document

- Bidders are expected to examine all instructions, forms, specifications, terms and conditions contained within this UNFPA solicitation document. Failure to comply with these documents shall be at the Bidder’s risk and may affect the evaluation of the bids, or may result in the rejection of the bid.

- Bidding documents consist of the following:

| Section I: | Instructions to Bidders |
| --- | --- |
| Section II: | Technical Specifications and Schedule of Requirements |
| Section III: | UNFPA General Conditions of Contract |
| Section IV: | UNFPA Special Conditions for Contracts |
| Section V: Section VI: | Bid Forms Contract Forms |

- Bidders are cautioned to read the specifications carefully (see Section II Technical Specifications and Schedule of Requirements), as there may be special requirements. The technical specifications presented herein are not to be construed as defining a particular manufacturer’s product. Bidders are encouraged to advise UNFPA if they disagree.

- The specifications are the minimum requirements for the products and related services. Products and services offered must meet or exceed all requirements herein. The products shall conform in strength, quality and workmanship to the accepted standards of the relevant industry. Modifications of or additions to basic standard products of less size or capability to meet these requirements will not be acceptable.

## Clarifications of solicitation document

- A prospective Bidder requiring any clarification on the bid solicitation documents may notify UNFPA in writing within two weeks from the date of issue of the bid. UNFPA shall respond in writing to any request for clarification received and circulate its response (including an explanation of the query but without identifying the source of enquiry) to all prospective Bidders who have received the bid solicitation documents. A copy of UNFPA’s answer shall also be posted on the UN Global Marketplace, http://www.ungm.org/.

## Amendments to UNFPA bid solicitation document

- At any time prior to the deadline for submission of bids, UNFPA may for any reason, whether at its own initiative or in response to a clarification requested by a prospective Bidder, modify the bidding documents by amendment.

- All prospective Bidders that have received the bidding documents shall be notified in writing of all the amendments to the bidding documents. In order to give prospective Bidders reasonable time to take the amendments into account in preparing their bids UNFPA may, at its discretion, extend the deadline for the submission of bids.

# Preparation of Bids

## Language of the bid

- The bid prepared by the Bidder and all correspondence and documents relating to the bid shall be written in English.

## Documents to be submitted with the bid

- Documents Establishing the Eligibility of the Bidder

To establish their eligibility, Bidders shall:

- Complete the Bid Submission Form, Section V, 2.

- Complete Bidders Identification Form, Section V, 3.

- Complete Joint Venture Partner Information Form, Section V, 7 and provide all documents as required in the Form in the event that the bid is submitted by a Joint Venture.

- Documents Establishing the Qualifications of the Bidder

To establish its qualifications, the Bidder shall submit to UNFPA’s satisfaction the following documents:

- Evidence that the Bidder is established as a company and legally incorporated in the country where it resides; e.g. through provision of certification of incorporation or other documentary evidence (this is not required for companies already registered in national, regional or international Stock Exchanges);

- Copy of valid manufacturing license from the country of manufacturing and/or a copy of company registration in the country of operation demonstrating that is duly authorized to supply these goods to the country of destination.

- In the case of a Bidder not doing business within the country of destination, the Bidder is or will be represented by an Agent in the country that is equipped and able to carry out the supplier’s maintenance, training, repair and spare parts-stocking obligations prescribed in the Section II, Technical Specifications and Schedule of Requirements

- Written confirmation from the Bidder that the Bidder is neither suspended by the United Nations system nor debarred by the World Bank group;

- The availability in the Beneficiary’s Country of spare parts and after-sales services for the equipment offered in the bid.

- Post qualification documentation outlined in Instructions to Bidders, Section 32

Failure to furnish all the information required for submission shall be at the Bidder’s risk as it may then be determined that the bid does not substantially respond to the UNFPA bid document in every respect.  This may result in a rejection of the bid.

- Documents Establishing the Eligibility and Conformity of the Goods and Related Services

Bidders shall submit:

- Documentary evidence that the goods conform to the Technical Specifications and standards specified in Section II Technical Specifications and Schedule of Requirements.

- Completed Product Item Overview Form, Section V, 5.

- Product catalogues containing pictures of the product(s)

- Manufacturer’s technical product specifications or datasheets

- Results of any testing carried out on the products

- Copies of current certificates such as GMP/quality, FSC/CPP, valid manufacturer’s ISO certificate per declared product, proof of compliance with EU Council Directive 93/42/EEC (MDD) or EU Regulations 2017/745 (MDR) with valid EC Certificate, based on the risk of the devices, USFDA 510k registration authorization, Japan QS standard, or other market authorization issued by a Stringent Regulatory Authority member of the IMDRF (Australia or Canada), proof of compliance with other relevant applicable safety, electrical and performance standards (test reports), as stated in the Technical Specifications and Schedule of Requirements Section II

- The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the goods during 5 years  following commencement of the use of the goods by UNFPA. Bidders must complete and submit with their bid the Excel table containing the individual item details, as per Form in Section V.5. Bidding Forms.

- Manufacturer’s warranty letter indicating the time of the warranty duration

- A copy of the user, service, spare parts manual and installation (when applicable), if not provided with the submission of the documents, a commitment letter to deliver such documents, is applicable

- Training program with the details of the activities and duration of the training, relevant topics. The training must be delivered to users and the technical service team

- Items requiring pre-installation (such as the operating light): the workshop layout for the installation of the light must be provided, and a commitment letter to ensure at least one field visit prior to the installation of the light will be perform by the potential bidder, to assure and confirm that operating theather fulfill all the requirements for the proper installation of the device: e.g: verification of ceiling support, sufficient space for the light arm's motions, connection to the electrical and safety protection system (emergency system support)

- Documents Establishing Sustainability Efforts of the Bidder

UNFPA requests Bidders to submit information on environmental and social policies and any related documentation in their bid.. In the long term it is UNFPA’s intention to incorporate environmental and social criteria considerations into the evaluation process, such as adherence to Global Compact requirements. More information can be accessed on the Global Compact web site, http://www.unglobalcompact.org/, or by contacting Procurement Services Branch at procurement@unfpa.org.  UNFPA encourages suppliers now to consider joining the UN Global Compact and to look into other ways to help reduce their environmental impact.

## Bid Currency and Prices

- All prices shall be quoted in US Dollars (USD).

- The Bidder shall indicate the unit prices (where applicable) and total bid price of the goods or services it proposes to supply under the contract. This price information shall be indicated on the Price Schedule Form, Section V, 6.

- Bidders are requested to quote the following based on INCOTERMS 2020:

- Price of goods FCA point of Departure and DAP to National Center for Maternal and Child Healthcare (NCMCH), “Huvisgalchdiin” street, Bayangol district, Ulaanbaatar 16060, Mongolia.

- Unloading and offloading the goods into rooms will be the suppliers’ responsibility.

- The terms FCA, DAP and other similar terms shall be governed by the rules prescribed in the INCOTERMS 2020, published by the International Chamber of Commerce.

- Where installation, commissioning, training or other similar services are required to be performed by the Bidder, the Bidder shall include an itemized list of the prices for the requested.

## Validity of Bid

- The prices of the bid shall be valid for 90 days after the closing date of bid submission as specified by UNFPA. A bid valid for a shorter period shall be rejected by UNFPA on the grounds that it is non-responsive.

- In exceptional circumstances, UNFPA may solicit the Bidder's consent for an extension of the period of validity under exceptional circumstances. The request and the responses shall be made in writing.

# Submission of Bids and Bid Opening

## Partial Bids

- Partial bids are allowed under this tender. Bidders can quote for all or some medical devices, however bidders must quote for full quantities per each item requested. UNFPA reserves the right to select and accept a part or parts of any bid. Bidders that will submit Bids for the majority of items may get priority at the bid evaluation and the contract awards stages.

## Alternative Bids

- Alternative bids will not be accepted. In the event of a supplier submitting more than one bid, the following shall apply:

- All bids marked alternative bids will be rejected and only the base bid will be evaluated.

- All bids will be rejected if no indication is provided as to which bids are alternative bids.

## Bids

- Bids shall be submitted in one envelope or transmitted in an email to a secure email address designated by UNFPA.

- The technical portion of the bid shall be prepared in accordance with Section II: Schedule of Requirements and Technical Specifications and shall include the requested documentation as per Instructions to Bidders Clause 10.

- The financial portion of the bid shall be prepared in accordance with the Price Schedule Form in Section V, 6 of the bid forms.

- Bids shall be signed by the Bidder or a person or persons duly authorized to bind the Bidder to the contract. A bid shall contain no interlineations, erasures, or overwriting except as necessary to correct errors made by the Bidder. In that case such corrections shall be initialled by the person or persons signing the bid.

## Electronic Submissions

- Bids may be submitted electronically. Please note the following guidelines for electronic submissions:

- Bidders shall make clear reference to the specific bid in the subject field as instructed, otherwise bids may be rejected. Clearly specify the following text in the subject line: ITB No. UNFPA/DNK/ITB/24/014, Bidder’s Name.

- The bid shall be submitted to bidtender@unfpa.org. Bids received at the bidtender@unfpa.org mailbox are kept undisclosed and shall not be opened before the scheduled opening date. Sending to any other email address will violate confidentiality and invalidate the bid.

- Email submission shall not exceed 10 MB, including the size of the cover email and attachment. It is recommended that all the bidding documents are consolidated into as few attachments as possible which shall be in commonly used file formats. If the bid consists of large electronic files, it is recommended to send these files separately before the deadline indicating the order of emails (email 1, email 2, etc.) after the bid reference number and the Bidder’s name in the subject line of each email.

It shall be the Bidder’s responsibility to ensure that bids sent by email are received by the deadline. All Bidders shall receive an auto-reply acknowledging the receipt of their email.  Bidders will receive an auto-reply acknowledging receipt of the first email. In the body of this first email, bidders are requested to list the number of messages which make up their technical offer and the number of messages which make up their financial offer. If you do not receive any auto-reply from UNFPA’s email system, please inform Tsetsenbaatar Batsuuri batsuuri@unfpa.org. Submission documents should be converted to PDF, submissions containing links to the files which need to be downloaded will not be considered. Bidders shall not receive responses to questions sent to bidtender@unfpa.org since it is a secure mailbox.

- In order to avoid last minute internet congestion it is recommended to send your bid as early as possible before the deadline.

## Bid Submission Deadline/Late Bids

- Bids must be delivered to the office on or before the date and time specified in the introductory letter of this solicitation document.  If any doubt exists as to the time zone in which the bid should be submitted please refer to www.timeanddate.com/worldclock, or contact the bid focal point.

- UNFPA may, under special and exceptional circumstances, extend the bid submission deadline and such changes shall be notified in UNGM before the expiration of the original period.

- Any bid received by UNFPA after the bid submission deadline shall be rejected. UNFPA shall not be legally responsible for bids that arrived late due to the Bidder’s problems with transmission of bid submissions via email and/or with the courier company.

## Withdrawal, Substitution and Modification of Bids

- A Bidder may withdraw, substitute, or modify its bid after it has been submitted by sending a written notice prior to the bid submission deadline. The modification shall be submitted in a sealed envelope or to the dedicated secured email.

- The Bidder may withdraw its bid after submission, provided that written notice of the withdrawal is received by UNFPA prior to the bid submission deadline requested to be withdrawn shall be shredded or shall be returned unopened to the Bidder.

- No bid may be withdrawn, substituted, or modified in the interval between the deadline for submission of bids and the expiration of the period of bid validity specified by the Bidder on the Bid Submission Form or any extension thereof.

## Storage of Bids

- Bids received prior to the deadline of submission and the time of opening shall be securely kept unopened until the specified bid opening date stated in the UNFPA’s solicitation document. No responsibility shall be attached to UNFPA for prematurely opening an improperly addressed and/or identified bid.

## Bid Opening

- UNFPA shall conduct the bid opening in public via online Teleconferencing on following date and time.

Date:  16th September 2024

Time: 10:00AM Copenhagen time, (reference: www.timeanddate.com/worldclock).

- Bids received electronically by the required deadline will be printed and a copy of the bids will be put in a sealed envelope that will be opened at the time and date specified in the bid document. Only the last received bid will be opened if multiple bids are sent by a same Bidder.

- UNFPA shall open all bids in the presence of at least two witnesses from UNFPA or another UN agency. The bids shall be opened publicly at the time and place specified in the ITB and an immediate record made thereof.

- Only those who have submitted bids may attend the bid opening. However, the Bidders may authorize a local agent, embassy or trade commission (also referred to as observers) to represent them. In order to be able to attend bid opening, agents representing Bidders must provide reasonable evidence (business cards, letter of authorization, etc.) confirming the name of the Bidder they represent.

- The report shall be available for viewing by Bidders for a period of thirty days from the date of the opening. No information that is not included in the bid opening report can be given to Bidders.

- No bid shall be rejected at bid opening, except for late bids. Bids that are not opened and read out at the bid opening shall not be considered further for evaluation, irrespective of the circumstances. Withdrawn bids will be shredded except for any bank securities, which will be returned to the Bidder.

# Evaluation and Comparison of Bids

## Confidentiality

- Information relating to the examination, evaluation, comparison, and post-qualification of bids, and recommendation of contract award shall not be disclosed to Bidders or any other persons not officially concerned with such process until the contract award is published.

- Any effort by a Bidder to influence UNFPA in the examination, evaluation, comparison, and post-qualification of the bids or contract award decisions may result in the rejection of its bid.

- Notwithstanding from the time of bid opening to the time of contract award, if any Bidder wishes to contact UNFPA on any matter related to the bidding process, it should do so in writing.

## Clarification of Bids

- To assist in the examination, evaluation and comparison of bids, UNFPA may ask Bidders for clarification of their bids. The request for clarification and the response shall be in writing by UNFPA and no change in price or substance of the bid shall be sought, offered or permitted.

## Responsiveness of bids

- UNFPA’s determination of a bid’s responsiveness is to be based on the contents of the bid itself.

- A substantially responsive bid is one that conforms to all the terms, conditions, and specifications of the bidding documents without material deviation, reservation, or omission. A material deviation, reservation, or omission is one that:

- affects in any substantial way the scope, quality, or performance of the goods and related services specified in the contract; or

- limits in any substantial way, inconsistent with the bidding documents, UNFPA’s rights or the Bidder’s obligations under the contract; or

- if rectified would unfairly affect the competitive position of other Bidders presenting substantially responsive bids.

- UNFPA considers material deviation to include, but to not to be limited to the following situations:

- During preliminary examination of bids (verification of formal criteria)

- Absence of bid form(s), change in the wording or lack of signature on key portions of the bid form when this is clearly specified in the tender document as a requirement. Any change in wording that is consistent with the standard format of the bid form(s) is not a material deviation;

- The Bidder indicates in the bid that they do not accept important contract conditions, i.e. related to Warranty, Force Majeure Applicable Law, Delivery Schedule, Payment Terms, General Conditions and Limitation of Liability;

- Non historical documents required in the solicitation document have not been provided, such as documents specifically related to the bidding process and that the Bidder could not be expected to possess before the solicitation document was issued;

- Non eligibility of the Bidder;

- Financial information is included in the technical bid when using the two-envelope method.

- During technical evaluation of bids and qualification of Bidders:

- Specifications of the item quoted vary in one or more significant respect(s) from the minimum required technical specifications.

- The Bidder does not meet the minimum conditions for qualification.

- During financial evaluation of bids:

- The Bidder does not accept the required price correction as Instructions to Bidders Clause 25.1, c.

- Required price components are missing.

- The Bidder offers less quantity than what is required.

- If a bid is not substantially responsive to the bidding documents, it shall be rejected by UNFPA and may not subsequently be made responsive by the Bidder by correction of the material deviation, reservation, or omission.

## Nonconformities, Errors, and Omissions

  - Provided that a bid is substantially responsive:

- UNFPA may waive any non-conformities or omissions in the bid that do not constitute a material deviation.

- UNFPA may request that the Bidder submit the necessary information or documentation within a reasonable period of time to rectify nonmaterial nonconformities or omissions in the bid related to documentation requirements.  Such omission shall not be related to any aspect of the price of the bid.  Failure of the Bidder to comply with the request may result in the rejection of its bid.

- UNFPA shall correct arithmetical errors on the following basis:

- If there is a discrepancy between the unit price and the line item total that is obtained by multiplying the unit price by the quantity, the unit price shall prevail and the line item total shall be corrected, unless in the opinion of UNFPA there is an obvious misplacement of the decimal point in the unit price. In that case the line item total as quoted shall govern and the unit price shall be corrected;

- if there is a discrepancy between words and figures, the amount in words shall prevail;

- if there is an error in a total corresponding to the addition or subtraction of subtotals, the subtotals shall prevail, and the total shall be corrected.

  - If the Bidder that submitted the lowest evaluated bid does not accept the correction of errors, its bid shall be rejected.

## Preliminary examination of Bids

  - UNFPA shall examine the bids to determine whether they are complete, that all documents and technical documentation requested as per Instructions to Bidders Clause 10 have been provided and to determine the completeness of each document submitted. UNFPA will also examine whether any computational errors have been made, whether the documents are properly signed, and whether the bids are generally in order.

## Examination of Terms and Conditions and Technical Evaluation

- UNFPA shall examine the bid to confirm that it does not contain any material deviations, reservation, or omission related to the conditions and requirements specified in the Section II Technical Specifications and Schedule of Requirements, Section III UNFPA General Conditions of Contract and Section IV UNFPA Special Conditions for Contracts.

- If after the examination of the terms and conditions and the technical evaluation UNFPA determines that the bid is not substantially responsive in accordance with Instructions to Bidders Clause 24, the bid shall be rejected.

## Conversion to Single Currency

- To facilitate evaluation and comparison, UNFPA will convert all bid prices expressed in the amounts in various currencies in which the bid prices are payable to US dollars at the official UN exchange rate on the last day for submission of bids.

## Domestic Preference

- Domestic preference shall not be a factor in bid evaluation.

## Evaluation of Bids

- UNFPA shall evaluate each bid that has been determined, up to this stage of the evaluation, to be substantially responsive.

- UNFPA’s evaluation of a bid will exclude and not take into account:

- Customs duties and other import taxes, sales and other similar taxes, which will be payable on the goods if the contract is awarded to the Bidder;

- Any allowance for price adjustment during the period of execution of the contract, if provided in the bid.

## Comparison of Price Bids

- UNFPA shall compare all substantially responsive bids to determine the lowest priced substantially responsive bid.

- Bid comparison will be made on the total cost, delivered to final destination. UNFPA reserves the right to compare freight prices of Bidders with rates of reputable freight forwarders and to consider such rates for the purpose of bid evaluation. In the event that Bidder’s freight prices are found to be less competitive than the rates offered by freight forwarders, UNFPA may issue a contract on FCA basis to the Vendor instead of DAP and issue a separate contract for freight to a freight forwarder if deemed in the best financial interest of UNFPA.

## Post-qualification of the Bidder

- UNFPA shall determine to its satisfaction whether the Bidder with the lowest priced, substantially responsive bid is qualified to perform the contract satisfactorily.

- The determination shall be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted in the bid. An affirmative determination shall be a pre-requisite in order to award the contract to the Bidder. A negative determination shall result in disqualification of the bid, in which event UNFPA shall proceed to the bid that was evaluated as the next lowest priced, substantially responsive bid in order to make a similar determination of that Bidder’s capabilities to perform satisfactorily.

- To determine the Bidder’s capacity to execute the contract , UNFPA shall consider the following elements:

- Performance Statement Form, Section V, 4, with documentary evidence

- Copy of last year audited company Balance and Financial Statements

- Financial Capability:

  - Annual sales turnover during any one of the last three years to be at least equal to the contract value (from Financial Statements)

  - Liquidity ratio: Current ratio (Current Assets/ Current liabilities) > 1.

  - Documentary evidence that the Bidder has successfully completed at least one similar contract within the last five years for supply of goods.

  - Provide contact details of commercial banks and names of contact persons from whom UNFPA could seek feedback.

- Experience and Technical Capacity:

  - Registration details of the company

  - Experience to undertake the contract

    - List of similar contracts executed for other clients, including contract details.

    - Evidence that the Bidder possesses experience in the geographical area required by the bid.

    - At least three years of experience in performing similar contracts.

  - Company’s managerial capability:

    - Details of company’s managerial structure.

    - Quality assurance systems in place.

  - Bidder must have manufactured and supplied satisfactorily similar goods to a similar extent of the quantity, as mentioned against each schedule during any one of the last three years and the goods should have been in use satisfactorily with no adverse report.

  - Client's certificates in support of the satisfactory operation of the goods as specified above.

  - Data to support that the Bidder has the production capacity to perform the contract and complete the supplies within the stipulated delivery period or data to support that it has an installed annual production capacity for the specific item to match the quantities required. To qualify for multiple schedules, the installation capacity requirement shall be the sum of requirements against the individual schedules.

  - Evidence that the Bidder is in the continuous business of manufacturing/supplying and providing after sale services for goods similar to those offered during the last three years prior to bid opening date.

  - Brief write-up, backed up with adequate data, explaining available capacity and experience in the manufacture and supply of the required products within the specified time of completion after meeting all their current commitments.

  - Confirmation that all the facilities exist at the factory for inspection and testing and these will be made available to the purchaser or his representative for inspection.

  - The Bidder shall disclose instances of previous past performance that may have resulted in adverse actions taken against the Bidder and the manufacturers whose products are being offered by the Bidder, in the last five years. Such adverse actions may be treated as unsatisfactory performance history while deciding the award of contract. If no instance of previous past performance has resulted into adverse actions, this must be clearly indicated in the Bidder’s bid.

  - A list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the equipment for a reasonable period of time following installation.

For non manufacturer Bidders:

  - Legally enforceable authorization from the manufacturer assuring full guarantee and warranty obligations as per the tender conditions for the goods offered; and

  - The Bidder, as authorized by the manufacturers, has supplied and provided after sales service for similar goods to the extent of at least 20 percent of the quantities indicated in the tender requirements in any one of the last three years, and the goods must be in satisfactory operation.

  - Financial Experience and Technical Capacity requirements of the manufacturer similar to those mentioned above.

  - Notwithstanding anything stated above, UNFPA reserves the right to assess the Bidder’s capabilities and capacity to execute the contract satisfactorily before deciding on award.

  - Even though the Bidders may meet the above qualifying criteria, they can be subject to disqualification if they have made misleading or false representations in the forms, statements and attachments submitted in proof of the qualification requirements, and/or record of poor performance such as, not properly completing contracts, inordinate delays in completion, litigation history, financial failures, etc.

## UNFPA’s Right to Accept Any Bid and to Reject Any or All Bids

- A bid that is rejected by UNFPA may not be made responsive by the Bidder by correction of the non-conformity. A responsive bid is defined as one which conforms to all the terms and conditions of the UNFPA’s bid solicitation documents without material deviations. UNFPA shall determine the responsiveness of each bid against the UNFPA solicitation documents.

- UNFPA reserves the right to reject any bid if a Bidder has previously failed to perform properly or complete on time in accordance with contracts or the Bidder who in UNFPA’s perspective is not in a position to perform the contract.

- The Bidders waive all rights to appeal against the decision made by UNFPA.

## UNFPA’s Right to Annul a Bidding Process

- UNFPA reserves the right to annul the bidding process and reject all bids at any time prior to award of purchase order, without thereby incurring any liability to the affected Bidder(s) or any obligation to provide information on the grounds for UNFPA’s action.

# Award of Contract

## Award Criteria

- In the event of a contract award, UNFPA shall award the Contract/Purchase Order to the lowest priced Bidder(s) whose bid has been determined to be substantially responsive to the bidding documents.

- If required, the Bidder shall permit UNFPA representatives’ access to their facilities at any reasonable time to inspect the premises that shall be used for the production, testing and packaging of the products. The Bidder shall also provide reasonable assistance to the representatives for such inspection, including copies of any test results or quality control reports as may be necessary. UNFPA may inspect the manufacturing facilities of the lowest evaluated responsive Bidder to assess his capability to successfully perform the contract as per the terms and conditions specified in the ITB.

- UNFPA reserves the right to make multiple arrangements for any item(s) where, in the opinion of UNFPA, the lowest priced substantially responsive Bidder cannot fully meet the delivery requirements or if it is deemed to be in UNFPA’s best interest to do so. Any arrangement under this condition shall be made on the basis of the lowest priced substantially responsive, the second lowest priced substantially responsive, the third lowest priced substantially responsive , etc.  Further, UNFPA reserves the right to prioritize higher priced substantially responsive options over lower priced substantially responsive options based on critical factors such as productions and delivery lead times proposed by the bidders in order to satisfy the best interest of UNFPA and partners.

## Right to Vary Requirements at Time of Award

- UNFPA reserves the right at the time of award of contract to increase or decrease by up to 20% the quantity of goods specified in this bid without any change in unit price or other terms and conditions.

## Signing of the contract

- Prior to the expiration of the period of bid validity, UNFPA shall send the successful Bidder the Contract/Purchase Order, which constitute the notification of award.  The successful Bidder shall sign, date the contract and return it to UNFPA within 10 days of receipt of the contract. After receipt of the contract, the successful Bidder shall deliver the commodities in accordance with the quantity, quality and delivery schedule outlined in its bid in conjunction with UNFPA terms and conditions.

## Publication of Contract Award

- UNFPA shall publish the contract award on United Nations Global Marketplace http://www.ungm.org, with the information of the awarded Bidder company name, contract amount or LTA and the date of the contract.

- Suppliers perceiving that they have been unjustly treated in connection with the solicitation or award of a contract may lodge a complaint directly to the Chief, Supply Chain Management Unit at procurement@unfpa.org, who will then make an assessment of the complaint and provide a reply to the Supplier within a week and, if required, advise the Supplier on further recourse.

# SECTION II: Technical Specifications and Schedule of Requirements

# 2.1.	Technical Specifications

| Item No | Item name | Requirement | Quantity |
| --- | --- | --- | --- |
| Item 1 (OR 1) | Activated Clotting Time (ACT) Measuring device | Product Description: Activated clotting time (ACT) device is intended for in vitro diagnostic testing in point-of-care and it is designed for a quick testing and monitoring the whole blood coagulation time of patients undergoing heparin therapy. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - Rechargeable batteries with autonomy of at least 45 test cycles at 150 sec per test. Technical specifications: - Portable system for point-of-care use. - Multiple test types available, at least: - Activated Clotting Time (ACT+ and ACT-LR), - Activated Partial Thromboplastin Time (APTT and APTT Citrate), - Prothrombin Time (PT and PT Citrate). - Clotting times measurement system: through disposable single-use cuvettes. - One cuvette per specific test. - Screen display of at least: - Time and date. - Type of test running. - Test result. - Battery status. - Operating and configuring menus. - Barcode Reader for cuvette lot information, and quality control lot information. - Keypad for manual entry information. - Internal memory for at least 500 patient results and 500 quality control results. - Data port to be connected to a PC or a printer. - Test Precision: ≤ 10% C.V. for whole blood samples. - Test Chamber: 1 - Timing Range: up to 1020 seconds - Incubation Temperature: 37 °C ± 1.0 °C - Incubation Warm-Up Time: 30 seconds to 90 seconds - Self-test. - All materials resistant to hospital-use disinfectants. Accessories: - One (1) Rechargeable battery - One (1) Battery charger for 1 battery - One (1) PC and printer cable. - Set of start-up consumables NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares parts and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/746 or Directive 98/78 EC and Agreement Letter signed with the NB demonstrating the on-going IVDR application, for IVD devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 61010-1: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements - IEC 61010-2-101: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment - IEC 61326-1: Electrical equipment for measurement, control and laboratory use - EMC requirements - Part 1: General requirements | 2 |
| Item 2 (OR 2) | Portable Blood Gas Analyzer | Product Description Device is intended for in vitro diagnostic in point-of-care for a quick testing of various analytes in whole blood or plasma. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - Rechargeable batteries with autonomy of at least 30 minutes of continuous use. Technical specifications: - Handheld portable system for point-of-care use. - Test available, at least: blood gases, electrolytes, chemistries, coagulation, cardiac markers (cTnl), Hematology. - One cartridge for all the available tests or combinations of tests. - Sample Type: arterial or venous whole blood. - Sample Volume: < 100 µl. - Time to Result: approx. 5 minutes or less. - Self-test. Autocalibration. - Integrated bar code scanning. - Internal memory for at least 50 patient results and 1000 quality control results. - LCD display - Low battery indication. - Keypad for manual entry information and operation. - All materials resistant to hospital-use disinfectants. Accessories: - One (1) Pack of rechargeable batteries - One (1) Battery charger - One (1) Simulator - One (1) Portable printer compatible with the system offered. Including data and power cables. - Set of start-up consumables as coagulation cartridges. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manuals must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/746 or Directive 98/78 EC and Agreement Letter signed with the NB demonstrating the on-going IVDR application, for IVD devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 61010-1: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements - IEC 61010-2-101: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-101: Particular requirements for in vitro diagnostic (IVD) medical equipment - IEC 61326-1: Electrical equipment for measurement, control and laboratory use - EMC requirements - Part 1: General requirements | 2 |
| Item 3 (OR 3) | Extra Corporeal Circulation (ECC) Machine | Product Description A device designed to support the Heart-Lung perfusion practice. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - UPS with autonomy of at least 20 minutes of continuous use at maximum power. Technical specifications: - Modular Design - Each module must work independently from one another: no central control, no central failure or breakdown of the whole machine. - Low maintenance pumps without belts and gearbox. - Panel displays must be exchanged during HLM function without compromising the procedure. - Timer: Counting range 0 - 999 min 59 sec - System Panel: 3 Display and Control Modules - Pumps - 3 Roller Pump: - Diameter of pump raceway Ø: 150 mm - Diameter of occlusion roller Ø: 30.5 mm - rotate in, at least 15° increments. - Double Roller Pump: - Two roller pumps with a in a single housing. Both roller pumps must be operated independently. - Each pump must be operated in continuous or pulse mode. - Diameter of pump raceway Ø: 85 mm - Diameter of occlusion roller Ø: 15 mm - Speed Range: 0 to 250 rpm (clockwise, counterclockwise) - Mast System: - Modular, it must allow at least 3 roller pumps to be mounted. - It must allow an oxygenator and tubing set to be mounted. - Swivel telescope, with infusion rack and castor. - At least 2 swivel arms - At least 2 horizontal masts - Transport locking arm. - Cardioplegia Module: To deliver cardioplegic solutions or blood cardioplegia during an operation. - Manual and automatic operation modes. - Sensors - Pressure - The system must measure and displays the pressure in the cardiopulmonary bypass circuit: - Measurement range mmHg -200 mmHg to +800 - Resolution 1 mmHg - Cardioplegia - Temperature - The system must measure and display up to four temperatures: Display range 0 °C to +50 °C. - Level - Regulate and stop the arterial pump flow whether the blood level in the reservoir reaches the target volume. Level sensor pad small, flexible, and transparent to fit and stick on every reservoir and allow a clear view of the blood level. - Bubble - Sensor to detect different bubble air sizes in the extracorporeal circuit. The system must have at least three alarm thresholds to be set up (4 mm, 5 mm, and 6.5 mm Ø). Accessories: - One (1) set of transducers, cable, and holders for the pressure module. - One (1) temperature probe the temperature module. - One (1) Complete level controller module with one set of cells. - One (1) Plexiglass pump protection - Six (6) Cover for quick fixing - Two (2) Cable retention system - Two (2) Cable retention system - One (1) DRP 85 double head roller pump with red, yellow, and black inserts. - One (1) CP5 Transonic flow sensor (3/8") - One (1) 3/8" air bubble detector sensor module with 620 mm articulated support - One (1) Level sensor support - One (1) Venous clamp with mechanical control for coarse and fine adjustments. - Four (4) Self-Adhesive Cells for Generation II Level Sensor - One (1) Thermal generator 3 tanks 220V with remote control on control panel, maintenance kit and connections - Two (2) Hansen female elbow fitting - One (1) Hansen male straight fitting - Two (2) Hansen female straight fitting - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 10993-1 Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process. | 1 |
| Item 4 (OR 4) | Hypo-hyperthermia Machine | Product Description A device used to circulate water through heat exchangers to warm or cool a patient during cardiopulmonary bypass procedures lasting 6 hours or less. The brand and model offered must be compatible with the Extracorporeal Machine (Item 03) to work together in cardiopulmonary bypass procedures. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - UPS with autonomy of at least 20 minutes of continuous use at maximum power. Technical specifications: - The system must be based on three water tanks: - Cold and Warm tanks: to control the warming and cooling cardioplegia. - Patient tank with two circuits. - Temperature Ranges: - Cold Cardioplegia Tank: 2 °C to 10 °C - Warm Cardioplegia Tank: 15 °C to 40 °C - Patient Tank: 2 °C to 40 °C - Temperature control with accuracy of 0.1 °C for settings. - Circuits for the patient and the circuit for cardioplegia can be switched off separately. - Flow indicator. - Vacuum gauge with scale - Control panel must allow separated operation and configuration of the three circuits: Two patient circuits and one cardioplegia circuit. - Aerosol Collection Set. - Alarms: - High-temperature fluid limit: Visual and audible - Low fluid: Visual and audible - Water must not come in contact with the patient's or the patient's blood during a procedure. - Compressor-based cooling with refrigerant covered by the Kyoto protocol. - Separate settings for warm cardioplegia, cold cardioplegia and patient temperature. - Independent safety system to prevent the water temperature from reach critical temperatures. - Two-chamber pumps - Castors with brakes for easy lock. - Operating Temperature: 10 °C through 30 °C - Service life expected: 10 years. Accessories: - One (1) Aerosol collection container holder - One (1) set of tubing connectors for tanks, drain, etc. - One (1) Vacuum extension line with connector - Ten (10) Water filters 0.2 um - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manuals must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, Class IIb devices, - or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | 1 |
| Item 5 (OR 33) | Patient Auto-Transfusion System | Product Description A device used for intraoperative recovery of blood, washing of blood collected in the postoperative period, and preoperative sequestration with indirect patient connection. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Automatic machine 4 bowl Sizes. - Color touch Screen user interface with large view: 8” or larger. - Hematocrit indicator of returned blood. - Plasma Free Hemoglobin Indicator - Centrifuge Speed: 1500 - 5600 rpm (steps of 100 rpm) setting in steps of 100 rpm. - Pump Speed: 25 - 1000 ml/min (adjustable) - Wash protocol available for high quality RBC, Heparin, and protein removal. - Specific protocol for fat removal - High RBC recovery rate - Display Type Graphic color LCD TFT: at least 8” diagonally. Accessories: - One (1) Vacuum Generator - Vacuum Range: -50 to -300 mmHg. - steps of 10 mmHg - Set of Start-up consumables NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manuals must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 10993-1 Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process. | 1 |
| Item 6 (OR 5) | Infusion Rack (Syringe Pump and Infusion Pump including racks and trolley with holder pole with hooks) | Product Description Devices designed to accurately deliver liquids through intravenous (IV) or epidural routes for therapeutic and/or diagnostic purposes. The infusion pumps are used when the solution to be administered must be delivered with greater accuracy than can be provided through a manually adjusted gravity administration set. The syringe pump is designed to precisely drive the plunger of a syringe down its barrel to infuse a solution when it must be administered with a high degree of volume accuracy and rate consistency. Because of the lower flow settings and flow resolution, it is especially appropriate for neonatal, infant, and critical care applications in which small volumes of concentrated drugs are to be delivered over an extended period. It can also be used to administer epidural analgesia. Adult, pediatric and neonatal use. The required Infusion Rack must include: - Eight (8) Infusion syringe pumps - One (1) Volumetric infusion pump - Three (3) connectable docking stations with four (4) spaces each one. - One (1) cover for the docking stations - One (1) Mobile Trolleys for the docking station tower. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E - Rechargeable battery with autonomy of at least 3 hours of continuous use. Technical specifications: General characteristic of the infusion pumps: - Robust design allows use in demanding environments, resistant against hospital-grade cleaning solutions, and fluid proof. - Drug library: at least 700 drugs available. - Automatic switch from mains to battery during power failure. - An integrated display indicating at least the following parameters / information: - Alarms. - Pumping status. - Volume infused. - Volume limit/Volume To Be Infused (VTBI). - Automatic calculation of third parameter when user enters other two parameters (volume, time, and flow rate). - Should be protected against uncontrolled gravity flow, a so-called free-flow protection. - Should be equipped with a "control lock-out" feature, preventing tampering by patients or visitors with the controls of the unit. - Visual and acoustic alarms for at least: - VTBI infused. - Time expired / Infusion complete notification. - Occlusion / High pressure - KVO finished. - Drive blocked. - Low battery - Incorrectly loaded set. - Calibration needed. Infusion syringe pump with the follow technical specifications: - Preselected Volume: - 0.1 – 99.99 ml in increments of 0.01 ml - 100 – 999 ml in increments of 0.1 ml - 1000 – 9999 ml in increments of 1 ml - Delivery Accuracy: ± 2 % - Syringes sizes: 50/60ml, 30/35ml, 20ml, 10/12ml, 5/6ml, 2/3ml - Adjustable delivery rates: - 0.1 – 99.99 ml/h in increments of 0.01 ml/h - 100 – 999 ml/h in increments of 0.1 ml/h - Bolus Rates: 1 - 1800 ml/h - Delivery accuracy with bolus: ± 2 % - Keep Vein Open (KVO) rate: - Delivery rate ≥ 10 mL/h: KVO-rate 3ml/h. - Delivery rate < 10 mL/h: KVO-rate 1ml/h. - Delivery rate < 1 mL/h: KVO-rate = set rate (default setting 0.1ml/h). - Occlusion alarm: 9 levels from 75 mmHg to 900mmHg - Incorrect delivery alarm: In case of incorrect dosage of 0.1 ml due a malfunction the pump turns off automatically. Volumetric infusion pump with the follow technical specifications: - Preselected Volume: - 0.1 – 99.99 ml in increments of 0.01 ml - 100 – 999 ml in increments of 0.1 ml - 1000 – 9999 ml in increments of 1 ml - Accuracy of set delivery rate: ± 5 % - Flow Rates: - 0.1 – 99.99 ml/h in increments of 0.01 ml/h - 100 – 999 ml/h in increments of 0.1 ml/h - 1000 – 1200 ml/h in increments of 1 ml/h - Accuracy bolus infusion: ± 5 % as of a bolus volume > 1 ml - Keep Vein Open (KVO) rate: - Delivery rate ≥ 10 mL/h: KVO-rate 3ml/h. - Delivery rate < 10 mL/h: KVO-rate 1ml/h. - Delivery rate < 1 mL/h: KVO-rate = set rate (default setting 0.1ml/h). - Detection of air bubbles ≥ 0.01ml - Air detector alarm. - Occlusion alarm: 9 levels from 225 mmHg to 900mmHg Docking stations with the follow technical specifications: - To accommodate the infusion and syringe pumps - With four (4) spaces for pumps. - Rails for the individual pumps. - With connectors for data lines and voltage supply for each pump. - With connectors for further docking stations, at least up to 24 pumps. - Mains power input. Cover for the docking stations with the follow technical specifications: - To cover the uppermost segment of the docking station tower. - With a carrying handle that supports all the docking stations. Mobile Trolleys with the follow technical specifications: - Made of steel with anti-corrosive epoxy coating, aluminum, AISI 304 stainless steel or higher quality material. - With brakes at least in two (2) of the wheels. - Solution holder pole with 2 hooks, minimum. Accessories: - Universal clamp rotatable for attachment to wall rails and infusion poles for each docking station. - Power supply cable for each pump and docking station. - Extension cables, to combine the docking stations (quantity as necessary). - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-24: Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers. | 4 |
| Item 7 (OR 10) | Monitor/defibrillator + specific spoons for children + specific options + accessories | Product Description Device design for defibrillation and monitoring on victims of cardiac arrest where there is apparent lack of circulation as indicated by unconsciousness, absence of breathing or absence of pulse. A versatile resuscitation system integrates essential components such as a defibrillator, ECG display, advanced monitoring functionalities, and Noninvasive Transcutaneous Pacing (NTP). Designed for various resuscitation scenarios with compact and lightweight design for facilitating easy patient transport within healthcare facilities. The device must include the followings operation modes: - Defibrillator Function in Manual Mode - The unit must be designed to be used for synchronized cardioversion of certain atrial under supervision of a qualified physician. - Real CPR Help - The unit must be designed to provide visual and audio feedback to encourage rescuers to perform chest compressions at a recommended rate. - ECG Monitoring - The unit must be designed to be used when ECG monitoring is indicated to evaluate the patient’s heart rate or ECG morphology, used by qualified personnel training in basic life and/or advanced life support, or other physician-authorized emergency medical training. - External Pacemaker - Noninvasive transcutaneous pacing (NTP) function is preferable for temporary external cardiac pacing in conscious or unconscious patients as an alternative to endocardial stimulation. - SpO2 Monitoring - Oximeter function indicated for the continuous, noninvasive monitoring of arterial oxygen saturation (SpO2) and pulse rate during both no motion and patient motion conditions for adult patients, and no motion conditions for Pediatric and neonatal patients in a hospital environment. - EtCO2 Monitoring - Indicated for the continuous noninvasive monitoring of end tidal carbon dioxide (EtCO2) and respiration rate in patients requiring ventilator support, in-hospital transport, or anesthesia. - NIBP - NIBP option indicated for the non-invasive measurement of arterial blood pressure for resting patients in critical care and in-hospital transport. Electrical Requirements: - Powered by both AC mains and a replaceable battery pack, the system ensures continuous operation and swift recharging capability. Additionally, compatible battery chargers enable efficient recharging and testing of the battery pack. - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Screen: - High-resolution, liquid crystal display (LCD) or TFT. - At least 6.5 inches (16.5 cm) diagonally size. - Battery: Rechargeable lithium-ion battery pack - Recharge time: 4 hours or less. - Battery Shelf Life - at least 3 months before retest and recharge. - Defibrillator: - Waveform - Rectilinear Biphasic - Energy selection (delivered to a 50Ω load) for adult and pediatric - At least 20 steps from 1 to 200 joules. - Charge Time less than 7 seconds. - Energy display - Screen shows selected and delivered energy. - Synchronizes defibrillator discharge to the patient’s R wave. - Paddles - Standard apex/sternum paddles. Adult and pediatric plates. - ECG Monitoring - Patient connection - 3-lead cable, 5-lead cable, paddles, or hands-free therapy electrodes. - Lead Selection - I, II, III, aVR, aVL, aVF, V, PADS and PADDLES - Input protection - Fully defibrillator-proof. - Implanted pacemakers display spike and pulse detection. - Bandwidth: 0.5 Hz to 21 Hz standard; 0.05 Hz to 150 Hz diagnostic - Heart rate range - 30 to 300 beats per minute - Heart rate alarm - Tachycardia and Bradycardia Accessories: - One (1) Sensor for each monitoring function - One (1) Power Cord Extension Cable (12”) - One (1) Power Charger - One (1) Power Battery - Two (2) sets of resuscitation electrodes - Two (2) sets of defibrillation electrodes. - Set of Start-up consumables NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-4:2010/AMD1:2018: Amendment 1 - Medical electrical equipment - Part 2-4: Particular requirements for the safety of cardiac defibrillators. | 2 |
| Item 8 (OR 11) | Suction system, surgery | Product Description A device designed for suctioning the oral cavity, pharynx, and bronchial system. Also clearing the airways of viscous secretion and food residues. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Suction capacity: At least 25 l/min - Max. Vacuum: At least -80 KPa - Short-term operation (5 min on. 20 min off) - Noise Level: 50 dB (A) or less. Accessories: - Set of Start-up consumables Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment – Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances – Requirements and tests - IEC 60601-1-11 – Medical electrical equipment – Part 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment. | 4 |
| Item 9 (OR 12) | Cardiovascular ultrasound imaging system | Product Description Device designed to deliver real-time, non-invasive imaging of blood vessels and heart structures and functionality, allowing the analysis and diagnosis of various cardiovascular pathologies, including heart diseases, valvular and vascular pathologies. To be used in adult and pediatric patients. Including cardiovascular software package and printer. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E - Rechargeable batteries with autonomy of at least 30 minutes of continuous use. Technical specifications: - High resolution in real-time mobile ultrasound system. - System mounted on stable, mobile trolley fitted with 4 caster antistatic wheels with brakes. - Displays images on screen and DICOM. - Main Monitor Full HD (at least 21 inch) on articulated arm for easy movements. - Control Panel: - Height adjustment. - Integrated (at least 12 inch) Full HD LCD touch screen. - Keyboard. - Trackball. - Support for at least 5 probes. - Support for probes cable. - Connection ports for at least 4 probes - ECG port connection - Color thermal printer. - USB port at least 2. - It must allow networking through WiFi. - Battery backup with sleep mode. - Measurement accuracy to be better than 2% over 10 cm distance. - Doppler display to indicate blood flow both numerically and in color. - Ultrasound penetration depth up to 36 cm. - The hardware and software included in the offer will allow at least the following application: - Cardiac and stress echo, including – TEE scanning capability. - Tissue differentiation to clearly show the walls of the left ventricle and regional wall motion abnormalities. - Peripheral and deep vascular. - Micro flow imaging. - Abdominal imaging. - Tissue imaging synchronization or equivalent technique. - 3D image for cardiac studies in adults, children, and infants. - High-definition zooms of the live image or the frozen image, with the possibility of cropping the zoom. - Quantification tools: measurement of longitudinal “strain” and automatic measurement of EF. - All materials resistant to hospital-use disinfectants. Accessories: - Probes to be included in the offer: all the necessary to cover the aforementioned applications in adult and pediatric patients (Convex, Linear, Phased, TEE, Micro TEE types of probes). - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-37 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. | 1 |
| Item 10 (OR 16) | Electrosurgical system | Product Description Device that uses high frequency electrical energy in a radiofrequency (RF) band to develop heat directly within soft-tissue cells (thermodynamic) for cutting and coagulating tissue typically during general surgical procedures. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Application Monopolar and bipolar - Monopolar and bipolar modes: cut, coagulation, including forced and spray coagulation mode. - Power activation controlled by foot switch, and by handswitch at the handpiece. - Self-test - Power output: - Bipolar: Max. 400 Watt - Monopolar: Max. 400 Watt - RF generator output: 350 kHz. - Instruments socket connection: at least two monopolar and one bipolar. - User interface: Power control and modes selection by touch high-resolution display. - Neutral electrode isolated from ground. - Neutral electrode contact monitoring system. - Automatic power limitation. - Power output blocking in case of active electrode or neutral electrode contact failure. - Visual and audible activation indicators. - Visual and audible alarms. Accessories: - Double-pedal footswitch for monopolar and bipolar modes, with connecting cable. - Two (2) neutral electrodes, reusable, sterilizable, monopolar with connecting cable. - One (1) monopolar electrode handle (pencils), finger switch controlled, reusable, sterilizable, with connecting cable. - One (1) monopolar electrode handle (pencils), foot switch controlled, reusable, sterilizable, with connecting cable. - One (1) set of different monopolar reusable electrodes: blade/knife, needle, ball, loop, spatula, and coagulation electrodes. - Two (2) bipolar forceps, foot switch controlled, reusable, sterilizable, with connecting cable. - Two (2) reusable connection cables for disposable split neutral electrodes. - Reusable adapters and connecting cables for monopolar and bipolar handpieces, as necessary. - Trolley with basket, made of steel with anti-corrosive epoxy coating, aluminum, AISI 304 stainless steel or higher quality material. With 4 antistatic castors, 2 with brakes. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia.. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-2: Medical electrical equipment - Part 2-2: Particular requirements for the basic safety and essential performance of high frequency surgical equipment and high frequency surgical accessories. - ISO 10993-1 Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | 1 |
| Item 11 (OR 18) | Hand-held ultrasound imaging system | Product Description Transportable ultrasound system designed to provide in real-time, visualization of the internal organs, anatomy, and fluid flow analysis, of the human body. To assist during puncture procedures. Pediatric and neonatal use. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E - Rechargeable batteries with autonomy of at least 4 hours of continuous use. Technical specifications: - Handheld ultrasound system. - To assist during puncture procedures. - Mobile device (tablet type) must be included, with at least the following characteristics: - Minimum 8 GB total storage for the specific software ultrasound package and patients records. - Color display, minimum 12 cm (4.7 in) - Touch interface. - Minimum resolution: 1280 x 800 - Integrated speakers - Full compliance with IEC 60950-1 or IEC 60601-1 - Full compliance with USB On-The-Go standard. - Operating system: Android 6.0 or later operating system - iOS devices: iOS 11 or later operating system - Wireless or cellular networking capability - Access to ports 80 and 443 - Audio capability - Front- and rear-facing cameras. - Connection port for probe: USB compatible with the transducer cable. - The hardware and software included in the offer will allow at least the following: - Imaging modes: 2D, PW, Color Doppler, M-mode, advanced XRES, multivariate harmonic imaging, SonoCT and B-line detection in lung imaging. - Applications: Abdominal, Carotid, Musculoskeletal (conventional and superficial), Pediatric, Peripheral Vessel, Small Organ, Lung. - Probe: - Broadband linear array probe and cable with USB connector. - Operating frequency range: 12-4 MHz - Aperture size: 34mm - Center line marker. - Ultrasound penetration depth up to 12 cm. - All materials resistant to hospital-use disinfectants. Accessories: - Power supply system for mobile device - Protective case for mobile device - Carry case. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-37 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. | 1 |
| Item 12 (OR 19) | Brain Monitoring System | Product Description Device designed for Brain monitoring. The system analyzes the electrical activity of the brain to provide real-time feedback on the depth of anesthesia during medical procedures. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Display Features: - BIS number - Trend graph of BIS values over time - Real-time raw EEG waveforms - Signal quality indicators like SQI and EMG - Alarm indicators and messages. - Sensor Technology: - Specialized sensor attached to the patient's forehead and temple. - Records brain's electrical activity through EEG - Data Analysis: - Proprietary algorithm analyzes EEG data to calculate BIS value. - Provides real-time information on anesthesia depth. Accessories: - One (1) Patient Interface Cable (PIC). - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical Electrical Equipment - General Requirements for Basic Safety and Essential Performance - IEC 80601-2-26 Medical electrical equipment: Part 2-26: Particular requirements for the basic safety and essential performance of electroencephalographs | 1 |
| Item 13 (OR 20) | Monitor for cerebral oximetry | Product Description The device is intended for use in hospitals, providing continuous non-invasive measurement of cerebral oxygen saturation and changes in regional oxygen saturation of blood in the brain or other body tissues beneath the sensor. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - Battery Type Lithium-ion - Operating time 1 hour under normal monitoring conditions. Technical specifications: - Sensor Application: Noninvasive monitoring of regional hemoglobin oxygen saturation in up to four site-specific areas. - Screen type TFT LCD: screen size - At least 25 cm (diagonally) - Technology: Real-time data on oxygen supply and demand balance, reflecting venous oxygen reserve. - Patient Monitoring: Intended for single patient use in monitoring regional oxygen saturation in neonatal, pediatric, and adult patients. - Two- or Four-Channel Monitoring: Capability to monitor either two or four channels simultaneously. - Signal Strength Indicator (SSI): Indicator showing the strength of the signal being received. - Alarms and Status Messages: Information on alarms and status messages displayed by the system. Accessories: - Two (2) Reusable sensor cable (RSC) - One (1) Adult regional saturation sensor - One (1) Pediatric regional saturation sensor - Two (2) Adult rSO2 sensor - Set of start-up consumables - Accessories to be reusable. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and greement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical Electrical Equipment - General Requirements for Basic Safety and Essential Performance - ISO 80601-2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. | 1 |
| Item 14 (OR 22) | Multiparametric TRANSPORT patient monitor for adult and pediatric patients | Product Description Device intended for use to measure basic physiologic parameters and track the status of adult, pediatric and neonatal patients. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - Built-in rechargeable battery, allowing at least 2 hours of continuous operation. Technical specifications: - For use in adult, pediatric and neonatal patients. - Automatic setting of alarm limits and cuff pressure limits for each type of patient. - Monitoring at least: ECG and Heart Rate (HR), Respiratory Rate (RR), Oxygen Saturation (SpO₂), non-invasive blood pressure, and Temperature. - LCD TFT display, 1280*800 resolution, possibly touchscreen, at least 12”. - Display of at least 4 waveforms and numeric parameters simultaneously. - Configurable display. - Defibrillator shock protection. - Trend storage of at least 120 hours. - Availability of a USB port for installing program updates and for recording parameters - Availability of a built-in thermal printer - ECG: - At least 5-leads: I, II, III, aVR, aVL, aVF, V. - ECG main cable (if applicable) and two (2) sets of patient cable terminals (1 neonatal/pediatric, 1 adult if applicable), included. - Simultaneously display a minimum of 2 ECG traces, and HR value. - Heart rate measurement range. Adult: at least 15 – 300 bpm, pediatric and neonatal: at least 30 – 300 bpm. Accuracy: ± 1% or ± 1 bpm, whichever is greater. - Signal amplification 5-10-20-40mm/mV. - S-T segment and arrhythmia analysis. - Pacemaker detection. - Lead off condition detected and displayed. - Adjustable sweep speed - Respiration: - Technique: transthoracic impedance - Measurement range: at least 6 – 120 rpm; 5-150 for neonatal. Resolution: 1 rpm - Display of waveform, and RR value. - Adjustable sweep speed - Noninvasive blood pressure: - Technique: oscillometric - Three (3) reusable blood pressure cuffs (1 neonatal, 1 pediatric, 1 adult) with hoses included. - Manual and automatic measurement, with configurable intervals. - Display diastolic, systolic, and average pressure. - Measurement range. - Adults. Systolic: at least 40 – 250 mmHg, Diastolic: at least 10 – 210 mmHg. Maximum mean error: ± 5 mmHg - pediatric. Systolic: at least 30 – 180 mmHg, Diastolic: at least 10 – 150 mmHg. Maximum mean error: ± 5 mmHg - Neonatal. Systolic: at least 30 – 130 mmHg, Diastolic: at least 10 – 100 mmHg. Maximum mean error: ± 5 mmHg - Pulse oximetry: - MASSIMO/NELLCOR technology - Measurement range: 0 – 100%. Resolution: 1%, SpO2. Accuracy at least ± 3% within the range 70 – 100% - Pulse rate: 30 – 250 bpm, resolution: 1 bpm, accuracy: 2% or 2 bpm, whichever is greater. - Display of percentage of oxygen saturation, plethysmography curve and heart rate. - Main cable of SpO2, if applicable, included. - Three (3) SpO2 sensors, reusable clip-on type, adult pediatric and neonatal sizes, included. - Temperature: - Cutaneous / abdominal. - Two temperature measurement channels: T1, T2, ∆T - Measurement range: at least 0 - 45°C. Accuracy: ± 0.1° C. Resolution: 0.1°C - Four (4) skin and esophageal temperature probes, reusable, 1 of each adult and 1 of each pediatric included. - Alarms: - Audio-visual alarms for all monitored parameters - Adjustable high and low alarm limits for all monitored parameters. - Temporary silence functions. - Leads-off or sensor disconnect. - Apnea alarm. - AC status and low battery - Automatic self-test. - All materials resistant to disinfection with hospital-grade products. - Indications and messages on the equipment must be in Russian language. Accessories: - Five (5) Reusable cuff for adult sizes M, L, XL; newborns - One (1) host for NIBP, in addition to that provided with the device. - ECG main cable (if applicable), in addition to that provided with the device. - Two (2) sets of ECG patient cable terminals (1 neonatal/pediatric, 1 adult if applicable), in addition to that provided with the device. - One (1) skin temperature sensor, reusable, adult size, in addition to that provided with the device. - One (1) Main cable of SpO2 (if applicable) in addition to that provided with the device. - At least three (3) SpO2 sensors adult size, reusable clip-on type, in addition to that provided with the device. - At least two (2) SpO2 sensors, pediatric size, reusable clip-on type, in addition to that provided with the device. - Two (2) SpO2 sensors, neonatal size (for children weighing from 1-5kg), reusable clip-on type, in addition to that provided with the device. - Twenty (20) SpO2 single use sensors. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Copy of the agreement between the bidder and the local service center for the provision of services - Manufacturer's authorization issued to the local service center for the right to provide maintenance services for the declared medical equipment (copy certified by the seal of the bidder and the local service center) - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests. - IEC 60601-2-49: Medical Electrical Equipment - Part 2-49: Particular Requirements for the Basic Safety and Essential Performance of Multifunction Patient Monitoring Equipment | 4 |
| Item 15 (OR 23) | Mounting system for monitoring arm or lateral support for ventilation machine | Product Description Adjustable arm for bringing the monitor away from the mounting point and closer to the working space offering the ability to adjust to the individual ergonomic requirements of the operator. This product does not fall under the category of medical devices. Technical specifications: - Monitor adaptation for patient monitor. - Quick Release tool-free disassembly - Arms with integrated cable management. - Load capacity - At least 22 kg / 48.5 lbs. Accessories: N/A Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). Regulatory approvals required: - CE-marked. Safety & product Standards: - ISO 9001 | 1 |
| Item 16 (OR 24) | Slave Monitor 28’’ + Wall fixation Arm | Product Description Monitor used in medical settings or surgery as a secondary monitor that duplicates the display of the primary monitor. It plays a crucial role in providing additional visual information during procedures, it does not fall under the category of medical devices. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Display size 28" or more. - True flat medical-grade full HD and UHD. - LED Technology based. - Brightness (cd/m²) - At least 500 cd - Viewing angle 178 ° - Touch screen. - Compatible Vesa mount - True-flat design with AR filter High brightness and high contrast ratio - 4K ultra high-definition resolution (3840 x 2160) - Ports: - 1 x HDMI 2.0 - 1 x DisplayPort 1.2 (3840 x 2160 resolution) - 1 x SDI (3840 x 2160 resolution) - 1 x DVI (1920 x 1080 resolution) - Power consumption - Max.: 158 W Accessories: N/A Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). Regulatory approvals required: - CE-marked Safety & product Standards: - ISO 9001 | 1 |
| Item 17 (OR 26) | Lamp, operating room, ceiling | Product Description A device designed to provide a specialized source of light during a cardiovascular surgical intervention. Double dome, operating lamp, LED technology. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Ceiling mounting system, with articulating arm. - Double dome. - LED technology. - Light life: at least 40,000 hrs. - Illumination level, at 1m distance: 160,000 lux. Without shadows. - Color temperature: between 4,000 and 5,600 K. - Color rendering index: at least 93. - Control panel located on the lamp, with on/off control and light intensity adjustment. - Intensity adjustment: at least 3 levels. - All materials resistant to hospital-use disinfectants. Movements: - Lamp head turning angle (inclination) at least 180° - Horizontal degree of freedom on all axles 360°. Accessories: - Two (2) removable and autoclave sterilizable handler - All necessary elements to anchoring and fixing the double dome operating lamp on the ceiling. - Set of start-up consumables. Accessories should be reusable when an option is available. Installation: - The supplier will carry out equipment installation, and safety and operational checks prior to delivery, leaving the equipment operating according to the manufacturer's specifications. The cost of installation should be included in the offer. NOTE: All the necessary accessories and spare for the correct installation and operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - Declaration of Conformity (DoC) according with regulation 2017/745 for Medical Devices Class I. - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-41: Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. | 1 |
| Item 18 (OR 27) | Video-laryngoscope | Product Description The video laryngoscope is a device designed for endotracheal intubation of patients who require assistance in the delivery of oxygen due to surgical or emergency events, especially for difficult airways. Electrical Requirements: - Operation: battery-powered with an autonomy of at least 100 minutes of continuous use. - Charger power requirements: according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E (if applicable) Technical specifications: - Portable tracheal intubation system, video assisted. - Provides a real-time view of the airways enabling quick intubation of adult, pediatric, and neonatal patients. - Monitor: - TFT LCD color display or higher. - Display size at least 2.4'' - Resolution of at least 320 x 240 - Integrated in a single piece with the handle or connectable. - Battery indicator status. - Acquisition module integrated or adaptable to the different handles. - Video camera with CMOS (Complementary Metal Oxide Semiconductor) or CCD (Charge Coupled Device) technology, - LED Light Source, - Set of handles and blades with the following characteristics: - Integrated handle-blade or with interchangeable blades. - In either case, at least four (4) different sizes of disposable and sterile blades, for tracheal intubation in adults, pediatrics and neonatal. - Additional hyperangulated curve blade, for complex tracheal intubation. - All materials resistant to hospital-use disinfectants. Accessories: - One (1) Carry case. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-1-12 Medical Electrical Equipment - Part 1-12: General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the emergency medical services environment. | 1 |
| Item 19 (OR 28) | Patient heater | Product Description Device designed to help the patient to meet the appropriate body temperature during surgical procedures. Pediatric and neonatal use. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Warming Unit type: Convective. - Normothermia System. - To be used with a pediatric warming blanket for surgical procedures. - The warming unit is attached to the warming blanket through a flexible hose. - Temperature adjustment: At least 4 temperature settings: - Ambient. - Low: 34°C ± 2°C - Medium: 40°C ± 2°C - High: 45°C ± 2°C - Control Panel: display and keypad. - Accuracy of displayed Temperature: < ± 2°C - Average time for contact surface (temp. to rise from 23 ± 2 a 37°C): ≤ 6 minutes. - Thermal protection threshold: Thermostat 49°C to 55°C - Unit inlet filter: At least MERV 14 or higher as HEPA. - Alarms: Visual and audible. At least for over temperature, under temperature and system failure. - Pediatric warming blanket dimensions: approx. 69 cm W x 124 cm L inflated. - All materials resistant to hospital-use disinfectants. Accessories: - Twenty (20) single-use warming blankets, pediatric use. - Cart / Trolley made of steel with anti-corrosive epoxy coating, aluminum, AISI 304 stainless steel or higher quality material. With 4 castors and at least 2 with brakes. - Other start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-35: Medical electrical equipment - Part 2-35: Particular requirements for the basic safety and essential performance of heating devices using blankets, pads and mattresses and intended for heating in medical use. | 2 |
| Item 20 (OR 29) | Blood / fluid warmer | Product Description A device designed to warm blood, blood products and intravenous fluids, and infuse them to the patient by transfusion, infusion, or irrigation. Adult and pediatric use. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Warmer Type: In-line - Infusion temperature: at least up to 40 °C (± 2 °C) - Gravity Flow Rate: at least up to 5.000ml /hr. - Control Panel: LCD Display, at least for temperature visualization. - Alarms: Visual and audible - Attachable to an IV Pole - All materials resistant to hospital-use disinfectants. Accessories: - Clamp for IV pole. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ASTM F2172: Standard specification for blood/intravenous fluid irrigation warmer - ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process | 2 |
| Item 21 (OR 31) | Patient Ventilator/Respirator | Product Description Device intended for use as an anesthesiology workstation, providing full flexibility for various spatial conditions and equipped with features like an electronic gas mixer with automatic fresh-gas capabilities. This device is intended to support lung protective ventilation and ensure precision in ventilatory support during medical interventions, specifically tailored for anesthesiology procedures. Designed for Adult & pediatric ventilation procedures. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. - Euro standard cable at least 3 meters long - Availability of at least 3 additional sockets for connecting auxiliary equipment (monitor, infusion pump). Technical specifications: - Patient type: Adults, pediatric and neonate patients - Ventilation modes - Volume controlled ventilation: ‒ VC-CMV ‒ VC-SIMV ‒ VC-AC ‒ VC-MMV - Pressure controlled ventilation: ‒ PC-CMV ‒ PC-BIPAP 1 / SIMV+ ‒ PC-SIMV ‒ PC-AC ‒ PC-APRV ‒ PC-PSV - Support of spontaneous breathing: ‒ SPN-CPAP/PS ‒ SPN-CPAP/VS ‒ SPN-CPAP ‒ SPN-PPS - Special procedures - Suction maneuver - Manual inspiration/hold - Medication nebulization - PEEPi - Therapy types - Invasive ventilation (Tube) - Non-invasive ventilation (NIV) - O2 -therapy - Respiratory rate (RR) - Adult 0.5 to 98/min - Neonates 0.5 to 150/min - Inspiratory time (Ti) - Adults 0.11 to 10 s - Neonates 0.1 to 10 s - Tidal volume (VT) - Adults 0.1 to 3.0 L - pediatric patients 0.02 to 0.3 L - Neonates 0.002 to 0.1 L - Inspiratory ﬂow (Flow) - Adults 2 to 120 L/min - pediatric patients 2 to 30 L/min - Maximum ﬂow during non-invasive ventilation of neonates (Flow max): 30 L/min - Inspiratory pressure (Pinsp) - Adults 2 to 120 L/min - pediatric patients 2 to 30 L/min - Pressure limitation (Pmax) - 2 to 100 mbar (or hPa or cmH2O) - Positive end-expiratory pressure (PEEP) - 0 to 50 mbar (or hPa or cmH2O) - Pressure support (Psupp) - 0 to 95 mbar (or hPa or cmH2O) - Pressure rise time (Slope) - Adults, pediatric patients 0 to 2 s Neonates 0 to 1.5 s - O 2 concentration (FiO2) - 21 to 100 Vol.% - Trigger threshold (Flow trigger) - 0.2 to 15 L/min - Automatic Tube Compensation - Inside tube diameter - Endotracheal tube ET Adults and pediatrics - Tracheostomy tube (adult and pediatric) - Degree of compensation 0 to 100 % - Flow Measurement - Minute volume measurement. - Expiratory minute volume, overall, not leakage corrected (MVe) - Inspiratory minute volume, overall, not leakage-corrected (MVi) Minute volume, leakage-corrected (MV) - Mandatory expiratory minute volume, overall, not leakage corrected. - Spontaneous expiratory minute volume, overall, not leakage corrected. - Tidal volume measurement - Tidal Volume, leakage-corrected - Mandatory inspiratory tidal volume, not leakage-corrected - Mandatory expiratory tidal volume, not leakage-corrected Spontaneous inspiratory tidal volume, not leakage-corrected - Respiratory rate measurement - Respiratory rate - Mandatory respiratory rate - Spontaneous respiratory rate - O2 measurement (inspiratory side) - Inspiratory O2 concentration (FiO2) - CO2 measurement in main ﬂow (adults and pediatric patients only) - End-tidal CO 2 concentration (etCO2) - Dynamic compliance (Cdyn): Range 0 to 1000 mbar/L/s (or hPa/L/s or cmH2O/L/s) - Leakage minute volume (MVleak): Range 0 to 99 L/min, BTPS - Rapid shallow breathing index (RSBI) - Adults 0 to 9999 (/min/L) - pediatric patients 0 to 9999 (/min/L) - Neonates 0 to 300 (/min/L) - Negative Inspiratory Force (NIF): Range -80 mbar to 0 mbar (or hPa or cmH2O) - Occlusion pressure P0.1: Range 0 to -25 mbar (or hPa or cmH2O) - Waveform displays: - Airway pressure Paw (t) - Flow (t) - Volume V (t) - CO 2 (t) - Alarms / Monitoring - Expiratory minute volume (MVe) - Airway pressure (Paw) - Inspiratory O 2 concentration (FiO2) - End-tidal CO 2 concentration (etCO2) - Respiratory rate (RR) - Volume monitoring - Apnea alarm time - Disconnection alarm time - Performance data - Control principle / Time-cycled, volume-constant, pressure-controlled. - Length of intermittent PEEP - Medication nebulization - Inspiratory ﬂow - Base ﬂow, adults, pediatrics, and neonates - Inspiratory valve - Gas supply - O2 - Air - Battery - Internal battery of ventilation unit - Automatic switch over from internal to external battery - Battery test available - Full HL7 data export to HIS. - Cybersecurity: design for preventing dangerous and damaging cyber-attacks based on NIST (National Institute of Standards and Technology) security best practices or equivalent. Accessories: - Two (2) vaporizers, one for Sevoflurane and one for Isoflurane. With a visual indicator of the filling level of the anesthetic agent, and adapter for filling the vaporizer if required. - One (1) Air pressure regulator, compatible with the medical gas system of the health unit. - One (1) O2 pressure regulator, compatible with the medical gas system of the health unit. - Hoses for Air and O2 with their respective connections compatible with the gas inlet of the equipment and the supplied pressure regulators. - One (1) Oxygen cell, if applicable. - Three (3) Flow sensors, if applicable. - One (1) Electronic Gas Mixer, if applicable. - One (1) Heated Breathing System, if applicable. - Five (5) Patients Circuits. - Set of start-up consumables NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Copy of the agreement between the bidder and the local service center for the provision of services - Manufacturer's authorization issued to the local service center for the right to provide maintenance services for the declared medical equipment (copy certified by the seal of the bidder and the local service center) - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests. - IEC 60601-2-13: Particular requirements for the safety of anaesthetic workstations. | 1 |
| Item 22 (OR 32) | Portable external pacemaker | Product Description External pacemaker is a device designed for temporary cardiac pacing, that is the application of an artificial electrical stimulus, in combination with a cardiac pacing lead system. Electrical Requirements: - Operation: Alkaline batteries powered external pacemaker. Technical specifications: - Dual-chamber device. - To be used with patients of any age. - At least for the following applications: - Pre-, intra- and postoperative pacing of patients with heart surgery - Treatment of arrhythmias and heart block - Symptomatic sinus bradycardia - Emergency pacing - Option with fast atrial stimulation up and atrial overdrive stimulation - Dial for quick parameter adjustment. - Visualization of the programmed settings. - Lead impedance monitoring. - Self-test. - Adjustable parameters range, at least: - Pacing modes (NBG code): DDD, D00, VDD, VVI, V00, VVT. - Rate: between 30 to 200 ppm - Sensitivity (A: Atrial): between 0.4 to 10 mV - Sensitivity (V: Ventricular): between 1 to 20 mV - AV delay: between 20 to 300 ms. - Atrial rate (A): between 80 to 800 ppm. - Pre-setting no-adjustable parameters, at least: - Pulse width: 1 ms. - Upper rate: 230 ppm - Atrial refractory period: between up to 500 ms. - In channel blanking: 110 ms. - Alarms acoustic and/or visual at least for: - Low battery indication. - Rate set > 180 ppm. - Lead impedance out of range. - Auto-test failure. - All materials resistant to hospital-use disinfectants. Accessories: - Two (2) Startup sets of alkaline batteries. - Two (2) Adapter to connect the pacemaker with patient cable, if necessary for the device offered - Two (2) Long patient cable/s for dual chamber, as necessary for the device offered. - One (1) Hook to hang from IV pole. - One (1) Protective cover, if necessary - One (1) Portable case. - Set of start-up consumables, if necessary. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-31: Medical electrical equipment - Part 2-31: Particular requirements for the basic safety and essential performance of external cardiac pacemakers with internal power source. | 3 |
| Item 23 (OR 34) | Patient drainage system | Product Description A device designed for continuous wound and chest drainage, including features like a fine control valve, fistula indicator, and bacterial filter. Electrical Requirements: - N/A Technical specifications: - Filling volume: At least 700 ml - Max. filling level up to overflow protection: At least 1000 ml. - Scaling in 10 ml steps up to max. - Max. vacuum: At least -10 to 0 kPa (-100 to 0 mbar) - Max. flow: At least 14 l/min - Weight: At most 3 kg (empty) - Dimensions (WxHxD): At most 300 x 350 x 200 mm Accessories: - Set of start-up consumables Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance | 1 |
| Item 24 (OR 37) | Sternal, oscillating saw | Product Description A device designed for sternotomy during a cardiovascular surgical intervention. Battery operated system. For pediatric and precision procedures. Electrical Requirements: - Operation: Rechargeable battery-powered. - Charger power requirements: according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Type: Sternum saw, battery-operated - Use: Must be able to be used for pediatric surgery. - Progressive Velocity: from 0 to at least 13.000 cpm. - Detachable and steam sterilizable saw system. - Single handpiece for multiple saw head. - Saw head included: - Oscillating pediatric saw - Sternum saw with a saw guide. - Cordless with removable battery for sterilization - Aseptic battery (no sterilization needed) - Trigger function with safety lock mode. Accessories: - Two (2) Rechargeable battery (preferable Li-ion battery) - Two (2) Autoclavable battery case or house - One (1) Autoclavable aseptic transfer shield for battery loading (if applicable) - One (1) Battery charger for 1 battery - One (1) Pediatric sternum saw guide or guard. - Two (2) Pediatric sternum saw blades. - One (1) Oscillating sternum saw blade. - One (1) Sterilization tray with fixations according to the offered system. - Set of start-up consumables, if necessary. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 7153-1: Surgical instruments — Materials — Part 1: Metals - ISO 10993-1 Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process. | 2 |
| Item 25 (OR 38) | Surgeon's headlamp | Product Description A device designed to be used as an adjunct light source during clinical or surgical procedures in a medical setting. Electrical Requirements: - Operation: Rechargeable battery-powered. Technical specifications: - LED lamp with optic - LED Expected Life Span > 50,000 hours - Illumination Range: 100.000 - 500.000 lux (25 cm) - Color temperature: 4500K. Preferable, adjustable between 4000K and 6100K. - Adjustable illumination spot size: beam angle up to 120 degrees - Uniform light distribution over the surgical field. - Mounting system: Adjustable and secure allowing firm and stable attachment of the lamp to the surgeon’s head. - Battery: - Autonomy of 8 hs. - Lithium based. - Fast and safe charging system. - Battery Expected Life Span > 300 Charges. - Material: - Lightweight and durable - Main Body: Aluminum alloy recommended. - Adjustable parts: High-quality plastic recommended. - Water Resistance: IPX7 or higher. - Weight: less than 350 g. Accessories: - One (1) Charger system at least 2 batteries - One (1) Battery pack - One (1) Charger Wires set. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal) in English and French. Preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class I devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests | 1 |
| Item 26 (ICU 3) | Dialysis Equipment for Neonate - Peritoneal dialysis system | Product Description A device intended to deliver a neonate peritoneal dialysis system. This sophisticated system offers a variety of therapy options, including hemodialysis, hemodiafiltration, and online hemodiafiltration, catering to a wide range of patient needs. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - HighVolumeHDF Capability: Ability to perform pre- and post-dilution hemodiafiltration for efficient removal of middle molecular toxins and cardioprotective hemodialysis. - AutoFlow Technology: Incorporation of AutoFlow for dialysate flow savings without compromising treatment efficacy, ensuring optimal use of resources. - EcoFlow Integration: Implementation of EcoFlow technology for efficient and sustainable workflows, contributing to resource optimization. - Heat Exchanger: Inclusion of a heat exchanger for temperature control during dialysis sessions, enhancing patient comfort and treatment effectiveness. - Service-Friendly Design: Emphasis on service-friendliness with features like interactive hydraulic flow charts for error diagnosis, easy maintenance access to machine components, and snap-lock technology for quick component exchange. - BCM-Body Composition Monitor: bioimpedance spectroscopy - Management of low volumes: - Sensitive machine settings and limits - Blood flow limited according to body weight. - Reduced bolus rate according to blood flow - Reduced blood flow of 50ml/min for a gentle start of dialysis - Adapted air and microbubble detection sensitivity. - Pediatric User setup - Pediatric AV blood line set with reduced fill volume of 114 ml. Accessories: - One (1) Online Clearance Monitor, if applicable. - One (1) Blood Temperature Monitor, if applicable. - One (1) Blood Volume Monitor, if applicable. - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Copy of the agreement between the bidder and the local service center for the provision of services - Manufacturer's authorization issued to the local service center for the right to provide maintenance services for the declared medical equipment (copy certified by the seal of the bidder and the local service center) - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 23500: Preparation and quality management of fluids for haemodialysis and related therapies — Part 1: General requirements. - ISO 11663: Quality of dialysis fluid for haemodialysis and related therapies. | 1 |
| Item 27 (ICU 18) | Anesthetic Monitor & Accessories | Product Description A device designed as an anesthetic gas monitor designed to meet respiratory gas monitoring needs. A portable anesthetic gas monitor that provides comprehensive respiratory gas monitoring for patients of all ages. It offers a range of monitoring capabilities including CO2 concentration, functional oxygen saturation and pulse rate among others. A device designed to meet the monitoring needs in anesthesia for various patient populations. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Anesthetic agent monitoring: CO2, and N2O - Sample Flow Rate: 200 ± 20 ml/min. - Accuracy for anesthetic agent: ± (0.20 Vol % + 15 % rel.) - Alarms (Visual and audible): - Apnea - inN2O - inCO2 high - etCO2 high and low - Battery Type: Rechargeable lithium-ion battery - Battery Operating Time: 60 minutes or more. Accessories: - One (1) CO2 sensor - One (1) SpO2 sensor - One (1) Anesthetic agent sensor - One (1) Battery pack - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical Electrical Equipment - General Requirements for Basic Safety and Essential Performance - IEC 60601-1-4: Medical Electrical Equipment - General Requirements for Safety - Collateral Standard: Programmable Electrical Medical Systems - IEC 60601-1-8: Medical Electrical Equipment - General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. - ISO 80601-2-55: Anesthetic Gas Monitors | 1 |
| Item 28 (ICU 26) | ICU Emergency & Transport ventilator with Acc & Options: “IPPV, SIMV, CPAP/ASB, Bibap, Mes. CO2” | Product Description Device designed to be used as time-cycled, volume-controlled, and pressure-controlled emergency and transport ventilator. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Ventilation Modes: The device offers VC-CMV, VC-AC, VC-SIMV, SpnCPAP, PC-SIMV+ - Settings: Pressure support, apnea ventilation, AutoFlow (optional), and NIV - Ventilation Respiratory Rate: Ranges from 2 to 60 breaths per minute. - Tidal Volume: Adjustable from 0.05 to 2.0 liters with high respiratory rate support - Gas Supply: Can be connected to a pipeline system or an O2 cylinder. - Dimensions and Weight: Reduce - Measures smaller than 300 x 200 x 200 mm and weighs less than 8 kg. - Display: Features a large, high-contrast display for easy reading of settings and parameters. - Wireless Connectivity: Supports IEEE 802.11g Wi-Fi standards for data export functionality. Accessories: - One (1) Breathing hose sets with flow measurement lines. - One (1) Ventilation valve - One (1) Pressure difference flow sensor - One (1) Elbow connectors - One (1) CO2 mainstream sensor - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 80601-2-12: This standard applies to the basic safety and essential performance of a ventilator, including its accessories, intended for patients dependent on mechanical ventilation in critical care environments or during transport within a healthcare facility. - ISO 10651-5: Specifies particular requirements for basic safety and essential performance of gas-powered emergency resuscitators. - ISO 17510: Pertains to medical devices for sleep apnea breathing therapy, including masks and application accessories. - ISO 80601-2-79: Focuses on the basic safety and essential performance of ventilatory support equipment for ventilatory impairment. - ISO 80601-2-80: Addresses the basic safety and essential performance of ventilatory support equipment for ventilatory insufficiency. | 1 |
| Item 29 (ICU 27) | ICU Ventilators, ADU/PED, mode, «IPPV,SIMV, CPAP /ASB, BIBAP Ventilation » | Product Description A turbine-driven ventilation system designed to provide quick response times and continuous high flow for both adult and pediatric patients. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E. Technical specifications: - Ventilation Modes - Volume controlled ventilation: ‒ VC-CMV ‒ VC-SIMV ‒ VC-AC ‒ VC-MMV - Pressure controlled ventilation: ‒ PC-CMV ‒ PC-BIPAP 1 / SIMV+ ‒ PC-SIMV ‒ PC-AC ‒ PC-APRV ‒ PC-PSV - Support of spontaneous breathing: ‒ SPN-CPAP/PS ‒ SPN-CPAP/VS ‒ SPN-CPAP ‒ SPN-PPS - Special procedures - Suction maneuver - Manual inspiration/hold - Medication nebulization - P0.1 - PEEPi - Therapy types - Invasive ventilation (Tube) - Non-invasive ventilation (NIV) - O2 -therapy - Respiratory rate (RR) - Adult 0.5 to 98/min - Pediatric patients, - Neonates 0.5 to 150/min - Inspiratory time (Ti) - Adults 0.11 to 10 s - Pediatric patients, - Neonates 0.1 to 10 s - Tidal volume (VT) - Adults 0.1 to 3.0 L - Pediatric patients 0.02 to 0.3 L - Neonates 0.002 to 0.1 L - Inspiratory flow - Adults 2 to 120 L/min - Pediatric patients 2 to 30 L/min - Maximum flow during non-invasive ventilation of neonates (Flow max) - 30 L/min - Inspiratory pressure (Pinsp) - Adults 2 to 120 L/min - Pediatric patients 2 to 30 L/min - Maximum flow during non-invasive ventilation of neonates (Flow max) - 0 to 30 L/Min - Pressure limitation (Pmax) - 2 to 100 mbar (or hPa or cmH2O) - Positive end-expiratory pressure (PEEP) - 0 to 50 mbar (or hPa or cmH2O) - Additional intermittent PEEP for sighs (ΔintPEEP) - 0 to 20 mbar (or hPa or cmH2O) - Pressure support (Psupp) - 0 to 95 mbar (or hPa or cmH2O) - O2 concentration (FiO2) - 21 to 100 Vol.% - Trigger threshold (Flow trigger) - 0.2 to 15 L/min - Automatic Tube Compensation - Flow Measurement - Minute volume measurement. - Expiratory minute volume, overall, not leakage corrected (MVe) - Inspiratory minute volume, overall, not leakage-corrected (MVi) Minute volume, leakage-corrected (MV) - Mandatory expiratory minute volume, overall, not leakage corrected (MVemand) - Spontaneous expiratory minute volume, overall, not leakage corrected (MVespon) Range 0 to 99 L/min BTPS. - Tidal volume measurement - Tidal Volume, leakage-corrected (VT) - Mandatory inspiratory tidal volume, not leakage corrected (VTimand) - Mandatory expiratory tidal volume, not leakage-corrected (VTemand) Spontaneous inspiratory tidal volume, not leakage-corrected (VTispon) - Range 0 to 5500 mL BTPS - Respiratory rate measurement - Respiratory rate (RR) - Mandatory respiratory rate (RRmand) - Spontaneous respiratory rate (RRspon) Range 0/min to 300/min - O2 measurement (inspiratory side) - Inspiratory O 2 concentration (in dry air) (FiO2) Range 18 to 100 Vol% - CO 2 measurement in main ﬂow (adults and pediatric patients only) - End-tidal CO 2 concentration (etCO2) Range 0 to 100 mmHg. - Displayed calculated values. - Range 0 to 1000 mbar/L/s (or hPa/L/s or cmH2O/L/s) - Dynamic compliance (Cdyn) - Range 0 to 1000 mbar/L/s (or hPa/L/s or cmH2O/L/s) - Leakage minute volume (MVleak) - Range 0 to 99 L/min, BTPS - Rapid shallow breathing index (RSBI) - Adults 0 to 9999 (/min/L) - Pediatric patients 0 to 9999 (/min/L) - Neonates 0 to 300 (/min/L) - Negative Inspiratory Force (NIF) - Range -80 mbar to 0 mbar (or hPa or cmH2O) - Occlusion pressure P0.1 - Range 0 to -25 mbar (or hPa or cmH2O) - Waveform displays: - Airway pressure Paw (t) -30 to 100 mbar (or hPa or cmH2O) - Flow (t) -180 to 180 L/min. - Volume V (t) 2 to 3000 mL - CO 2 (t) 0 to 100 mmHg - Alarms / Monitoring - Expiratory minute volume (MVe) / High / Low - Airway pressure (Paw)/ - High - Inspiratory O 2 concentration (FiO2) / High / Low - End-tidal CO 2 concentration (etCO2) / High / Low - Respiratory rate (RR) / High - Volume monitoring (VT)/ High / Low - Apnea alarm time (Tapn)/ 5 to 60 seconds, Oﬀ - Disconnection alarm time (Tdiscon) / 0 to 60 seconds - Performance data - Control principle / Time-cycled, volume-constant, pressure-controlled. - Length of intermittent PEEP / 1 to 20 expiratory cycles - Medication nebulization / For 5, 10, 15, 30 minutes, continuously (∞) - Inspiratory ﬂow / Max. 180 L/min, BTPS - Base ﬂow, adults / 2 L/Min - Base ﬂow, pediatric patients 3 L/Min - Base ﬂow, neonates / 6 L/Min - Inspiratory valve Accessories: - Set of start-up consumables NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits) in English and French. Preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIb devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 80601-2-12: This standard outlines the basic safety and essential performance requirements for a ventilator in critical care environments, ensuring compliance with ventilation accuracy, alarms, monitoring capabilities, electrical safety, mechanical performance, and environmental requirements. - ISO 18562-4: Focuses on biocompatibility evaluation of breathing gas pathways in healthcare applications, specifically testing for leachables in condensate to ensure patient safety. - ISO 10651-5: Specifies particular requirements for basic safety and essential performance of gas-powered emergency resuscitators. - ISO 17510: Addresses medical devices for sleep apnea breathing therapy, focusing on masks and application accessories. | 3 |
| Item 30 (ICU 28) | Lamp, examination | Product Description A device designed to provide a specialized source of light during examinations, diagnosis, or treatments of the patient. LED technology. Wall mounted. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E Technical specifications: - Wall mounted type: To be fixed to the wall adjacent to ICU bed/Station - Arm: flexible type gooseneck for high and position adjustments. - LED technology. - Light life: at least 50,000 hrs. - Illumination level: at least 50,000 lux at 0,5m of distance. - Color temperature: between 4,000 and 5,000 K. - Color rendering index: at least 90. - Adjustment of light Intensity. - All materials resistant to hospital-use disinfectants. Accessories: - All necessary elements to anchoring and fixing the lamp on the wall. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty"). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services"). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - Declaration of Conformity (DoC) according with regulation 2017/745 for Medical Devices Class I. - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - IEC 60601-2-41: Medical Electrical Equipment, Part 2-41: Particular Requirements for the Safety of Surgical Luminaires and Luminaires for Diagnosis. | 3 |
| Item 31 (ICU 29) | Non-invasive ventilator | Product Description Non-invasive respiratory support for spontaneously breathing patients. It consists of applying a humidified and heated high flow of air/oxygen, in adult and pediatric patients. Electrical Requirements: - Power requirements according to Mongolia standards: 220 VAC, 50 Hz. Power cord with plug type E - Rechargeable batteries with autonomy for at least 2 hours. Technical specifications: - Non-invasive respiratory support equipment should consist of at least: - Interface with the patient: through nasal cannula, sealed facial mask, nasal mask, or tracheostomy mask with adapter, according to the needs and age of the patient. - Inspiratory or respiratory tubing (for heated and humidified gas), which prevents the accumulation of condensation. - Air-Oxygen Mixer and Heater-Humidifier, all integrated in one device. - Capacity to generate high flow rates of air-oxygen mixtures, with adjustment in temperature and humidity. - Volumetric flow range: at least 3 l/min to 60 l/min - FiO2: 21% to 100% - Temperature range: at least 33°C to 37°C - External connection for air and oxygen source, from the hospital network or by cylinder. Preferably built-in compressor or turbine to provide compressed air. - Visual and acoustic alarms for at least: - Temperature. - Water level. - Disconnection of the patient circuit. - Oxygen supply. - System failure. - Mobile trolley for the entire system and its accessories, made of steel with anti-corrosive epoxy coating, aluminum, AISI 304 stainless steel or higher quality material. With: - Brakes at least in two (2) of the wheels. - Support for oxygen cylinder, - Solution holder pole with 2 hooks, minimum. Accessories: - Set of start-up consumables. NOTE: All the necessary accessories for the correct operation of the product, including all standard tools, cleaning, and lubrication materials, even if they are not included in these required technical specifications must be included. Accessories should be reusable when an option is available. The accessories will be used for both neonates and pediatric patients. Documentation requirement: - User manual must be provided (including operation instructions, maintenance and/or procedures for decontamination, storage conditions, safe disposal). In English, French and preferably also in Mongolian. - Service manual must be provided (including preventive maintenance and calibration procedures, equipment necessary for preventive maintenance and repair, diagrams, and circuits). In English, French and preferably also in Mongolian. - List of common spares and accessories with part numbers must be provided. - Manufacturer authorization. - Commitment Manufacturer letter, including: - at least two (2) years of full onsite warranty (according to the general requirements in the bid document under the clause “Warranty”). - at least five (5) years of Spare Parts availability (according to the general requirements in the bid document under the clause “Post-sale Services”). - training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian. (according to the general requirements in the bid document under the clause “Training”). Regulatory approvals required: - Valid Mongolian official Import license from the main manufacturer for receiving, installing, and maintaining medical equipment in Mongolia. - And at least one of the following regulatory approvals and certificates: - European Certificate of Conformity (CE) with Regulation 2017/745 or Directive 93/42 EC and Agreement Letter signed with the NB demonstrating the on-going MDR application, for Class IIa devices, or - FDA (Food and Drug Administration) of the USA that certifies marketing permission in the United States, or - Other regulatory bodies of an IMDRF founding member country such as Australia, Canada, or Japan. Safety & product Standards: Bidder shall furnish the documentary evidence to demonstrate that the good it offers meet the international safety & regulatory standards, providing in addition a signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the follow standards: - ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes. - IEC 60601-1: Medical electrical equipment - Part 1: General requirements for basic safety and essential performance - IEC 60601-1-2: Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests - ISO 80601-2-74: Medical electrical equipment Part 2-74: Particular requirements for basic safety and essential performance of respiratory humidifying equipment. | 1 |

General condition for all items:

- Whether it is included specifically in the technical specification or not, the installation, commissioning and training of personnel should be carried out by the supplier for all items.

# 2.2.	Schedule of Requirements

The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the goods for 5 years following commencement of the use of the goods by UNFPA.

| - List of Goods and Delivery Schedule |  |  |  |  |
| --- | --- | --- | --- | --- |
| Line Item | Description of Goods | Quantity | Unit of measure | 'Desired Delivery Schedule from date of Contract'' |
| Item 1 (OR 1) | Activated Clotting Time (ACT) Measuring device | 2 | Each | 50-70 days |
| Item 2 (OR 2) | Portable Blood Gas Analyzer | 2 | Each | 50-70 days |
| Item 3 (OR 3) | Extra Corporeal Circulation (ECC) Machine | 1 | Each | 50-70 days |
| Item 4 (OR 4) | Hypo-hyperthermia Machine | 1 | Each | 50-70 days |
| Item 5 (OR 33) | Patient Auto-Transfusion System | 1 | Each | 50-70 days |
| Item 6 (OR 5) | Infusion Rack (Syringe Pump and Infusion Pump including racks and trolley with holder pole with hooks) | 4 | Each | 50-70 days |
| Item 7 (OR 10) | Monitor/defibrillator + specific spoons for children + specific options + accessories | 2 | Each | 50-70 days |
| Item 8 (OR 11) | Suction system, surgery | 4 | Each | 50-70 days |
| Item 9 (OR 12) | Cardiovascular ultrasound imaging system | 1 | Each | 50-70 days |
| Item 10 (OR 16) | Electrosurgical system | 1 | Each | 50-70 days |
| Item 11 (OR 18) | Hand-held ultrasound imaging system | 1 | Each | 50-70 days |
| Item 12 (OR 19) | Brain Monitoring System | 1 | Each | 50-70 days |
| Item 13 (OR 20) | Monitor for cerebral oximetry | 1 | Each | 50-70 days |
| Item 14 (OR 22) | Multiparametric TRANSPORT patient monitor for adult and pediatric patients | 4 | Each | 50-70 days |
| Item 15 (OR 23) | Mounting system for monitoring arm or lateral support for ventilation machine | 1 | Each | 50-70 days |
| Item 16 (OR 24) | Slave Monitor 28’’ + Wall fixation Arm | 1 | Each | 50-70 days |
| Item 17 (OR 26) | Lamp, operating room, ceiling | 1 | Each | 50-70 days |
| Item 18 (OR 27) | Video-laryngoscope | 1 | Each | 50-70 days |
| Item 19 (OR 28) | Patient heater | 2 | Each | 50-70 days |
| Item 20 (OR 29) | Blood / fluid warmer | 2 | Each | 50-70 days |
| Item 21 (OR 31) | Patient Ventilator/Respirator | 1 | Each | 50-70 days |
| Item 22 (OR 32) | Portable external pacemaker | 3 | Each | 50-70 days |
| Item 23 (OR 34) | Patient drainage system | 1 | Each | 50-70 days |
| Item 24 (OR 37) | Sternal, oscillating saw | 2 | Each | 50-70 days |
| Item 25 (OR 38) | Surgeon's headlamp | 1 | Each | 50-70 days |
| Item 26 (ICU 3) | Dialysis Equipment for Neonate - Peritoneal dialysis system | 1 | Each | 50-70 days |
| Item 27 (ICU 18) | Anesthetic Monitor & Accessories | 1 | Each | 50-70 days |
| Item 28 (ICU 26) | ICU Emergency & Transport ventilator with Acc & Options: “IPPV, SIMV, CPAP/ASB, Bibap, Mes. CO2” | 1 | Each | 50-70 days |
| Item 29 (ICU 27) | ICU Ventilators, ADU/PED, mode, «IPPV, SIMV, CPAP /ASB, BIBAP Ventilation » | 3 | Each | 50-70 days |
| Item 30 (ICU 28) | Lamp, examination | 3 | Each | 50-70 days |
| Item 31 (ICU 29) | Non-invasive ventilator | 1 | Each | - days |

| - Consignee Address and Consignee-wise Quantity Distribution |  |  |  |  |
| --- | --- | --- | --- | --- |
| Line Item | Consignee Address | Contact person | Quantity | Unit of measure |
| Item 1 (OR 1). | National Center for Maternal and Child Healthcare (NCMCH), “Huvisgalchdiin” street, Bayangol district, Ulaanbaatar 16060, Tel: 976-11 362205 | Tsetsenbaatar Batsuuri, Procurement associate, UNFPA Mongolia, Tel: 976-11-353503 (ext3355), batsuuri@unfpa.org | 2 | Each |
| Item 2 (OR 2). |  |  | 2 | Each |
| Item 3 (OR 3). |  |  | 1 | Each |
| Item 4 (OR 4). |  |  | 1 | Each |
| Item 5 (OR 33). |  |  | 1 | Each |
| Item 6 (OR 5). |  |  | 4 | Each |
| Item 7 (OR 10). |  |  | 2 | Each |
| Item 8 (OR 11). |  |  | 4 | Each |
| Item 9 (OR 12). |  |  | 1 | Each |
| Item 10 (OR 16). |  |  | 1 | Each |
| Item 11 (OR 18). |  |  | 1 | Each |
| Item 12 (OR 19). |  |  | 1 | Each |
| Item 13 (OR 20). |  |  | 1 | Each |
| Item 14 (OR 22). |  |  | 4 | Each |
| Item 15 (OR 23). |  |  | 1 | Each |
| Item 16 (OR 24). |  |  | 1 | Each |
| Item 17 (OR 26). |  |  | 1 | Each |
| Item 18 (OR 27). |  |  | 1 | Each |
| Item 19 (OR 28). |  |  | 2 | Each |
| Item 20 (OR 29). |  |  | 2 | Each |
| Item 21 (OR 31). |  |  | 1 | Each |
| Item 22 (OR 32). |  |  | 3 | Each |
| Item 23 (OR 34). |  |  | 1 | Each |
| Item 24 (OR 37). |  |  | 2 | Each |
| Item 25 (OR 38). |  |  | 1 | Each |
| Item 26 (ICU 3). |  |  | 1 | Each |
| Item 27 (ICU 18). |  |  | 1 | Each |
| Item 28 (ICU 26). |  |  | 1 | Each |
| Item 29 (ICU 27). |  |  | 3 | Each |
| Item 30 (ICU 28). |  |  | 3 | Each |
| Item 31 (ICU 29). |  |  | 1 | Each |

| - List of Related Services and Completion Schedule |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| No. | Description of Service | Qty | Place where Services shall be performed | Final Completion Date(s) of Services |
| 1 | Installation and commissioning. The supplier performs the installation and commissioning for all items. | 1 | NCMCH | Within 1 weeks upon completion of delivery |
| 2 | Training on use, cleaning, disinfecting for healthcare staff and basic maintenance technical staff, in English and Mongolian for all items. Please also review the requirement included in the technical specification of each item. | 1 | NCMCH | Within 2 weeks upon completion of delivery and installation |
| 3 | After sale service. A minimum of two years of in-country after-sale services by an authorized local service provider. | 1 | NCMCH | 2 years |

# SECTION III: UNFPA General Conditions of Contract

UNFPA General Conditions of Contract can be found at:

http://www.unfpa.org/resources/unfpa-general-conditions-contract

# SECTION IV: UNFPA Special Conditions for Contracts

| CONTRACT PRICE | The prices charged for the Goods supplied and the related Services performed shall not be adjustable. |
| --- | --- |
| PERFORMANCE SECURITY | A Performance Security “Shall” be required. “The Performance Security in original shall be submitted within ten working days from the date of the contract” The amount of the Performance Security shall be 10 % of the Contract Price. The Performance Security shall be unconditional and irrevocable and in the form of either: - An unconditional Bank Guarantee - A Demand Draft - A Cashier’s Cheque - A Certified Cheque In the event of Suppliers submitting the Performance Security in the form of a Cheque or Demand Draft in favour of UNFPA, such documents shall be accompanied by a signed statement from the issuing bank on its letterhead indicating the validity period and confirming irrevocability of the Cheque or Demand draft during the required period. Banks issuing Performance Securities must be acceptable to the UNFPA Comptroller, i.e. they have to be banks certified by the Central bank of the country to operate as commercial bank. The Performance Security shall be denominated in the currencies of payment of the Contract, in accordance with their portions of the Contract Price, and shall have a validity period of forty-five (45) days after the date of delivery indicated in the Contract days. UNFPA reserves the right to request an extension of the Performance Security. Discharge of the Performance Security shall take place upon expiry of the Performance Security or upon confirmation of receipt of the Goods by the Consignee. The Performance Security shall then be returned to the Supplier by UNOPS. |
| WARRANTY | The warranty period shall be 24 months. Details on Warranty Services required are included in Section II: Technical Specifications and Schedule of Requirements. |
| GOODS AND SERVICES DEFINED | Goods are hereinafter deemed to include, without limitation, equipment, spare parts, commodities, raw materials, components, customized and standard software as required, intermediate products and products which the Supplier is required to supply under the Purchase Order. Services are to include design, installation and commissioning, training services, technical assistance and warranty services as required to supply in the Purchase Order. |
| AFTER SALES SERVICES | A minimum of two years of in-country after-sale services by an authorized local service provider. |
| TRANSPORTATION AND FREIGHT | Responsibility for transportation of the Goods shall be as specified in the INCOTERMS. Partial shipment is allowed. Transshipment is allowed. |
| SHIPPING AND PAYMENT INSTRUCTIONS | Access the following link for shipping and payment instructions: Shipping Instructions |
| LIQUATED DAMAGES | In the event of a Contract being issued and in case the Vendor fails to deliver all the goods by the date or dates of delivery specified in the Purchase Order, UNFPA reserves the rights to claim liquidated damages from the Vendor and deduct 1% of the value of the goods pursuant to the Purchase Order per additional week of delay, up to a maximum of 10% of the value of the Purchase Order. The payment or deduction of such liquidated damages shall not relieve the Vendor from any of its other obligations or liabilities pursuant to any current Long-Term Agreement or Purchase Order. |

# SECTION V: Bidding Forms

The following checklist is provided as a courtesy to Bidders. Please use this checklist while preparing the bid to ensure that your bid contains all required information. This checklist is for the Bidder’s internal reference and does not need to be submitted with the bid.

| ACTIVITY | LOCATION | YES / NO/ NOT APPLICABLE | REMARKS |
| --- | --- | --- | --- |
| Have you read and understood all of the Instructions to Bidders in Section I of the bidding documents? | Section I |  |  |
| Have you reviewed and agreed to the UNFPA General Conditions of Contract? | Section III |  |  |
| Have you reviewed and agreed to the UNFPA Special Conditions for Contracts? | Section IV |  |  |
| Have you completed the Bid Confirmation Form? | Section V, 1 |  |  |
| Have you completed the Bid Submission Form? | Section V, 2 |  |  |
| Have you completed the Bidder’s Identification Form? | Section V, 3 |  |  |
| Have you completed the Performance Statement Form? | Section V, 4 |  |  |
| Have you completed the Product Item Overview Form? | Section V, 5 |  |  |
| Have you completed and signed the Price Schedule Form? | Section V, 6 |  |  |
| [Delete if not applicable] Have you completed the Joint Venture Partner Information Form? | Section V, 7 |  |  |
| Have you reviewed all of the relevant contract form(s)? | Section VI |  |  |
| [Delete if not applicable] Have you prepared a copy of your valid manufacturing license from the country of manufacturing? | Section 1, Sub-Clause 10.2, b. |  |  |
| [Delete if not applicable] Have you prepared a copy of your company’s registration in the country of operation? | Section I, Sub-Clause 10.2, b. |  |  |
| Have you prepared a copy of the previous year’s audited company Balance and Financial Statements? | Section I, Sub-Clause 10.2, d. |  |  |
| Have you provided written confirmation that your company is neither suspended by the United Nations system nor debarred by the World Bank Group? | Section I, Sub-Clause 2.4 |  |  |
| Have you prepared documentary evidence that the goods conform to the technical specifications and standards specified in Section II Technical Specifications and Schedule of Requirements? | Section I, Sub-Clause 10.3, a. |  |  |
| [Delete if not applicable]Have you prepared product catalogues containing pictures of the product(s)? | Section I, Sub-Clause 10.3, c. |  |  |
| [Delete if not applicable] Have you prepared the manufacturer’s technical product specifications or data sheets? | Section I, Sub-Clause 10.3, d. |  |  |
| [Delete if not applicable] Have you provided the results of any testing carried out on the products? | Section I, Sub-Clause 10.3, a. |  |  |
| [Delete if not applicable] Have you provided any copies of current certificates such as GMP/Quality, FSC/CPP, manufacturer’s ISO certificate for the product, manufacturer’s CE certificate, USA510k, Japan QS standard, etc. as stated in the Technical Specifications and Schedule of Requirements, in Section II? | Section I, Sub-Clause 10.3, g. |  |  |
| [Delete if not applicable] Have you provided a copy of the valid authorization letter issued by the manufacturer for each product, if you are not the manufacturer? | Section I, Sub-Clause 10.3, h. |  |  |
| Have you furnished a list of full particulars, regarding the available sources and current prices of space parts, special tools, etc., necessary for the proper and continuing functions of the goods within the Product Item Overview Form, Section V, 5? | Section I, Sub-Clause 10.3, i. |  |  |
| Have you provided a copy of any of your company’s environmental or social policies, and any related documentation? | Section I, Sub-Clause 10.4 |  |  |
| Have you reviewed the UN Global Compact requirements? | Section I, Sub-Clause 10.4 |  |  |
| Have you sealed and marked the bids according to Instructions to Bidders Clause 16 (hard copy bids) or Clause 17 (electronic bids)? | Section I, Sub-Clause 16 & 17 |  |  |
| If submitted electronically, is the file size of the bid less than 10MB? (If the file size is above 10MB, refer to Instructions to Bidders Sub-Clause 17.4) | Section I, Sub-Clause 17.4 |  |  |
| Have you noted the bid closing deadline? | Cover letter, #5 |  |  |
| Have you provided information on annual sales turnover during any one of the last three years to be at least equal to the contract value (from Financial Statements)? | Section I, Sub-Clause 32.3 Financial Capability, a. |  |  |
| [Delete if not applicable] Have you provided documentary evidence that the Bidder has successfully completed at least one similar contract within the last five years for supply of goods? | Section I, Sub-Clause 32.3 Financial Capability, c. |  |  |
| [Delete if not applicable] Have you provided contact details of commercial banks and names of contact persons from whom UNFPA could seek feedback? | Section I, Sub-Clause 32.3 Financial Capability, d. |  |  |
| [Delete if not applicable] Have you provided sufficient documentation of your company's ability to undertake the contract, i.e., - List of similar contracts executed for other clients including contract details. - Evidence that the Bidder possesses experience in the geographical area. - At least three years of experience in performing similar contracts. | Section I, Sub-Clause 32.3 Experience & Technical Capacity, b. |  |  |
| [Delete if not applicable] Have you provided sufficient documentation of your company’s managerial capability? - Details of company’s managerial structure. - Quality assurance systems in place. | Section I, Sub-Clause 32.3 Experience & Technical Capacity, c. |  |  |
| [Delete if not applicable] Have you demonstrated that your company has manufactured and satisfactorily supplied similar goods to a similar extent of the quantity as mentioned against each schedule during any one of the last three years and the goods should have been in use satisfactorily with no adverse report? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, d. |  |  |
| [Delete if not applicable] Have you supplied Client's certificates in support of the satisfactory operation of the goods as specified above? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, e. |  |  |
| [Delete if not applicable] Have you supplied data to support that your company has production capacity to perform the contract and complete the supplies within the stipulated delivery period or data to support that it has an installed annual production capacity for the specific item to match the quantities required? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, f. |  |  |
| [Delete if not applicable] Have you provided evidence that your company is in the continuous business of manufacturing/supplying and providing after sale services for goods similar to those offered during the last three years prior to bid opening date? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, g. |  |  |
| [Delete if not applicable] Have you provided a brief write-up, backed up with adequate data, explaining available capacity and experience in the manufacture and supply of the required products within the specified time of completion after meeting all their current commitments? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, h. |  |  |
| [Delete if not applicable] Have you provided confirmation that all the facilities exist at the factory for inspection and testing and that these facilities will be made available to the purchaser or his representative for inspection? | Section I, Sub-Clause 32.3 Experience & Technical Capacity,i. |  |  |
| [Delete if not applicable] Have you disclosed any instances of previous past performance that may have resulted in adverse actions taken against the Bidder and the manufacturers whose products are being offered by the Bidder, in the last five years? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, j. |  |  |
| [Delete if not applicable]Have you given a list of full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the equipment for a reasonable period of time following the installation? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, k. |  |  |
| [Delete if not applicable] Have you provided any legally enforceable authorization from the manufacturer assuring full guarantee and warranty obligations as per the tender conditions for the goods offered? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, l. |  |  |
| [Delete if not applicable] Have you provided documentation of your company’s experience, as authorized by the manufacturers, in supplying and providing after sales service for similar goods to the extent of at least 20 percent of the quantities indicated in the tender requirements in any one of the last three years (and that the goods have been in satisfactory operation)? | Section I, Sub-Clause 32.3 Experience & Technical Capacity, m. |  |  |

# 1. Bid Confirmation Form

[Complete this page and return it prior to bid opening]

|  |  | Date: |
| --- | --- | --- |
| To: | UNFPA TSETSENBAATAR BATSUURI | email: batsuuri@unfpa.org |
| From: | [Company name] |  |
|  | [Contact person] |  |
|  | [Telephone] |  |
|  | [Email address] |  |
|  | [Postal address] |  |
| Subject: | ITB No.: UNFPA/DNK/ITB/24/014 |  |

YES, we intend to submit a bid.

NO, we are unable to submit a bid in response to the above mentioned Invitation to Bid due to the following reason(s):

( )	The requested products and services are not within our range of supply

( )	We are unable to submit a competitive bid for the requested products at the moment

( )	The requested products are not available at the moment

( )	We cannot meet the requested specifications

( )	We cannot offer the requested type of packing

( )	We can only offer FCA prices

( )	The information provided for quotation purposes is insufficient

( )	Your ITB is too complicated

( )	Insufficient time is allowed to prepare a quotation

( )	We cannot meet the delivery requirements

( )	We cannot adhere to your terms and conditions (please specify: payment terms, request for performance security, etc)

( )	We do not export

( )	Our production capacity is currently full

( )	We are closed during the holiday season

( )	We had to give priority to other clients’ requests

( )	We do not sell directly, but through distributors

( )	We have no after-sales service available in the recipient country

( )	The person handling bid is away from the office

( )	Other (please specify)

If UNFPA has questions to the Bidder concerning this NO BID, UNFPA should contact Mr./Ms._________________, phone/email ________________, who will be able to assist.

# 2. Bid Submission Form

[The Bidder shall fill in this form in accordance with the instructions indicated. No alterations to its format shall be permitted and no substitutions shall be accepted.]

Date: [insert date (as day, month and year) of Bid Submission]

ITB No.: UNFPA/DNK/ITB/24/014

To: TSETSENBAATAR BATSUURI, UNFPA

Dear Sir / Madam,

We the Undersigned have examined and have no reservations to the Bidding Documents No. UNFPA/CC/YY/NNN and amendments We hereby offers to supply, in conformity with the Bidding Documents and in accordance with the Delivery Schedules specified in the Schedule of Requirements, the following goods and related services ___________________________________________________ which are subject to UNFPA General Conditions of Contract and other terms and conditions as specified in the document.

We agree to abide by this bid for a period of [Select between 30-90 days depending on the type of good/commodity] days from the date fixed for opening of bids in the Invitation to Bid, and it shall remain binding upon us and may be accepted at any time before the expiration of that period.

We have no conflict of interest in accordance with Instructions to Bidders Sub-Clause 2.3;

Our firm, its affiliates or subsidiaries—including any subcontractors or suppliers for any part of the contract—have not been declared ineligible by UNFPA, in accordance with Instructions to Bidders Sub-Clause 2.4;

We understand that you are not bound to accept the lowest evaluated bid or any other bid that you may receive.

Dated on .............day of ......................................[year].

| Signature: | ……………………………………………………………… [insert signature of person whose name and capacity are shown] |
| --- | --- |
| In the capacity of: | ……………………………………………………………… [insert legal capacity of person signing the Bid Submission Form] |
| Name: | ……………………………………………………………… [insert complete name of person signing the Bid Submission Form] |
| Company: | ……………………………………………………………… [insert name of company] |

# 3. Bidders Identification Form

Bid No. UNFPA/DNK/ITB/24/014

- Organization

| Company/Institution Name |  |
| --- | --- |
| Address, City, Country |  |
| Telephone/FAX |  |
| Website |  |
| Date of establishment |  |
| Legal Representative: Name/Surname/Position |  |
| Legal structure: natural person/Co.Ltd, NGO/institution/other (please specify) |  |
| Organizational Type: Manufacturer, Wholesaler, Trader, Service provider, etc. |  |
| Areas of expertise of the organization |  |
| Current Licenses, if any, and permits (with dates, numbers and expiration dates) |  |
| Years supplying to UN organizations |  |
| Years supplying to UNFPA |  |
| Production Capacity |  |
| Subsidiaries in the region (please indicate names of subsidiaries and addresses, if relevant to the bid) |  |
| Commercial Representatives in the country: Name/Address/Phone (for international companies only) |  |

- Quality Assurance Certification

| International Quality Management System (QMS) |  |
| --- | --- |
| List of other ISO certificates or equivalent certificates |  |
| Presence and characteristics of in-house quality control laboratory (if relevant to bid) |  |

- Expertise of Staff

| Total number of staff |  |
| --- | --- |
| Number of staff involved in similar supply contracts |  |

- Contact details of persons that UNFPA may contact for requests for clarification during bid evaluation

| Name/Surname |  |
| --- | --- |
| Telephone Number (direct) |  |
| Email address (direct) |  |

P.S.: This person must be available during the next two weeks following receipt of bid

# 4. Performance Statement Form

(For the last five years)

Bid No. UNFPA/DNK/ITB/24/014

Name of Bidder: ____________________________

| Order No. & Date | Client | Contact person/phone | Description & quantities of ordered items | Value of order (USD) | Date of completion |  | Satisfactory completion |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | As per contract | Actual |  |

To be attached: Documentary evidence (client’s letter or certificate) in support of satisfactory completion of above orders.

| _____________________________________ | ______________________________ |
| --- | --- |
| Signature and seal of the Bidder | Date |
| _____________________________________ | ______________________________ |
| Countersigned by and seal of Chartered Accountant | Date |

# Product Item Overview Form

| Item No. | Description and minimum /mandatory specifications | Description of items offered and Bidder’s statements on deviations (To be completed by the bidder) | Compliant? (Y/N) (To be completed by UNFPA during evaluation) |
| --- | --- | --- | --- |
| 1 | [Bidder is asked to copy here technical specifications from Annex I which is offered] | [Bidder is asked to write here technical specifications from Annex I for each offered item and statements on deviations] |  |

# 6. Price Schedule Form

[The Bidder shall fill in these Price Schedule Forms in accordance with the instructions indicated. The list of line items in column 1 of the Price Schedules shall coincide with the list of goods and related services specified by UNFPA in the Schedule of Requirements.]

| BIDDER’S TOTAL PRICES (Price & Currency to be entered by Bidder): |  |
| --- | --- |
| TOTAL FIRM FCA PRICE |  |
| TOTAL FIRM DAP PRICE |  |
| TOTAL PRICE FOR SERVICES (if applicable) |  |
| FREIGHT COST PER 20/40 FT CONTAINER (if applicable) |  |

| BIDDER’S PRICES FOR GOODS (Price & Currency to be entered by Bidder): |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| ITEM/ LOT | DESCRIPTION OF THE GOODS | QTY (a) | CURRENCY: |  |  |  |
|  |  |  | UNIT PRICE FCA (b) | UNIT PRICE DAP (c) | TOTAL PRICE FCA (a)x(b) | TOTAL PRICE DAP (a)x(c) |
| 1. |  |  |  |  |  |  |
| 2. | Insert more rows if necessary |  |  |  |  |  |
| 3. | or delete if too many |  |  |  |  |  |
| 4. |  |  |  |  |  |  |
| 5. |  |  |  |  |  |  |

| BIDDER’S PRICES FOR SERVICES (Price & Currency to be entered by Bidder): |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| ITEM/ LOT | DESCRIPTION OF THE SERVICES | COUNTRY OF ORIGIN | QUANTITY AND PHYSICAL UNIT (a) | UNIT PRICE (b) | TOTAL PRICE PER SERVICE (a)x(b) |
| 1. | e.g. Comprehensive Annual Maintenance Contract |  |  |  |  |
| 2. | Insert more rows if necessary |  |  |  |  |
| 3. | or delete if too many |  |  |  |  |
| 4. |  |  |  |  |  |
| 5. |  |  |  |  |  |

| #### BIDDER’S DELIVERY DATA |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Country of origin of offered products: | Item 1 |  |  |  |  |
|  | Item 2 | Insert more rows in each section if necessary |  |  |  |
|  | Item 3 | or delete if too many |  |  |  |
|  | Item 4 |  |  |  |  |
|  | Item 5 |  |  |  |  |
| FCA point(s) of delivery for offered products: | Item 1 |  |  |  |  |
|  | Item 2 |  |  |  |  |
|  | Item 3 |  |  |  |  |
|  | Item 4 |  |  |  |  |
|  | Item 5 |  |  |  |  |
| Delivery time (FCA from date of order): | Item 1 |  |  |  |  |
|  | Item 2 |  |  |  |  |
|  | Item 3 |  |  |  |  |
|  | Item 4 |  |  |  |  |
|  | Item 5 |  |  |  |  |
| Shipment dimensions of offered products (including package): |  | Gross weight | Total volume | Containers (if applicable): |  |
|  |  |  |  | Number | Size |
|  | Item 1 |  |  |  |  |
|  | Item 2 |  |  |  |  |
|  | Item 3 |  |  |  |  |
|  | Item 4 |  |  |  |  |
|  | Item 5 |  |  |  |  |
|  | Total |  |  |  |  |

| BIDDER'S SIGNATURE AND CONFIRMATION OF THE ITB |  |
| --- | --- |
| PROVIDED THAT A PURCHASE ORDER IS ISSUED BY UNFPA WITHIN THE REQUIRED BID VALIDITY PERIOD, THE UNDERSIGNED HEREBY COMMITS, SUBJECT TO THE TERMS OF SUCH PURCHASE ORDER, TO FURNISH ANY OR ALL ITEMS AT THE PRICES OFFERED AND TO DELIVER SAME TO THE DESIGNATED POINT(S) WITHIN THE DELIVERY TIME STATED ABOVE. |  |
| Exact name and address of company COMPANY NAME ADDRESS PHONE NO. FAX NO. EMAIL ADDRESS OF CONTACT PERSON OTHER EMAIL ADDRESSES | AUTHORIZED SIGNATURE DATE NAME OF AUTHORIZED SIGNATORY (TYPE OR PRINT) FUNCTIONAL TITLE OF SIGNATORY WEB SITE |

# 7. Joint Venture Partner Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below.]

Date: [insert date (as day, month and year) of Bid Submission]

ITB No.: UNFPA/CC/YY/NNN

Page ________ of ______ pages

1. Bidder’s Legal Name: [Insert Bidder’s legal name]

2. JV’s Party Legal Name: [Insert JV’s Party legal name]

3. JV’s Party Country of Registration: [Insert JV’s Party country of registration]

4. JV’s Party Year of Registration: [Insert JV’s Part year of registration]

5. JV’s Party Legal Address in Country of Registration: [Insert JV’s Party legal address in country of registration]

6. JV’s Party Authorized Representative Information Name: [Insert name of JV’s Party authorized representative] Address: [Insert address of JV’s Party authorized representative] Telephone/Fax numbers: [Insert telephone/fax numbers of JV’s Party authorized representative] Email Address: [Insert email address of JV’s Party authorized representative]

7. Attached are copies of original documents of: [Check the box(es) of the attached original documents] € Articles of Incorporation or Registration of firm named in 2, above, in accordance with Instructions to Bidders Sub-Clauses 3.1 and 3.2. € JV Agreement, or letter of intent to enter into such an Agreement, signed by the legally authorized signatories of all the parties

# SECTION VI: Contract Forms

# Bank Guarantee for Advance Payment

[Insert one of the following: No advance payment shall be made. / The bank, as requested by the successful Bidder, shall fill in this form in accordance with the instructions indicated.]

Date: [insert date (as day, month, and year) of Bid Submission]

ITB No: UNFPA/DNK/ITB/24/014

[bank’s letterhead]

Beneficiary:	[insert legal name and address of UNFPA]

ADVANCE PAYMENT GUARANTEE No.: [insert Advance Payment Guarantee no.]

We, [insert legal name and address of bank], have been informed that [insert complete name and address of Supplier] (hereinafter called "the Supplier") has entered into Contract No. [insert number] dated [insert date of Agreement] with you, for the supply of [insert types of goods to be delivered] (hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the contract, an advance is to be made against an advance payment guarantee.

At the request of the supplier, we hereby irrevocably undertake to pay you any sum or sums not exceeding in total an amount of [insert amount(s) in figures and words] upon receipt by us of your first demand in writing declaring that the supplier is in breach of its obligation under the Contract because the supplier used the advance payment for purposes other than toward delivery of the goods.

It is a condition for any claim and payment under this guarantee to be made that the advance payment referred to above must have been received by the Supplier on its account [insert number and domicile of the account]

This guarantee shall remain valid and in full effect from the date of the advance payment received by the supplier under the contract until [insert date].

_____________________ 
[signature(s) of authorized representative(s) of the bank]

# Performance Security

[Insert one of the following: No Performance Security shall be requested.

Or

The bank, as requested by the successful Bidder, shall fill in this form in accordance with the instructions indicated]

Date: [insert date (as day, month, and year) of Bid Submission]

ITB No. and title: [insert no. and title of bidding process]

Bank’s Branch or Office: [insert complete name of Guarantor]

Beneficiary: [insert legal name and address of UNFPA]

PERFORMANCE GUARANTEE No.:	[insert Performance Guarantee number]

We have been informed that [insert complete name of Supplier] (hereinafter called "the Supplier") has entered into Contract No. [insert number] dated [insert day and month], [insert year] with you, for the supply of [description of Goods and related Services] (hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the Contract, a Performance Guarantee is required.

At the request of the Supplier, we hereby irrevocably undertake to pay you any sum(s) not exceeding [insert amount(s) in figures and words] upon receipt by us of your first demand in writing declaring the Supplier to be in default under the Contract, without cavil or argument, or your needing to prove or to show grounds or reasons for your demand or the sum specified therein.

This Guarantee shall expire no later than the [insert number] day of [insert month] [insert year], and any demand for payment under it must be received by us at this office on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees, ICC Publication No. 458, except that subparagraph (ii) of Sub-article 20(a) is hereby excluded.

[signatures of authorized representatives of the bank and the Supplier]

# Contract Forms

The following sample contract forms are available on the UNFPA procurement website:

- Purchase Order

- Contract for Professional Services

- Long Term Agreement

----------

Reference: www.timeanddate.com/worldclock

The bank shall insert the amount(s), either in the currency(ies) of the contract or a freely convertible currency acceptable to UNFPA.

Insert the delivery date stipulated in the contract delivery schedule. UNFPA should note that in the event of an extension of the time to perform the contract, UNFPA would need to request an extension of this guarantee from the bank. Such request must be in writing and must be made prior to the expiration date established in the guarantee. In preparing this guarantee, UNFPA might consider adding the following text to the Form, at the end of the penultimate paragraph: “We agree to a one-time extension of this guarantee for a period not to exceed [six months/one year], in response to UNFPA’s written request for such extension, such request to be presented to us before the expiry of the guarantee.”

The Bank shall insert the amount(s) specified in the SCG and denominated, as specified in the SCG, either in the currency(ies) of the Contract or a freely convertible currency acceptable to UNFPA.

UNFPA should note that in the event of an extension of the time to perform the Contract, UNFPA would need to request an extension of this Guarantee from the Bank. Such request must be in writing and must be made prior to the expiration date established in the Guarantee. In preparing this Guarantee, UNFPA might consider adding the following text to the Form, at the end of the penultimate paragraph: “We agree to a one-time extension of this Guarantee for a period not to exceed [six months] [one year], in response to UNFPA’s written request for such extension, such request to be presented to us before the expiry of the Guarantee.”