## Specification final
Spicification for medical Equpement
NO | Item Description | Unit of Measure | Quantity
1 | Anesthesia device regulator | unit | 12
Medical gas regulator to regulate the pressure: rate of medical gas supplied frome cylender to a equipment, maximum 4 bare, for medical oxygen.
The regulator body should be Chrome plated brass body , Steel or equivalent.
The product to be supplied shall be new, unused and conform to the standards as specified in technical specifications.
The equipment proposed shall be of highest quality and produced by well known manufacturers. The equipment shall carry the name and quality label of the manufacturer and fulfill the standards in force.
2 | Anaesthesia Machine, mobile station with Ventilator and Monitor | unit | 13
Anesthesia Workstation to be used for delivering anesthesia agents to the patients during surgery and monitors the vital signs and ventilates the patient.
The machine should be suitable for all ages and body weights of patient, the unit should consist of gas supply and control circuit
breathing and ventilation circuit, scavenging system and patient monitoring capabilities
Pipeline gas inlets: three gas system with three gas Oxygen, Nitrous Oxide and Medical Air
Gas Cylinder Yokes: O2, N2O and Air
Double scale flowmeter with high and low flow and minimal flow provisions.
System should be designed such that all components are integrate
Should have an independent Oxygen flow meter for Oxygen delivery and an integrated variable flow suction unit.
Anesthesia machine should have high grade reinforced fiber frame free from oxidation
Station should have at least three drawers and an additional writing surface that can be easily accessed
Drawers shall all have the ability to lock, and shall be easily removed for the purposes of cleaning and sterilization
Pipeline, cylinder and Airway pressures should all be displayed on color coded gauges and be visible at all times during operation
Should have provision to attach 2 cylinders 1 each for O2 and N2O.
Should have facility of delivering basal flow of oxygen on switching on the machine.
System should have a second user accessible port for extraction of Anesthetic gas when using a nonrebreathing patient circuit.
System should also provide the option of returning sample gas to the scavenging system with a dedicated port.
A single pneumatic/electric on/off switch should activate the gas flow and vaporization.
The unit should have a battery backup facility for the ventilator in the event of power loss and should operate for a minimum of one hour.
In the event of complete power loss and battery failure it shall still be possible to manually ventilate and deliver an aesthetic agent.
Should have unlockable Oxygen flush to deliver oxygen flow of approximately 40l/min.
Should have built in safety features like O2 failure alarm, N2O cutoff, Low O2 pressure etc.,
The frame should have integrated power outlets to supply a minimum of four external devices.
Should have locking of the front castors by a single central brake mechanism.
System should be designed to mounted on pendants.
AGSS – Anesthetic Gas Scavenging System, should be in place
Gas Flow
The unit shall have a mechanical hypoxic guard system to control the ratio of Oxygen and Nitrous
It shall be possible to deliver Air with only basal flow oxygen independent of the above mentioned hypoxic control
Gas flow shall be controlled mechanically to avoid errors during power failure and electronic malfunction
Visual display of the gas flow shall be by physical means independent of electrical power.
Cascade or dual flow tubes should be available for all gases to allow suitable resolution and accurate control at low total fresh gas flows
Flow meters should have backlight and antiglare illumination.
The unit should have an independent measurement and display of fresh gas flow offering safety for low and minimal flow anesthesia.
A bag arm with height and positional adjustment shall be available
Vaporizers
The unit should accommodate two vaporizers for anesthetic agent delivery to allow easy selection of agent to be used - Two vaporizers, Agents: Isoflurane and Halothane with interlock
Vaporizer should be selectable type; tool free installation and vaporizer of our choice can be mounted with interlocking facility to allow operation of only one vaporizer at one time.
Vaporizers supplied with the unit shall be maintenance free for the life of the product.
Breathing System
All parts of the breathing system that are in contact with patient gas should be latex free and autoclavable
Should not require tools when dismantled for cleaning and sterilization.
Should accept large and small volume absorber canisters.
The ventilator bellows shall be clearly visible and should ascend on expiration to provide a quick visual indicator for system leaks
Breathing system should have the option of C02 Absorber bypass control that will allow the absorber canisters to be removed without introducing system leaks.
Should have bag / vent selecting valve integrated onto the absorber and should automatically turn on the ventilator when positioned to vent mode.
Ventilator
Ventilator should be pneumatically driven, electronically controlled and should be ascending bellows type.
Ventilator should automatically change drive gas should there be a gas depletion.
Ventilator shall have a large color display with touch screen user interface.
Ventilator should have the following ventilation abilities, volume control, decelerating flow pressure control, SIMV with pressure support and pressure support.
Ventilator should be capable of ventilating diverse range of patient groups from neonates to patients with restrictive airways with tidal volume range between 20 ml to 1500 ml with single bellows system.
Assisted modes of breathing should be flow triggered.
Ventilator shall have an active proportional exhalation valve to prevent the potential of over delivery during pressure modes of ventilation.
Ventilator should have a leak and compliance test that can be done independently of the full system
On switching on, the ventilator system should be able to and shall give the user a choice of doing a unit test or bypassing in the case of an emergency.
Ventilator shall compensate for fresh gas flow and compliance of the entire circuit dynamically.
Measurement at the patient end of the circuit (sensor at the patient end) should be provided to compensate for small leakages and compressible volume variability that occur during ventilation.
User should also have the option of setting a pre set compliance correction where similar circuits are used constantly
Should provide constant fresh gas flow into the breathing circuit during the inspiratory phase as mandatory
Ventilator should have the ability to set and store a hospital default as well as individual user preferences for easy selection of ventilation parameters and include screen layout, alarm preferences and settings.
Apnea alarms must be user adjustable to allow for all operating conditions and phases during Anesthesia.
Ventilator should have the ability to display and store Patient spirometry loops including Flow Volume and Pressure Volume curves.
Ventilator should also display waveforms for flow and airway pressure.
Ventilator shall display measured fresh gas independent of the flow meters.
Integrated Monitoring system:
Anesthesia Monitoring system should be of modular type and capable of monitoring adult, pediatric and neonatal patients. Should be from the same manufacturer as of the anesthesia system.
Monitor should have minimum 19” independent flat panel display with multi color touch screen user interface to ensure all parameters are visible simultaneously.
Module rack / housing should be independent and shall be able to be placed near to the patient.
Should have facility to monitor ECG, NIBP, SpO2, Respiration, Invasive pressures, temperatures (2), Capnography and Bispectrality index. Should have Cardiac output port enabled
Should have automatic identification and measurement of anesthetic agents, EtCo2, O2 and N2O and MAC value.
Cardiac output monitoring facility using thermo dilution technology with all accessories.
ECG should have capability for 3, 5 and / or 10 lead monitoring and should have built in arrhythmia monitoring on all 12 leads
Inbuilt ST segment analysis and arrhythmia detection for all the leads should be available.
Should have hemodynamic, oxygenation and drug dose calculations.
EtCO2 should have both mainstream and side stream in one module.
Respiration should be available with Cardiovascular Artifact filter.
OCRG(oxy cardio respire gram) should be available for monitoring neonates.
Alarm parameter should flash red in the presence of high priority alarms (e.g. ventricular fibrillation and asystole) and flash yellow in the presence of medium or low priority alarms (e.g. noisy signal, etc.)
24 hours trend data should be displayed.
Position of the displayed waveforms and color of the waveform must be user configurable.
Monitor shall permit the optional ability to receive and display information from other patient devices such as ventilators, infusion pumps and other standalone devices.
Should be supplied with necessary accessories for adult, pediatric and neonatal accessories.
Accessories and spares
ECG / respiration: 5 lead ECG cable and lead wire set and 10 lead ECG cable and lead wire set per monitor
NIBP: Adult 2 sizes and Pediatric 2 sizes and neonatal 1 size per monitor
SPo2 Sensor: Adult sensor with cable, pediatric sensor with cable and neonatal sensor with cable per monitor
IBP: Include 10 nos of disposable pressure transducers with bracket and interface cable per monitor
Temperature: Skin and nasopharyngeal probes per monitor
BIS: 25 nos of disposable sensors per monitor
3 | Bedside cabinet and overbed table | Each | 38
Bedside cabinet: Laminated top, resistant to water and heat Body is resistant to corrosion Has one drawer and one cabinet. With four casters (two with brake) WxLxH ≥ 40x 40x 70 cm Overbed table: Table is covered by ABS material (resistance and tough) The table has 4 wheels which can rotate 360 around Surface is rust & corrosion-free. WxL ≥ 40x80 cm Adjustable height ≥ 85 - 110 cm.
4 | Bench-top Steam Sterilizer | Each | 14
Volume: ≥40 liters
Temperature range: 100°C to 134°C
Interior chamber: Stainless steel
Programmable microprocessor control system
Digital display for temperature and time
Timer: 1 min. to 100 hours
Proportional heating system
Manometer for chamber pressure
Electrical and mechanical pressure safety systems
2 shelves
Mains power 220- 240V ±6% , 50 Hz British Standard 3 Pin Power Plug / Cable
5 | Binaural Stethoscope | Each | 124
Stethoscope, Binaural type, Cardiac and Pulmonary examination
Receptor shall be of dual membrane in single monobloc Stainless Steel casing
Adult membrane shall be 45mm Dia. Children membrane shall be 36mm Dia.
Membrane shall be of high resolution type (Epoxy glass fiber). Polypropylene rings shall be inserted in the assembly to decrease cold effect
Tubing in moulded PVC, Monobloc
Ear frame shall be made of Stainless Steel and orientable as well as interchangeable. Nickel-steel spring
Ear caps made of elastomer rubber
Indicative acoustic precision: Cardiac: -3dB.50/500Hz) and Pulmonary: -8dB.600/1500Hz
Supplied with
4 Ear caps made of elastomer rubber
1 Storge Case
6 | Blood Bank Refrigerator | Each | 18
Technical Specifications
1 Temp range-should have adjustable temperature control range from +1 degree to +8 degree C, factory preset at 4 degree C.
2 Capacity should accommodate 300 standard blood bags for each of 350ml/450 ml capacity and storage internal volume should be 700 liters.
3 Refrigerator system-
a)The system should have high density CFC –free urethane foam insulation to protect cabinet from ambient temperature fluctuation.
b)The system should have positive, forced, air circulation to maintain temperature uniformity at all shelf levels, with quick recovery +/- 1 degree C.
c)The system should have sensors for activating automatic defrost cycle to minimize the frost build up.
d)The system should have automatic condensate removal with no requirement for separate drainage lines.
4 Internal construction should be made up of high grade stainless steel (min 22 G) External construction Corrosion resistant sheet at least I mm thickness.
5.Internal Temp Control
a)System should have temperature control range from +1 degree C to +8 degree C.
b)Temperature control resolution should be better than 0.1 degree C.
c)Cooling down time of max of 150 min on half load.
6 External ambient temp should perform in ambient temp up to +43 degree C.
7 Door System should lockable double glass doors for better safety
8 Safety system:
a. system should have large and clear Digital displays for the set/run parameters.
b. The system should have weekly chart recorder temperature chages
c. The system should have key operated set point for the added security.
9 . Alarms.
a)System should have audible/visual warnings for over-temperature under temperature and power failure with visual status reports on critical functions.
b)System should have battery backup and connections for remote alarm contacts
10 Should have adjustments for uneven bases. The adjustments should be easy to use like rotating a screw at the legs in the base.
11 Scratch resistant internal lining of the cabinet (stainless
steel or aluminium).
Standards and Safety
1 Should be FDA , CE,UL or BIS approved product
2 Should comply with WHO/UNICEF Specification Reference: BTS/RF.1
3 Test and inspections as per WHO Procedure reference: Laboratory Test Procedure: Standard Test Procedure: BTS/ Proc/ 3.
4 Should comply with International Electromagnetic Compliance standards like IEC OR EMC Directives.
7 | Bowl Washing st.st 150 Mm Dia x 120 Mm Deep | Each | 53
• Basic container for nursing and surgical care • Contaner, round. • Smooth surface. • Inner diameter: approx. 75-85 mm. • Height: approx. 36-48 mm. • Capacity: approx. 150 to 190ml. • Thickness: approx. 0.75-0.85mm. • Material: Austenitic stainless steel. • Austenitic stainless steel composition: 18-20% chromium, 8-10% nickel. • Reusable.
8 | Bowl Washing st.st 250 Mm Dia x 120 Mm Deep | Each | 48
• Basic container for nursing and surgical care • Contaner, round. • Smooth surface. • Inner diameter: approx. 85-95 mm. • Height: approx. 46-58 mm. • Capacity: approx. 250 to 300 ml. • Thickness: approx. 0.75-0.85mm. • Material: Austenitic stainless steel. • Austenitic stainless steel composition: 18-20% chromium, 8-10% nickels. • Reusable."
9 | Bowl Washing st.st 350 Mm Dia x 120 Mm Deep | Each | 92
• Type Bowl, round, stainless steel • Container, round. • Smooth surface. • Internal diameter: approx. 100 - 120 mm. • Height: approx. 60-75 mm. • Capacity: approx. 350to 450 ml. • Thickness: approx. 0.75-0.85mm. • Material: Austenitic stainless steel. • Austenitic stainless steel composition: 18 - 20% chromium, 8 - 10% nickel. • Reusable." • Materials:"High resistance to corrosion (tropical environment) • Support column, ring structure and bowl: Austenitic stainless steel 18/10
10 | Ceasarian Set | Sets | 70
| B.P.Handle No.3 | 4
| Dissecting Forceps 23 cm Serrated | 8
| Dissecting Forceps 23 cm 1x2 teeth | 8
| Forceps Artery Spencer Wells 12.5 cm | 32
| Forceps Artery Spencer Wells 22.5 | 10
| Forceps Kocker STR 16.5 cm | 8
| Forceps Tissue Allis | 8
| Forceps Towel Cross Action | 12
| Green Armytage Forceps | 20
| Needle Holder Mayo's | 4
| Needle Holder Thomson Walker 20 cm | 4
| Perforator Scissor | 2
| Placenta Willet Forceps | 2
| Retractor Doyen's | 4
| Retractor Morris | 4
| Retractor Small | 4
| Safety Pin | 12
| Scissors Mayo's CRV 15 cm | 4
| Scissors Mayo's STR 16 cm | 4
| Scissors Mclntosh's CRV 16 cm | 4
| Sim's Speculum Medium Size | 4
| Spong Holding Forceps Ramply | 16
| Suction Tube Nozzle | 4
| Umbilical Scissor | 2
| Instrument Tray with cover" | 1
11 | Automatic hematology Analyzer complete blood counter (CBC) | Each | 38
Fully Automatic and compact hematology analyzer, 5-part differentiation of WBC.
Open vial sampling mode requiring maximum 20ul blood.
23 + 4 parameters including RDW-SD.
Throughput: at least 60 samples/hour
At least 4 quality control programs and 60 files storage
Storage Capacity: 40,000 samples with data and graph
Technology: Laser scatter, Flow cytometry, Impedance.
Should have facility to analyses CBC, CBC + Diff mode.
Fully customized report formats including microscopic exam results
Automatic diluent dispensing for capillary samples.
Should have syringe based measurement technology.
Linearity range Performance
WBC: 0 - 99.99 x 109/L
RBC: 0 - 8.00 x 1012/L
HGB: 0-250 g/L
HCT: 0 - 67%
PLT: 0 - 1000
Reproducibility performance
WBC: ≤ 2.0%
RBC: ≤ 1.5%
HGB: ≤ 1.5%
MCV: ≤ 1.0%
PLT: ≤ 4.0%
Supplied with Personal Computer, printer and UPS, preferably sine wave based with maintenance free batteries with duration of two hours.
Hematology Analyzer kits/reagents/solutions - supplier to include list of kits, reagents, solutions, accessories and consumables for 3-month operation. The list to include description, part number and price of each for selection.
Power 220 V, 50 Hz
Service manual, operation manual and technical data in English with all necessary passwords without any obligation.
12 | Couch examination ,fold headrest mattress | Each | 54
Couch Examination, Fold. Headrest Mattress
Shall be 2 sections type
Adjustable backrest thru manual ratchets made of chrome-plated steel.
Rigid top, 2 sections foam-padded mattress with fire safety protection. Removable, washable and protected by Durable cover, anti-bacterial, anti-static, acid-resistant and waterproof for easy Clean
Chrome-Plated heavy duty tubular frame , Floor protection by rubber studs
13 | Cupboard, metal, 2 doors, Lockable | Each | 107
Cupboard metal Steel 2-Doors Storage
composition from one piece , contain of 2 doors , entirely metallic
Equipped with 6 metallic adjustable shelves , Heavy- gauged metal framing with clip and bolted assembly , can status each one in different level .
made from entirely metallic , from matter Heavy-gauged , Epoxy-Painted and oven-baked components
Cupboard contain from 2 metal lockable doors with heavy duty hinges , to make easy to open and close doors and accuracy
Equipped the doors key kind one upset , doorknob each other door designer and position form proportionate to make easy to open & close doors
height : 180 cm at least
width : 80 cm at least
depth : 40 cm at least
14 | Delivery beds | Each | 47
Table, Delivery, with accessories
Three sections, metal construction, epoxy coated, with four castors of diameter five inch with central brake, heavy duty
ABS mattress base, 80 mm durable hygienic mattress, collapsible side rails, leg section mattress, adjustable foot supports, retractable foot section. adjustable foot supports, retractable foot section.
Adjustments:
* Height: from 550 to 900 mm
* Back section: 70° gas spring assisted
* Trendelenburg angle: + 12° & -3°
Complete with mattress, sliding st/st bowl, leg rest with straps, side rails, adjustable IV stand, 5 protective bumpers.
Dimensions of bed: ≥1900 * 900 mm
15 | Delivery bed hydrolic | Each | 60
For Delivery & Obstetric Use.( Hydraulic Type)
Dimensions: 65 - 92cm Heihgt, 170cm Length Including Auxiliary Table.
Tilt Angle: Pelvis Up/Down 20 Degrees, Back70 Degrees & Lateral 2 Degrees.
Manual adjustment of back section with ratchets with safety stops
Made of steel, Epoxy-Painted with sturdy sampling of material
Fully equipped with proprietary mattresses (2 sections), Removable soft pads
Supplied with: Legs support, Arm support, Shoulder support , Placenta basin made of Stainless Steel and 2 grab handles
16 | Delivery Set | Sets | 128
| B.P. Handle No.3 | 1
| Cross Action Towel Forceps | 4
| Delivery Forceps (Midwifery) Anderson | 16
| Forceps Dissecting 1x2 teeth 22 cm | 2
| Forceps Dissecting non toothed 22 cm | 4
| Forceps Haem. Green Armytage | 4
| Forceps Artery Spencer Wells 12.5 cm | 2
| Forceps Artery Spencer Wells 23.5 cm | 4
| Needle Holder Mayo | 4
| Scissor Mayo STR | 2
| Scissors Mayo CVD | 4
| Scissors Mcintosh CVD | 4
| Sim's Speculum Medium Size | 4
| Sponge Holding Forceps Ramply 23 cm | 4
| Umbilical Scissor | 8
17 | Digital Glucometer |  | 78
Manual coagulation instrument with extended test and data management capabilities. The test manu includes all routine coagulation test. (PT,APTT,TT, Fibrinogen and Factorassays).Operation is simplified by audio and visual prompts that indicate ending of warming and activation cycles .Storage wells with stirring capability allow for preheating of reagents . Built-in data management
Sample volume one drop
Test time less than 10 seconds
With memory
Set complete with prick pen
Lancets
100 glucose strips
battery operated
18 | Thermometer, clinical | Each | 126
Safety standard CE mark approval or ISO Design & quality Compact , heavy duty and high quality Type Thermometer Clinical (C) For oral, underarm, or rectal use Prob : Oral probe, Probe well Measurement type less than 30s. Measurement rang +32 ~ +43 Measurement accuracy ± 0.1°C\ display Liquid crystal display Ton sound when final temperature is reached Required Automatic shut off. Required
19 | Foetal heart detector | Each | 98
Microprocessor controlled equipment. LCD display with visualization of at least fetal heart rate. Integrated fetal heart processing software. Ultrasound working frequency in the range 2MHz -10% to 3MHz +10%. Sensitivity to detect fetal heart beats of at least a 10-12 weeks fetus. At least two high sensitivity equipment compatible probes provided: 2 and 3 MHz. Heart rate measurement range not smaller than 50-210 bpm with resolution not higher than 2 bpm. Audio output reproduction of the fetal heart rate with integrated speaker and with headphones. Audio volume control system integrated. At least 1 of system compatible headphones provided. At least one integrated serial port for PC connection and data transmission. Memory storage capacity of at least 4 hours of working data. Cable for data transmission. 1 pair of spare system compatible headphones. Carry case for easy transportation.
. 1 transducer 2 MHz ± 10%
. Gel, electroconductive. 2x250g
. 3 rechargeable batteries (if the model proposed functions with rechargeable batteries)
. Operated by rechargable battery with Charger
Mains power 220- 240V ±6% , 50 Hz British Standard 3 Pin Power Plug / Cable
20 | Double Foot Steps | PCS | 72
Material of Construction Heavy duty Stainless Steel 304 or better.with elecrostatic powder, two steps
Steps surface covered by anti slip material & Floor protection by rubber studs
Guiding Dimentions:( 50* 50 * 45 cm) at least
21 | Drum, Sterilization.Medium | PCS | 113
Drum Sterilizing Instruments, "Material: Stainless Steel, Medium Cylinder st.st 340 Dia X 235 Mm H.
22 | Drum,Sterilization.Large | PCS | 110
Drum Sterilizing Instruments, "Material: Stainless Steel, large Cylinder st.st 390 Dia X 280 Mm H.
23 | ECG Machine 6 channel | PCS | 27
Safety standard FDA CLEARANCE Required OR CE MARK (MDD) Required with treding sales in Europe Union, USA, Astralia, Canda & Japan Design & quality "Multi Channel ECG Machines, With trolley with Four casters 2 with brakes New model, able for high load & hard work" Heavy duty and high quality A Resting Electrocardiograph Unit It shall have the following minimum features: 12-channel ECG recorder with data storage capability Display: Minimum 7” or more, color LCD Displayed screen data : Heart rate, patient name and ID number, time, display of 12 channels, speed, amplitude and filter settings, alerts. Patient data entry including : Name, ID number, age Operating modes: automatic and manual. Gain: 5, 10 and 20 mm/mV. Storage: not less than 50 ECGs. With Vectorcardiography (preferable) Arrhythmia analysis and ECG interpretation: Comprehensive ECG interpretation software package shall be included Advanced detection and classification of different types of arrhythmias for adults and pediatrics Shall detect and classify arrhythmias including (but not limited to): Ventricular Fibrillation, Ventricular Tachycardia, Supraventricular Tachycardia, Ventricular bigeminy, Sinus Bradycardia, Sinus Tachycardia, Asystole, ST segment analysis: Advanced monitoring and analysis of ST segment deviation Continuous for all monitored leads Networking and connectivity: Shall be networkable to share data with hospital networks such as Hospital Information System (HL7 compliant) and ECG Management Systems via Local area Network (LAN) and/or wireless networks (coordinate with client and IT contractors) Alarm Visual indication of: Battery status Loose electrode contact. System status Built-in Recorder: Capability of simultaneous printing of 12 ECG channels or more High-resolution, digital thermal writer, Speed: 12.5, 25 and 50 mm/s. Prints out reports including: heart rate, P duration, QRS duration and PR interval. Filters Line frequency and Muscle filtration through digital filters Defibrillator protection Standard Accessories Shall be included : Original mobile cart with at least one drawer, and articulated cable holder 12 lead ECG cable adults and pediatrics Reusable electrodes kit for adults and pediatrics Adult and pediatric disposable electrodes Printer paper Case, carrying Packs, thermosensitive. 50mmx20m Electrodes, Limb. clamp type Electrodes, Suction Adaptors, Clip.Alligator.4mm Gel, electroconductive. POWER SUPPLY : LINE POWER, VAC : 100-240, 50/60 Hz Battery: With internal rechargeable batteries Battery run time: 120 minutes or more . Installation Installation & Commissioning must be done by manufacturer engineer Training Service Training for one MWC Bio-Engineer shall be provided within the first year of warranty User /Nurses training, by Specialist from the Suplier. Warranty/After Sale Service Two Years or more comprehensive onsite warranty of entire system (Spares and labour) Essential requirement: • The model should be FDA approved and/ or CE marked with treding sales in Europe, USA, Canda & Japane • That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model. • Spare parts list with code NO • The supplier must ensure the availability of expertise service and maintenance. • Uninterrupted availability of spare parts and repair of next ten years must be assured. • Mention the number (with addresses and phone numbers) of installations of quoted units in Yemen Maintenance: "Preferred less maintenance needed. 2 years free maintenace.or more" Service manual operation manual {Hardcopy & Softcopy} application software and interface connection Included. Other specification provide all items according to the supplier's catalog
24 | Electronic Baby scale |  | 64
Safety standard FDA approval or CE marking Design Heavy duty and high quality. Type Infant Clinic Min patient weight 20 Kg Safety trayechanical. Required Auto zero switch Required Electric adapter and battery operated Re-chargeable battery preferred Auto calibration Required Warranty Minimum of 1year Power supplay 100 to 240 V ~ ±10%, 50/60 Hz Single phase Electrical Safety class 1 Other specification provide all items according to the supplier's catalog
25 | Electrosurgical cautery | Unit | 12
Type Electrosurgical Unit with cart
Machine must have independent HF generator for monopolar & Bipolar, so that both can be operated simultaneously
It must have frequency compensation circuit it maintain the output
Should have 400 W of cutting power, 300W of coagulation power & 50 W of Bipolar coagulation power
Should have digital display for monopolar & Bipolar power setting
Should have opinion of bleeding the current at least 5 steps
Should have Silicon rubber patient plate to prevent patient burns
Should have two paddle excision protected foot switch one for cut & second for coagulation
Should have automatic start mode for Bipolar coagulation
Must have separate audible tones for cut & coagulation
Unit must be capable to cut off the current delivery in case of disc- connectivity of patient plate with the unit. Should give audio & visual alarm in that case
Unit should have separate connector to connect the resectoscope
Power: 220 VAC, 50/60 Hz
The Unit should be quoted along with the following Accessories
1) Monopolar set of accessories (1 set), consist of
Hand control pencil
Double pedal Footswitch
Patient plate
Electrode set of 5
2) Bipolar Footswitches
3) Bipolar Forceps
4) Bipolar Cable
26 | Emergency Cart. (Crash Cart ) with complete accessories | Each | 34
For Storage and transportation of equipment and medications of emergency use
Designed with drawers and cupboard and two universal rails on both sides.
Two Oxygen cylinders at the side of cart one of them with pressure regulator,oxygen
flowmeter, humidifier , oxygen mask and silicon hose, the other with pressure regulator,
suction injector, 2 liters autoclavable collecting bottle with over float protection, and silicon patient hoses.
Mobile On at least 10cm castors.
27 | Episiotomy Set | Each | 64
Technical Specifications:
Material: Stainless Steel
Reusable or Disposable: Reusable
Sterile or Non-Sterile: Non-Sterile
Latex-Free: Yes
Rust Prevention Procedure: Passivation
Ultrasonic Cleaned: Yes
Matt-Polished: Yes
Lubricate: Yes
Usage: Left Hand or Right Hand
Tests Performed: Boil Test, Performance Test, Shape Test
QC Passed: Yes
Quality: high quality , Safety, Ease of Use.
Each set contains the following:
1) Amino hook 19 cm (1),
2) Instruments Basket for sterilization wired 25x15x5 cm (1),
3) Artery Forceps, Kocher, 14 cm, 1x2 teeth straight (2),
4) Tissue Forceps, stand, 1x2 teeth, straight 14.5cm (1),
5) Needle Holder, 18cm, standard (1),
6) Mayo Scissor
28 | Examination delivery bed | Each | 54
Heavy duty
3-Sections Chrome stainless steel
Collapsible side rails ( perfarred )
leg section mattress, adjustable foot supports, retractable foot section
Adjustmented back section 0 to 80 degree gas spring ( perferred ) ot gradated assisted.
Safe for mother ( Anti-fall-down base )
Complete with:
- Washable mattress 5 to 7cm height minimum.
- Sliding st/st bowl
- Leg rest with straps
- Arm support
- Adjustable IV stand
- Stainless Steel Bowl
- Foot steps ( Two Stairs ) stainless steel or high quality material
29 | Examination lamp mobile on castors | Each | 71
Operating Light, mobile, AC/DC complete with Emergency Power
Portable O.T Light LED with adjustable height,cool and shadow-less light
Luminous intensity: ≥100,000 lux and a working field of 150 mm
Colour rendering index: Ra 95
Red-colouring index R9: > 97
Light colour: ~4,500 K
Continuous Dimming .10 – 100 %
Backup Battery >2 hrs.
Heat absorbing color correcting filters provided with plastic diffusers for shadow reduction. Provided with sterilizable handle Power: 220V and built-in backup battery
30 | Fingertip pulse oximeter | Each | 97
• Displays SpO2%, pulse rate, pulse waveform and pulse intensity
• Auto power off
• Universal finger clip suitable for adults and children
• Requires 2 x AAA batteries
• SpO2 accuracy: ±2% (70 to 100%)
• Pulse rate measurement range: 25 to 250 bpm
• Pulse rate accuracy: ±1% or ±1 bpm whichever is greater
• Operating temperature: +5 to +40°C
• Storage temperature: -20 to +55°C
• CE Marked
31 | Gynecology Examination Set | Sets | 103
| Instrument Tray With Lid Stainless Steel 225x125x50mm | 1
| Kidney Basin With Lid Stainless Steel 200x30 | 1
| Speculum Vaginal, Cusco's - Medium | 2
| Speculum Vaginal, Cusco's - Small | 2
| Speculum Vaginal, Cusco's - Large | 1
| Speculum Vaginal, Ducbill, Simm's Double - Ended Small | 1
| Speculum Vaginal, Ducbill, Simm's Double - Ended Medium | 1
| Curette Uterine Blunt- Sharp Combined 240 mm | 2
| Curette Uterine Simm's Pliable Shaft 250mm Small 6mm | 2
| Curette Uterine Simm's Pliable Shaft 250mm Medium 10mm | 2
| Curette Uterine Simm's Pliable Shaft 250mm Large 13mm | 2
| Dilator Cervix Hegar Double Ended 3-18mm | 2
| Forceps Ovum Greenhaigh's 240mm | 2
| Sound Uterine, Simposon's 290mm | 2
| Sterilizing and storing case St. St. with filters 260x150x50 mm | 1
32 | Hospital Bed Mattres | Each | 169
2 Sections mattress ( Heavy duty, washable, not less than7- 10 cm height, ) Mattress made of Polyurethane foam with washable case. Mattress: Pressure-reducing mattress for full body pressure redistribution. The mattress should meet the flammability standards. Mattress-padding waterproof, washable plastic and flame retardant. Cover unzips for cleaning.an anti-microbial agent to be contained in the surface of the mattress to assist in prohibiting growth of bacteria and fungus.
33 | Hospital bed with mattress | Each | 144
2 Sections are moved by two mechanical levers Steel Constraction stainless steel
Heavy duty desgine.
Rubber Bumpers at all 4 corners
Removable plastic ends ( Head & Legs )
Siutable castors with lock
Folding-down Sides C2(pair)
Supplied with:
IV rod, 2 hooks with adjustable height
Sheet ( easy to remove for cleaning )
Pillow Foam
Pillows made of Polyurethane foam.
About 350 X 700mm
With water proof plastic or PVC cover
With Pillow cases made of cotton. Dimensions: 400X750mm ± 10%
Supplied with:
Washable Cover Cotton Cases ( 2 pcs. For each pillow )
Sheet ( easy to remove for cleaning )
Mattress made of Polyurethane foam with washable case.
2 Sections mattress heavy duty, washable, not less than thickness not less than (7-10 cmheight)
Removable plastic ends ( Head & Legs )
Bedside cabinet and overbed table
Laminated top, resistant to water and heat
Body is resistant to corrosion
Has one drawer and one cabinet.
With four casters (two with brake)
WxLxH ≥ 40x 40x 70 cm
Overbed table:
Table is covered by ABS material (resistance and tough)
The table has 4 wheels which can rotate 360 around
Surface is rust & corrosion-free.
WxL ≥ 40x80 cm
Adjustable height ≥ 85 - 110 cm.
34 | Infant incubator | Each | 23
Infant incubator
An infant incubator provides a closed, controlled environment that warms an infant by circulating heated air over the skin. The heat is then absorbed into the body by tissue conduction and blood convection. Ideally, both the skin and core temperatures should be maintained with only minor variations.
Operational Requirements
High quality with humidity and servo controlled double walled with cabinet incubator.
Microprocessor controlled, easy access control panel with feather touch switches
With a facility to elevate base to offer adjustable range
Facility with both servo control as well as air temperature control and servo humidifier
Accommodates shelves and IV poles.
The quality of the material used should very high and crystal transparent
Super quality microprocessor based control system - self test functions are performed
System required complete with Oxygen port with tubing and Gel Mattress.
Technical Specifications
Continuous bed tilt up to 8° on either sides
Head end raise facility with auto lock.
Both visual and audible alarms for Patient and control and high / low temperature alarm and Air circulation / probe / system / power failure alarm.
Humidity control alarm.
Facility to take x-ray and weight without removing baby.
Facility to display and trends of temperature information on compatible monitors with other physiological parameter
Iris port for tubing, probes, leads. 4cm thick gel mattress, easily cleanable.
With at least 4” diameter caster wheel with swivel in all directions and with front lockable wheels. Two shelves cabinet with door. Weight 90-100 kg.
Patient control (Servo) mode – 35° -37°C. and Air Control (Manual mode)- 20°C to 39°C
Air velocity less than 10 cm/sec with inner wall.
Temperature variability less than +/-0.2°C. and Temperature resolution 0.1°C
Average oxygen input concentration range 5-15 liters/min or 25-70%.
Humidification adjustable electronically with digital display Standard
Double wall canopy with Six hand ports with elbow operated flaps with separate ports for tubing.
CO2 flushing, according to lEC 601-2-19 / 105.1 Maximum C02 concentration inside incubator 0.2%
Servo control for Oxygen with integrated monitoring Lifecare Limited
Air filter :- 0.3 micron
Built in weighing scale with sensitivity of + 1 gm
Mattress should be radiolucent
Provision for X ray cassette holders
2 drawer storage facility and two platforms for keeping monitors, able to bear at least 5 kg weight each.
System Configuration Accessories, spares and consumables
Two sets of extra non disposable temperature sensors and humidification sensors.
Power: 220-240VAC, 50/60Hz with suitable UPS with 30 min Backup for complete system
35 | Infant Warmer | Unit | 25
A mains electricity (AC-powered) mobile device that contains an infrared (IR) heating element(s) designed to emit controlled, evenly distributed overhead heat to the body of a newborn/infant patient requiring supplemental heat. This device is equipped with wheels so that it can easily be moved to different areas of a room, ward, or department.
Detailed requirements
*Controlled by microprocessor or microcontroller. *Radiation type heater operated by both timer and skin temperature regulation, selectable between the two (Manual and Servo controlled). *Heater power switch with a range of 0 to 100%. *Self-test facility on power on required. * Patient temperature range control from 35°C/95°F to 37°C/98.6°F * Temperature resolution at least ±0.5 deg C * Apgar Score monitor. *Trendelenburg and inverse trendelenburg positions. * Maximum heater power output not less than 200 W.
Displayed parameters
*Visual and audible alarms for patient high/low temperature and probe / system / power failure. *Heater power indicator to be clearly visible. *Skin Temperature display to be clearly visible. *Visual/audible alarm for Patient Check Reminder.
User adjustable settings
*Patient temperature range control from 34°C/93.2°F to 38°C/100.4°F *Temporal alarm silencer.
Components(if relevant)
*Bassinette to allow tilting of infant bed,clear view of infant and provide easy access to the infant from at least three sides. *Swing side panels to access infant table. *Height adjustable infant table, minimum height of which to be at least 80 cm. *Tilting table mechanism > 12° *Mattress made by a material flame retardant, washable, antibacterial and resistant to: corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium. * Equipment compatibility with heated mattresses. * Bassinet size not less than 65 x 40 cm. *Drawer or shelf to be included for storage *Mounting fittings for separate suction pump and bottled oxygen supply *Examination light with dedicated power switch to be included X-Ray chasis tray underneath bassinette. *At least one unit integrated monitor shelf which could support up to at least 20 Kg. * Equipment composed by at least: an open-bassinet, heater unit and control unit.
Accessories
Two extra mattresses. Three extra sets of sensors. Base for external oxygen cylinder. Neonatal manual resuscitator.
Consumables / reagents
Oxygen bottle of approximately 10 litters, 200 bars, portable and provided with at least the following accessories: flux meter, humidifier and oxygen tubes. Reflective sensor patch (box of 50).
Spare parts
Two sets of spare fuses (if replaceable fuses used) One complete set of warmer element(s) Replacement examination light bulb
Other components
Reusable Resuscitator: a) integrated handle resuscitator with self-inflating bag, used in pulmonary resuscitation of infants and children. b) maximum volume delivered not less than 300 ml; c) oxygen reservoir bag capacity not less than 1500 ml. 100% latex-free; d) transparent valve, transparent face mask.
Warranty:Up to 24 months
Documentation requirements
User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided Trade-in with manufacturer if available. Plastic recycling.
36 | Infusion pump |  | 20
Volumetric Infusion Pump complete unit with pole clamp Open system, accepts any infusion line (admin set)
Flow rate should be from 0.1 to 999 ml/hr with 0.1ml/hr increment
It should have choice of Micro & Macro drops
Volume Limit should be more than 9900ml *
It should use conventional tubing (4 mm), PVC
Immediate display of choice of programming after purge
Must have function for immediate forward & reverse of screen/programming parameters
Programming: ml/hr,time x volume limit. Weight x concentration x dose
Must have display of flow rate observable from a distance of 10-12 feet
Must have Occlusion Pressure from 0 to 120 kPa
KVO should be 0.1 ml to 5 ml/hr
Bolus rate should be more than 900 ml/hr
Must have Special functions like Titration , Fluid Balance
Time limit from 1min to 99hrs & 59 minutes.
Should have pre alarms – Ultrasonic sensor type
Visual / sound alarms like End of Battery Power , Incorrect flow rate, Door open, No drops, Occlusion , programming error, Running in KVO, Fluid balance etc.
Flow rate deviation with standard tubing should not be more than +/- 5% of the programmed flow rate.
Rechargeable battery with life ≥ 4 hours @ 125 mL/hr
Line power: 220 VAC 50/60Hz
37 | Instrument Trolley st.st shelves | PCS | 88
Safety standard FDA Approval or CE Mark Material of Construction Heavy duty, Stainless Steel medical grade 304 or better, Anti-corrosion and easy to clean and high quality Shelves three shelves in the top and bottom , edge around the shelves . Plat form 2-3 side railing on top shelf round pipe guardrail for three side Dimensions Depth = 45-50 Cm Approx. Width = 60-70 Cm Approx. Height = 80-90 Cm Approx. loading capacity 150-200 Kg Castors Four casters ,two with brakes and Durable, noiseless and stable casters Accessories & options Please specify
38 | IUD insertion set | Each | 99
Technical Specifications: Material: Stainless Steel Reusable or Disposable: Reusable Sterile or Non-Sterile: Non-Sterile Latex-Free: Yes Rust Prevention Procedure: Passivation Ultrasonic Cleaned: Yes Matt-Polished: Yes Lubricate: Yes Usage: Left Hand or Right Hand Tests Performed: Boil Test, Performance Test, Shape Test QC Passed: Yes Quality: high quality , Safety, Ease of Use. IUD insertion set Stainless Steel ,1) Cup/bowl/gallipot,(1) Forceps, Schroeder-Braun uterine tenaculum, 9.75 inches (24.8 cm), (1) Forceps, Rampley, sponge, Foerster, straight, 9.5 inches (24.1 cm),(1) Sound, uterine, Sims, 13 inches (33 cm),(1) Scissors, operating,Mayo, curved, 6.75 inches (17.1 cm), (1) Bivalve Speculum, vaginal, Graves, medium (1.38 inches [3.5 cm] x 4 inches [10.2 cm])
39 | IV Stand | Each | 176
Quality: high quality, Safety, Ease of Use.
Drip Stand or Rod for I.V. procedures , Adjustable height
Very stable stand with 5 castors, 5 stars-shaped, smooth running
Rod made of chrome stainless steel tubing,.
Rod equipped with 4 hooks perfarred
40 | Microscope binocular: | Each | 14
Microscope teaching. inclined binocular head , Body-Inter changeable, inclined Binocular body, 360º rotatable head Large sturdy stand with coarse and fine focusing knobs 30°-45° inclined eye piece tube Interpupillary distance adjustment. Diopter compensation on both eyepieces Wide filed eye point eye pieces 10. Built-in mechanical stage. with low positioned coaxial controls. Low positioned coaxial controls. Bright filed Abbe condenser (NA 1.25) with iris diaphragm and filter holder. Built-in low voltage illuminator 230 V 150Hz 12 VI 20 W or 6 V20 W halogen lamp. Plano concave mirror, adaptable to illumination exit. Quadruple reversed nosepiece. 4 semi plan achromatic objectives: 4, 10, 40, spring loaded 100 ' (oil) sprig loaded. Complete with: Ground glass filter Spare fuses. Main cable , One spare lamp. Bottle of immersion oil. dust cover Original wooden or leatherette case. Instruction manual in English. - Power supply: 220 - 240 V/50 Hz.
41 | Laboratory Stool Swivel | Each | 26
Chrome-plated tubular framed stool with adjustable height
Wooden seat
Adjustable backrest, hand-controlled and foot circular rest
Shall be mobile on 5 stars stand, 5 castors
42 | Bench-top Centrifuge | Each | 18
Tabletop model with swing out rotor head 16 tubes of 15 ml. glass tubes.
Digital speed indicator with 60 min. Count down timer.
Speed range 500-5500 rpm with RCF 2750 with rotor head .accuracy of 100 rpm .
Complete with: Rotor head. Centrifuge glass tubes 50nos.
Lid interlock required, locking lid while motor is running.
Brushless motors.
Closed lid security system during operation.
Microprocessor control with digital display.
Digital count down timer.
Dynamic breaks, Imbalance detector, cut-off in case of uneven load.
Step less speed regulator & safety lid interlock to prevent lid opening during operation.
Mains power 220- 240V ±6%, 50 Hz British Standard 3 Pin Power Plug / Cable
Technical Data
Max.speed:6000 RPM
Max.RCF:5070 'g'
Max Capacity ML.16*15ml
43 | Laryngoscope for adult & child | Each | 28
Safety standard FDA or CE certificate Laryngoscope type Laryngoscope set for adults and children Laryngoscope set with blades for adults and pediatric patients Blades All blades should be of reusable type, no disposable blades are allowed Handle for adult and pediatric use Handle made of either chromium-plated or stainless steel Handle fitting various sizes and types of blades A set of four stainless steel blades, with halogen bulb or LED technology Maclntosh type: Curved Nr 2,( length approx. 110 mm ) Curved Nr 3, (length approx. 135 mm ) Curved Nr 4, (length approx. 155 mm ) Miller type: Straight Nr 1,( length approx. 100 mm ) Supplied with 1 x durable protective plastic box or padded vinyl case 4 x spare halogen bulb (one for each blade) or LED technology Should be provided with batteries Electrical characteristics Usable with standard disposable commercial batteries Warranty/After Sale Service • Two Years or more comprehensive onsite warranty of entire system (Spares and labour) Maintenance: "Preferred less maintenance needed." 2 years free maintenace.or more by the manufacturer's agent Other specification Please specify other specification
44 | Mayo Table Stainless Steel | Each | 24
Trolley for Instruments, fully made from Stainless Steel
Mobile on 4 antistatic castors, 2 of which shall have foot brakes
The unit shall be height adjustable with rebated edges
45 | Medical oven | Each | 25
Input power : 750W interior Dimension (mm): 420W ×350D ×50H Overall Dimension (mm): 710W ×520D ×520H Gross/Net Weight(kg): 48/42 kg Material: Inner: Stainless steel, outer :High quality steel , thickness 1mm- 1.5mm Circulation : Air inlet can be switch , low-noise blower Shelves : Adjustable 2 Shelves Door : Double safety glass window door Electric supply : 120V 60HZ or 220 V 50HZ Temperature range : RT+10℃ to 200℃ Temperature Precision/Stability : 0. 1℃/±1.0℃ Display : LED 4 digits display Controller : Microprocessor PID multi-function controller Timer : 0-9999min
46 | Medical Refrigerator | Each | 34
Vaccine Solar Refrigerator about 100Litres Low voltage 12 or 24 V Low wattage about 50 W Two parts configuration one part for vaccine, the other part to freezing ice packs. With digital control for two parts Supplied with: all attachment for solar system ( Solar panels, Regulator, Batteries,cables .... etc.) Electric adapter
47 | Mercury Sphygmomanometer | Each | 178
Sphygmomanometer, Desk type, Mercury column-Based
Measuring scale shall be made from injected transparent polymethacrylate encases in metallic structure protected with Epoxy paint or baked-enamel
Range of measurement shall be 0 to 300mmHg as standard with at least a precision of +/-3mmHg (static measure). Display of digits shall be easily readable (4mm high minimum)
Protecting tube with appropriate black color to enhance the reading
Mercury tank shall be monobloc and mercury proof. Safety tap shall be compulsory
Cuff shall be made of PVC, and shall ensure a constant and precise decompression, Velcro fixation
The unit shall be recessed in a closable metallic container
CE marked
Supplied with:
. 1 Cuff for Adult
. 1 Cuff for Children
. 1 Container
48 | Micro automatic pipettes-digital | Each | 72
Safety standard FDA or CE certificate
Design
Used for multiple techniques to accurately dispense required volumes of liquid
New model, able for high load & hard work
Heavy duty and high quality
Type Pipette, digital, 10-100ul, single channel
Technical specifications:
Manual air displacement micropipette
A handle with digital setting and read-out of delivered volumes
Built-in tips ejector
Compatible with universal fit pipette tips ranging 10-100ul
Autoclavable, and Resistant to UV sterilization
Sturdy and chemical resistant parts
Warranty – 3 years
with 500 Strips
Other specification provide all items according to the supplier's catalog
49 | Mobile hospital screan with three section | Each | 77
Ward Screen 3 section with wheel chrome plated
Metal construction,painted with elecrostatic powder, three folds, mobile on six twin castors
complete with pvc curtain, Guiding Dimentions: (180*170 cm) at least
50 | Operating Light, mobile, AC/DC complete with Emergency Power | Each | 16
Technical Specification . Portable O.T Light LED with adjustable height . Light source, type LED . Number of lightheads 1 . Number of bulbs or elements Please specify Number of bulbs or elements . Volts of bulbs 12 V or 24 V . Life, hr of bulbs ≥40,000 Hr . Color temperature, K 3,000-5,000 . Color adjustable Required . Color rendering index >90 . Field size, cm Diameter ≥20 . Focal length, cm ≥ 70 cm .Illumination level, maximum lux at 1 m 140.000 -160.000 Lux at 1.0m . Mounting Mobail .CONTROLS Required .Dimmer With Dimmer . Focus Adjustable . Field size Adjustable .Alarms main lides failure indicator . On sterile handle On sterile handle sterilizable. . Rotation 360 . Vertical adjustment range, cm ≥80 .Maximum irradiance at 1 m, W/m2 <1,000 . Heat-to-light ratio, mW/m2.lux ≤3 .Light head material high qualety Light head material . Shadow control Shadow less . Backup Battery >2 hrs. . Sterilizable handle Required Maintenance: "preferred less maintenance needed. 1 year free maintenace Warranty Service manual operation manual spare parts list with code NO. Including maintenance tools." Battery & charger "Shall run on mains and battery . When used on Battery (DC) the decreasing of intensity shall be less than 10% Battery shall be equipped with an automatic charger with uninterruptible connection to mains" Casters Four casters ,two with brakes Power supply 100 to 240 V ±10%, 50 Hz, ( power cable Compatible with the Hospital electric outlet, plug ), Electrical Safety class 1.
51 | Neonate ambo bag | set | 57
Design & qualit :Heavy duty & high quality Type Resuscitator, hand-operated, adult (body weight > 5 kg), set. Materials Latex-free Materials life All parts must be manufactured from high-strength, long-life materials and require No special maintenance or storage conditions Bag Made of transparent silicone rubber capacity: 220 – 280 ml Compressible self-refilling ventilation bag silicone rubber. Non-rebreathing patient valve with pressure limiting valve polycarbonate/polysulfone/silicone Inlet valve with nipple for O2 tubing polycarbonate/polysulfone Oxygen reservoir bag plastic Accessories Size 0 for preterm and low-weight baby, round type, diameter 35–50 mm Size 1 for term baby, round type, diameter 50–65 mm translucent Reservoir bag of approx. 600 ml capacity Carrying case Safety standard FDA approval or CE marking Warranty Minimum of 1 yearOther specification Please specify other specification
52 | Neonate bed for delivery room | Each | 66
Quality: high quality, Safety, Ease of Use.
Heavy duty desgine.
Cot, designed for Baby
Epoxy-Coated main frame constructed from heavy duty section steel tube
Sliding sides with safety system (Space between vertical bars shall not exceed 110mm)
Mobile on 4 castors, Diameter 100mm minimum. With 2 brakes
Bumpers at all 4 corners
dimentions of bed: 90 * 45*100 cm
Supplied with:
Mattress, 5-8cm height covered Made from foam, Polyurethane. High density, Protected with vinyl cover anti-bacterial, anti-static, acid-resistant and waterproof for easy use and care
st/st IV rod, 2 hooks with adjustable height
53 | Operating table | Each | 9
Gyneacology Operation Table
Gyn. operating table for Caesarian & Gyn. surgeries
foam upholstery pads, electrically conductive. Lectrolyte or similar
Elevation of the table top actuated hydraulically with a pedal
Leg plates (Separate) shall be infinitely adjustable
Overall length including head support: at least +/- 1900mm
Trendelenburg / reverse Trendelenburg position: at least - 20° /+ 20°
Supplied with:
All Accessories for Caesarian & Gyn. surgeries
Accessories if any (to put machine/ system into operation)
Standard Technical Condition:
Installation & Commissioning must be done by Supplyer engineer
Operator training After Installation
2 Years comprehensive warranty, from the date of installation and commissioning
Product Information:
Model release date/ latest version
Origin of the product CE Market high quality
Special feature of the product (other than mentioned on list)
54 | Operation Light (Ceiling) | Each | 8
Operating ceiling light AC/DC, complete with Emergency Power
Shadowless Cold light LED lamp, ceiling mounted, double dome.
No of dome: Double Dome
Major Dome: 160,000 Lux at 1 m
Satellite Dome: 90,000 lux at 1 m
Single bulb Technology with switch over to reserve bulb automatically when main bulb fails (operating bulb should be in center).
Colour Temperature should be 4200 K
Colour Rendering Index should be min 93
Light field Diameter
Major Dome- 155 mm to 270mm
Satellite Dome- 130 mm to 205mm
Depth of Illumination (L1 + L2) should be minimum 1350mm
Irradiance / illumination intensity should be min 3,5mW/m²lx.
Working column with homogeneous light distribution should be min 700mm
Switch on intensity should be 71% of max intensity automatically
Average life cycle of illuminant should be min 1,000 hrs.
Parabolic reflector with min 3000 facets
Light should have completely shadow less
Possibility to field upgrade to wall control panel in future
Integration of the surgical light to central OR control systems possible via RS232 port.
Power Supply, 100-220VAC/50 /60 Hz.
Emergency Power Supply shall be ensured by means of Stand-by Batteries. Stand-by automatic charging unit. Controls of the latest equipment being visible in the Operating Room
55 | Oxygen cylinder 40 liter with regulator 02 with flowmeter and humidifier | Each | 145
Oxygen Cylinder 40 Liters Cap ( Bull Nose British standard)
Trolley Regulator O2 with Flowmeter and Humidifier Flow range: From 0.15 L/min to 15 L/min Maximum possible flow: 35 L/mn above range Input pressure: Not less than 3.5 bar +/- 0,7 Suitable for all types of fitting: wall outlet, trolley. Autoclavable, unbreakable Clear vision of bubbles Humidification at 4 L/mn - 20° C: ~3 ml/h Humidification at 12 L/mn - 20° C: ~9 ml/h Capacity: 100 ml (distilled water) Porosity of diffuser: No more than 70 microns Humidifier, Autoclavable, for multipurpose uses in Medical and Surgical Sections. Shall have the capability to be associated with Flowmeters (O2) or Air , O2 Reserve Cylinders and O2 Generators Fixation to flowmeters or any compatible devices shall be easy and fast (1/4 twist shall be acceptable) Supplied With Tubing And Face Mask Safety standard CE marking approval or ISO
56 | Oxygen regulators | Each | 122
Flow range: From 0.15 L/min to 15 L/min
Maximum possible flow: 35 L/mn above range
Input pressure: Not less than 3.5 bar +/- 0,7
Suitable for all types of fitting: wall outlet,
Autoclavable, unbreakable
Clear vision of bubbles
Humidification at 4 L/mn - 20° C: ~3 ml/h
Humidification at 12 L/mn - 20° C: ~9 ml/h
Capacity: 100 ml (distilled water)
Porosity of diffuser: No more than 70 microns
57 | Patient Monitor | Each | 30
Patient monitor system should be of modular type and capable of monitoring adult, pediatric & neonatal.
Monitor should have 17” independent flat panel display.
Touch screen user interface.
Module rack / housing should be independent and shall be able to be placed near to the patient.
Monitor must be capable of simultaneously monitoring the following parameters which should be present as standard: ECG, NIBP, SpO2, invasive pressures (2), temperatures (2)
24 hours trend data should be displayed.
Monitor shall provide the capability to receive and display real-time waveforms, trended data and alarm status from other bedside or telemetry units on the patient monitoring network.
Alarm limit status (ON/OFF) must be indicated on-screen for each parameter and actual parameter alarm settings must be displayed on-screen when alarms are on.
Line Power: 230 VAC Built-in rechargeable battery with operating time of at least ≥60 minutes.
Patient monitor system should be of modular type and capable of monitoring adult, pediatric & neonatal patients.
Position of the displayed waveforms must be user configurable.
Waveform color changing should be user configurable.
Should be supplied with necessary accessories for adult, pediatric and neonatal accessories.
Supplied with accessories and spares
i. ECG / respiration: 5 lead ECG cable and lead wire set per monitor
ii. NIBP: Adult: 2 sizes and Pediatric 2 sizes and neonatal, 1 size per monitor
iii. SPo2 Sensor: Adult sensor with cable, pediatric sensor with cable and neonatal sensor with cable
iv. IBP: Include 10 no’s of disposable pressure transducer with bracket and interface cable per monitor
v. Temperature: Skin and nasopharyngeal probes per monitor.
58 | Patient Transfer Trolly | Each | 53
Patient Trolley, 2 Sections
Adjustable backrest by means of manual rachetting (Both sides)
All around or corners' rubber bumpers
Mobile on 4 orientable castors (150mm dia. minimum. Rubber bands) 2 brakes
Equipped with Polyurethane foam mattress
Supplied with:
I.V.Hanger Rod & Storage Tube For I.V.
Basket And Oxygen Cylinder Holder For 500L.
With Head Set.
Construction & features Constructed from steel frame epoxy powder coated painting Lever operated back raise Bed side rails Sliding aluminum or chrome plated steel Mattress 7cm urethane foam approx. Bumpers Perimeter Trendelenburg position IV POLE Sockets, number ≥4, at least 1 mountable at each corner Crank-operated high-low position Oxygen tank holder Required CASTERS Diameter, cm ≥12.7, 20.3 preferred Brakes ≥2, diagonal MAXIMUM PATIENT WEIGHT, kg ≥200 kg STRETCHER TYPE Adjustable height Dimension: (L x W), cm (210 x 80) or better. Height, cm Height adjustment by hydraulics system (50 – 120) Warranty Minimum of 1 year Other specification Please specify other specification
59 | Phototherapy infant | Each | 36
Wavelength: 400 - 500nm
Irradiance: > 18µW/cm2/nm at 40cm
Effective surface area :500mm x 300mm
Lamps: 4 Blue CFLs and 2 White CFLs.
Height adjustability: 130cm - 180cm from floor
Electrical supply: 220 V, 50/60 Hz,
Power: 108W
Max. Height: 160cm
1 FDA Required certificate 2 CE MARK Required certificate 3 IF THE MACHINE HAVE CE CERTIFICATE PRODUCT ONLY NEEDS TO HAVE MARKET CLEARANCE FOR EITHER: Australia, Canada, EU, Japan, USA, Canda 4 Copy of the above mentioned certificates shall be included in the offer General specification The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl. 1 Designed: Heavy duty designed, new Model. 2 Quality: high quality, Safety, Easy of Use 3 Supplier the supplier is required to be a certified distributor or agent of the company. Technical Specification 1 Irradiance at 400 to 500nm, µw/cm2/nm Specify 2 Angular adjustments Required 3 Bulbs type LED 4 Bulbs number Specify 5 Bulbs expected life >7000 hours 6 Cooling fan Required 7 Hour meter Required 8 Timer Required Estimated life-span of the device: 10 Years Essential requirement: 1 That the bidder has the capability for corrective and preventive maintenance of the unit. 2 That the bidder/supplier has the engineer/s trained and capable for corrective and preventive maintenance for the model bidded. 3 Service engineer should be presently employed by the bidder/supplier or authorized by the manufacturer. 4 To submit a spare part pricelist valid for 10 years starting date after the warranty completion 5 That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model. 6 That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract. 7 Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment. 8 Technical support from the manufacturer incase the agent or distributor doesn’t response when needed. 9 "Supplier must include full installation, acceptance testing, safety checks and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the responsibility of the supplier to provide. The installation engineer/technician as well as the end-user should be trained and authorized by manufacturer to do the installation." 10 Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number. 11 Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion. Training Service Training for one MWC Bio-Engineers Maintenance: 1 Warranty 2 years free maintenance by Bio medical Engineer of manufacturer's agent and another years by MOH Engineer Power supply 1 240 V ~ ±10%, 50/60 Hz Single phase Electrical Safety class 1 2 Built-in protections against over-voltage and over-current line conditions. Other specification Please specify other specification
60 | Portable Ultrasonic Nebulizer | Each | 45
Ultrasonic Nebulizer Type (Used For Hospitals)
Operating Frequency : ≥ 1.5 MHz ±%5 .
Particle Size : ≥ 0,5 μm .
Average Nebulization Rate : ≥ 5 ml/min ±%5 .
Aerosol Tank : ≥ 250 ml .
Heater Max. : 35 - 38 ºC .
Operating Temperature Range : 10 ºC ~ 38 ºC .
Operating Type : Continuous or Timer.
Standard Accessories
filters, tubings, face mask, etc.
Any needed accessories should be included to ensure full operation.
Power supply OWER , VAC : 100-240, 50/60 Hz
Warranty/After Sale Service One Year or more comprehensive onsite warranty of entire system (Spares and labour)
Essential requirement:
• The model should be FDA approved and/ or CE marked
• The supplier must ensure the availability of expertise service and maintenance.
• Uninterrupted availability of spare parts and repair of next ten years must be assured.
• On site Training should be provided.
• Mention the number (with addresses and phone numbers) of installations of quoted units in Yemen
Other specification provide all items according to the supplier's catalog
60 | Portable Ultrasonic Nebulizer | Each | 55
Ultrasonic Nebulizer Type (Used For Hospitals) Operating Frequency : ≥ 1.5 MHz ±%5 . Particle Size : ≥ 0,5 μm . Average Nebulization Rate : ≥ 5 ml/min ±%5 . Aerosol Tank : ≥ 250 ml . Heater Max. : 35 - 38 ºC . Operating Temperature Range : 10 ºC ~ 38 ºC . Operating Type : Continuous or Timer. Standard Accessories filters, tubings, face mask, etc. Any needed accessories should be included to ensure full operation. Power supply.VAC : 100-240, 50/60 Hz Warranty/After Sale Service One Year or more comprehensive onsite warranty of entire system (Spares and labour) Essential requirement: • The model should be FDA approved and/ or CE marked • The supplier must ensure the availability of expertise service and maintenance. • Uninterrupted availability of spare parts and repair of next ten years must be assured. • On site Training should be provided. • Mention the number (with addresses and phone numbers) of installations of quoted units in Yemen Other specification provide all items according to the supplier's catalog
61 | Resuscitation set (adult) | Each | 56
Resuscitation Adult including:
(Bag resuscitator with: patient valve, relief valve, oxygen reservoir and 3 different sizes of face masks)
Supplied with Spare Valve
62 | Resuscitation set (pediatric) | Each | 27
Resuscitation pediatric including:
(Bag resuscitator with: patient valve, relief valve, oxygen reservoir Tube and face mask)
Supplied with Spare Valve
63 | Resuscitation set (Infant ) | Each | 36
Resuscitation Infant including:
(Bag resuscitator with: patient valve, relief valve, oxygen reservoir Tube and face mask)
Supplied with Spare Valve
64 | Scale, Adult, with Height Measurement | Each | 45
It shall have a solid base made of steel with tempered spindles of 6 mm and reinforced bearings in V shape. Chromed equilibrating arm with easy reading graduation indications. Steel platform with non slipping ABS cover. the pillar shall be separated from
Power Source:Battery Function:Weight Measuring Display Type:Digital Display Material:Good material with anti-skid surface Division:100g Capacity:180kg Shape:Rectangle Scale Type:Digital
65 | Semi Automatic Biochemistry Analyzer | Each | 24
Safety standard FDA or CE with treding sales in Europe Union, USA, Australia, Canda & Japan Design New model, able for high load & hard work Heavy duty and high quality Type Semi-Automatic single-beam filter photometer ASSAYS Please Specify Analysis Mode Absorbance, Kinetic, Fixed time, End point or better Control Microprocessor controlled programmable, semi auto analyzer to perform routine biochemistry tests with Multistranded curve calibration & memorization etc. Light Source Tungsten/ halogen or higher grade with one additional bulb. Wavelength --Filters Photometer system with at least 7-Filters : 340 - 670 nm, One free position for optional filter Accuracy : 2nm Photometer Range 0.001 to 2.300Abs Quality Control Required Flow Cell / Cuvettes Supports Flow cell as well as cuvettes. Aspiration Mode Automatic by Peristaltic Pump. Biochemistry tests on whole blood, serum, plasma, or urine samples. Display Color LCD/LED (Preferred Touchscreen) Printer Inbuilt printer or External Printer to be supplied Storge Please Specify Equipped with : "500 ml of reagents for urea, S. creatine, S. bilirubin, sugar, cholesterol, and Quality control 5ml one each for normal, abnormal." Two additional filters and the rate for additional two filers must be quoted separately POWER SUPPLY : LINE POWER, VAC : 100-240, 50/60 Hz Installation Installation & Commissioning must be done by manufacturer engineer Training User /Nurses training, by Specialist from the Supplier. Warranty/After Sale Service Two Years or more comprehensive onsite warranty of entire system (Spares and labour) Accessories Suitable UPS 1 x Set of spare fuses, if applicable 1 x Plastic protective dustcover Instructions for assembly, Service manual operation manual {Hardcopy & Softcopy} in English Spare parts list with code NO Electrical cable Essential requirement: • The model should be FDA approved and/ or CE • That the equipment is a brand new unit and not a discontinued model or a demo model & not refurbished model. • Spare parts list with code NO • The supplier must ensure the availability of expertise service and maintenance. • The supplier hereby commits to the availability of spare parts & consumables for the indicated life-span. •That the bidder has the capability for corrective and preventive maintenance of the unit. •Technical support from the manufacturer incase the agent or distributor doesn’t response when needed. • Quick guide card intended to describe the basic operations and routine maintenance in practical applications for the equipment. Maintenance: "Preferred less maintenance needed. " 2 years free maintenace.or more by the manufacturer's agent and after that by MOH BIO medical Engineer. Other specification Please specify other specification
66 | Steam Sterilzer Autoclave | Each | 48
Safety standard: CE, or ISO approval Type : vertical, gase and electric working. Capacity: 18-25 Liters safety: safety valve, door gasket, index pressure. steel type: Stainless steel 316 chamber safety: steam pressure regulator. Water level switch etc. Spare parts: Spare parts for 1 year operation such as heating elements, safety valves,door gasket, Fuse. power suplay: Electrical 220 Volts 50Hz, 1Phase and flame gas. Other specification provide all items according to the supplier's catalog
67 | Suction pump mobile |  | 37
Safety standard FDA approval or CE marking Type Electric , silent , mobile , general purpose operated for operation eoom vane pump or piston, High quality Motor : - Flow rate L/ Min. Not less than approx. 30 liter / min. - Vacuum , Hg. Not less than 650 mm. Vacuum regulate: Not less than 650 mm. Bottles "One or more collecting jars (1.5-2) liters fitted with sterilizable rubber bung & fluid valve. Change-over block from jar to other" Tube Suction tube between operation field and suction machine) reusable length 3m Material silicone, Autoclavable Castors The apparatus should be mounted on four anti-static tired rubber castors two with brakes Filters The apparatus should be available with bacterial filters 2 Spare bottles, 10 spare bacterial filters and 1 tube Must be included in the price Warranty Minimum of 1 year Power supply 100 to 240 V ~ ±10%, 50 Hz ( power cable Compatible with the Hospital electric outlet, plug ), Electrical Safety class 1,with indicators for power Other specification Please specify other specification
68 | Suture Instrument and Dressing Set | Set | 48
| Tray Instrument with Lid 225X 125 X 50 mm Deep | 1
| Forceps artery Kocher COF BJ.S/S 180mm. | 6
| Forceps Dissecting Bonney's 1X 2 Teeth S/S 150 mm | 1
| Forceps Dissecting Plain serrated jaws S/S 150 mm | 1
| Forceps Dressing Bryants serrated jaws S/S 125mm | 2
| Forceps Towel Holding cross Action S/S 100mm | 4
| Forceps Sponge Holding Rampley BJ S/S 240 mm | 2
| Handle size 3 for Scalple blades size 10-15 | 1
| Needle Holder Kilner's 165 mm | 1
| Scissors Mayo Straight SJ. S/S 200mm. | 1
| Scissors Iris straight, Sharp points 115 mm | 1
| Sterilizing and storing case St. St. with filters 230x130x50 mm | 1
69 | Syring Pump with holder | Each | 29
Manufacturer: Please specify manufacturer and country of origin. Model number: Please specify model number of the offered equipment. Country of Origin for the offered model: Please specify.4 Safety standard FDA Required certificate CE MARK Required certificate IF THE MACHINE HAVE CE CERTIFICATE PRODUCT ONLY NEEDS TO HAVE MARKET CLEARANCE FOR EITHER: Australia, Canada, EU, Japan, USA, Canda Copy of the above mentioned certificates shall be included in the offer Genaral specification The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofnessl. Designed: Heavy duty designed, new Model. Quality: high quality, Safety, Easy of Use Supplier the supplier is required to be a certified distributor or agent of the company. Performance data: . Flow rate setting: 0.1-999 ml/hr or better .Volume (VTBI) setting 0.1-999 ml or better .Drug dose rate calculation function .Automatic and manual bolus administration that is adjustable in rate and volume . KVO (keep vein open) function with rate: ≤ 5 ml/hr or better . Prime (purge) function . Accuracy ± 2% (or better) of displayed flow rate and volume .Adjustable limit for occlusion pressure.
1. Volume of syringes accepted range from 5 ml up to 50 ml (or better) 2. Open system, with a wide selection of accepted and compatible syringe types from several manufacturers (major suppliers) while maintaining accuracy 3. Pump Mechanism: Syringe driver 4, Control Microprocessor controlled unit with start-up self-test capability with an error code system 5. Medication database: Pre-programmed drug protocols 6, Display Clear and easy to read LCD display of the following: 7. Flow rate 8. Volume infused and volume to be infused 9.Time remaining 10. AC or Battery operation (with Battery capacity ) 11. Bolus delivery 12. Medication name and settings (when drug library is being used) 13. Alarms and error codes 14. Easy to clean, soft touch keys control panel with visible and clearly identified controls 15. Battery support for 4 hours and more and mains AC power supply indicators 16. Built in maintenance free rechargeable battery that is easily accessible with capacity not less than 5 hours operation at 10 ml/h flow rate when fully charged. 17. Alarm type Visual and audible alarms with adjustable alarm volume control. 18. Alarm founection The following alarms should be included (at least but not limited to): 19. High pressure occlusion 20. Syringe unlocked 21. Malfunction alarms 22. Low battery alarm 23. Empty and near empty syringe 24. Syringe size Automatic syringe size identification Included 25. Safety standards The unit shall meet at least the following standards: 26. Defibrillation proof type CF 27. IEC 60601-2-24: safety and essential performance of Syring Pump 28. IPX 2 (or higher): Fluid ingress protection 15 Maintenance: 1 "Preferred less maintenance needed. Warranty 2 years free maintenaceor more" 2 Service manual operation manual {Hardcopy & Softcopy} 3 application software and interface connection Included.
16 Training Service Training for one MWC Bio-Engineer and user shall be provided within the first year of warranty 17 Power supply 1. 110-240V, 50Hz single-phase electrical source.
70 | Tray instrument with lid st.st | PCS | 114
Safety standard CE marking approval or ISO Type Tray,instr,ss,310x195x60mm,w/cover Specifications "Tray, instruments. S/S tray with cover. Rectangular with rounded corners, smooth surface. Length: approx. 300-320.mm. Width: approx. 185-205.mm. Height: approx. 55-65.mm. Thickness: approx. 0.75-0.85.mm. Material: Austenitic stainless steel. Austenitic stainless steel composition: 18-20% chromium, 8-10% nickel. Reusable" Materials: "High resistance to corrosion (tropical environment) Support column, ring structure and bowl: Austenitic stainless steel 18/10 Other specification provide all items according to the supplier's catalog
71 | Ultrasound | Each | 29
Safety standard FDA or CE certificate Type Portable ultrasound unit, digital channels technology with convex and Vaginal probes, B/W printer, trolley and accessories Imaging focusing: Adjustable focal depth, either manually or automatically synchronized to the selected scanning depth. Zooming capability with automated image optimization. Depth range selection: Capable of multiple depth range selection, synchronized with automatic focal zone selection. Field of view: Imaging depth at least 20 cm. Image orientation: Capable of lateral and vertical inversion (in B-mode). Image modes at least: 2-D imaging M-mode B/M mode Dual 2-D/colour image mode with cine loop Software applications that include at least: Obstetrics/gynaecological measurements and calculations, including a) gestational sac mean diameter b) femur length, c) crown-rump length, d) biparietal diameter and abdominal circumference, e) enabling estimation of gestational age f) foetal weight (preferable) Small parts/soft tissues Easy selection of callipers Measurements capabilities a) distance b) area c) circumference by ellipse and trace method Capability to be upgraded with additional software applications. Screen annotations capture patient data, date and time, scanning protocols, probes. Text annotations and body markers and image orientation indicator. Transducer ports: At least three active transducer ports permanently available Capability of electronic switch between probes. Monitor and display Screen monitor: 1) high-definition (HD) digital black and white display (LCD) monitor 2) At least 12 inches 3) With reflection filter or adequate alternative. Screen monitor protection available User-friendly control panel: Easy to use, logical and orderly control panel for quick and easy location of most common functions Back lighting or highlighting of application knobs/buttons/controls. Communication and storage Data communication Storage transfer interface a) USB minimum b) high-definition multimedia interface (HDMI) preferable DICOM 3.0 conformity Digital image storage capacity: At least 200 GB of cine memory for static images and dynamic clips Cine loop: Freeze and cine-loop functions Image grey scale: At least 256 shades of grey and video output of 625 lines/frame 150 dB full time dynamic range Capability for database of patient images and information Accessories Transducers (multi-frequency): a) Convex probe 2.5 - 5 MHz multi frequency (Priced probe) b) Vaginal probes Trolley, with gel bottle holders, drawer, or dedicated space for accessories, as well as a place for scanner positioning and easy orientation. Printer (W/B) for medical ultrasound applications (must be specified manufacturer and model) Online UPS (must be specified manufacturer , model and details) 2x gel bottle, 1000ml Electrical characteristics A 220-240V, 50Hz single-phase electrical source Compliant with IEC60601 or equivalent Built-in protections against over-voltage and over-current line conditions. Maintenance: "Preferred less maintenance needed. " 2 years free maintenace.or more by the manufacturer's agent Other specification provide all items according to the supplier's catalog Essential requirement: The product to be supplied shall be new, unused and conform to the standards as specified in technical specifications. That the terms and conditions stated in the contract shall be honored by the manufacturer in the event that a change of exclusive distributorship will occur during the duration of the said contract. Transport, handle and store all products and materials in accordance with the manufacturer’s recommendations and in a manner that prevents damage or deterioration or excessive distortion. The equipment proposed shall be of highest quality and produced by well known manufacturers. The equipment shall carry the name and quality label of the manufacturer and fulfill the standards in force. Technical offers must include brochures, data sheets and technical complete technical specifications. Upon delivery of the system, the supplier should show official documents showing the country of origin and the date of manufacturing the system as stated by the manufacturer, associated with the serial number. All equipment is to be brand new (previously used for demonstration or loan). Must not include previously used and/or refurbished components, latest in design and technology. The supplier must include complete installation, acceptance testing, calibration, safety checks, user training and commissioning of the system for clinical use as per the manufacturer’s recommended procedures. All the equipment and tools to be used in these steps shall be the supplier's responsibility. The installation engineer/technician and the end user should be trained and authorized by the manufacturer to do the installation. The supplier should deliver with the machines all operator manuals, service manuals, engineering schematics, and all documents, software media relevant to the machines and a spare part list with cod numper To submit a spare part valid for 10 years starting date after the warranty completion The supplier should have medical equipment work shoppe Technical support from the manufacturer incase the agent or distributor doesn’t response when needed. Suppliers should have spare parts stores for medical equipment.
72 | Vaginal Speculum Set | Each | 83
Safety standard CE MARK (MDD) / FDA CLEARANCE : Required Technical Specifications: "Material: Stainless Steel Reusable or Disposable: Reusable Sterile or Non-Sterile: Non-Sterile Latex-Free: Yes Rust Prevention Procedure: Passivation Ultrasonic Cleaned: Yes Matt-Polished: Yes Lubricate: Yes Usage: Left Hand or Right Hand Tests Performed: Boil Test, Performance Test, Shape Test QC Passed: Yes Quality: high quality , Safety, Ease of Use." Each set contains the following: Supplied with: 3 diferinte size Other specification Please specify other specification
73 | Ventilator (adult, pediatric, and neonatal) |  | 8
Design: Compact , robust design on mobile, Heavy duty designed, new Model. Safety: The unit shall be safe to use both for -the operator and the patient. Class I, Protection against leakage current, protection against electric shocks: class I, Waterproofness, FDA or TUF approval. Ventilation modes: (Adult, pediatric, and neonatal ventilation) -Adaptive (APVcmv /APVsimv) Adaptive pressure ventilation + CMV). -Pressure (CMV/ P-SIMV /SPONT /DuoPAP/ APRV). -Noninvasive (NIV/ NIV-ST /nCPAP-PS 1). -Oxygen therapy (High flow oxygen therapy). -Back up mode for apnea -Full alarm system for all ventilator settings and monitored values. -Pulse oximetry measurement (SpO2 sensor) -Integrated pneumatic nebulizer -Storage and display of up to 900 events with date and time stamp -Automatic maneuver for static compliance assessment and lung recruitment including transpulmonary pressure -Automatic leakage compensation. Ventilation parameters: -Tidal volume (200 – 2000 ML (Adult patient). a. 50 to 300 ML (Pediatric PC mode). -Respiratory rate (5 – 100 BPH) -Pressure (0 – 100 cm H2O) -Inspiratory Peak Flow(4 – 100 1/min) -Minute volume (ASV)% (-25% to 350%) -Oxygen Concentration (21 –100 %) -Inspiratory pause ( 0.1 – 5.5 sec) -Expiratory trigger sensitivity (ETS) (5% to 70% of inspiratory peak flow) -PEEP/CPAP Display: -14“ or higher size XGA, TFT color, LCD touchscreen -Display of more than 50 monitoring parameters -Up to 8 waveforms at a glance based on: Paw, Pes, Paux, Volume, Flow, FetCO2, PetCO2, SpO2. Service Life: Under normal usage, the unit shall have a service life of not less than five years (excluding routinely replaced items such as a battery, accessories, etc.). Warranty: Parts and labor shall be furnished under warranty for ten full years. Maintenance: To need less maintenance,1 years free maintenance. Service manual, operation manual {book & CD} - spare parts list with code NO - software CD and interface cable Included. Quality: Top quality machine, Safety, Ease of Use, continuously operating. Accessories: -SPO2 sensor. -Patient circuit(Adult reusable) complete set. -Patient circuit (Paediatric reusable) complete set . - Nebulizer Ultrasonic Complete set. - Humidifier. -O2 Pressure Regulator with hose . - 5 meters (conversion kit) -Hose for O2 connection with connector 5 mts. - Hose for compressed air with connector 5 mts. -Test lung. - All other needed accessories. Power supply: 100 to 240 V ~ ±10%, 50/60 Hz, automatic range selection ( power cable Compatible with the Hospital electric outlet, plug ), Electrical Safety class 1, Rechargeable battery with built in charger with indicators for charging and low battery indicator & easy accessible from exterior.
74 | Wheel chair | PCS | 38
Hospital use type for adult Patient Transfer
Tubular Framing in Stainless Steel.
Heavy duty, foldable.
2 small orientable castors and 2 large castors, Rubber bands
with back case and hand break, Foot action. Directly accessible by staff
Collapsible armrests
Seat and back made of foam upholstery. Vinyl.
Removable legrest
75 | Microplate Shaker | Each | 12
Safety standard CE marking certificate Type Bowl, round, stainless steel Specifications "Container, round. Smooth surface. Internal diameter: approx. 100 - 120 mm. Height: approx. 60-75 mm. Capacity: approx. 350to 450 ml. Thickness: approx. 0.8 mm. Material: Austenitic stainless steel. Austenitic stainless steel composition: 18 - 20% chromium, 8 - 10% nickel. Reusable." Materials: "High resistance to corrosion (tropical environment) Support column, ring structure and bowl: Austenitic stainless steel 18/10 Other specification provide all items according to the supplier's catalog
76 | Water bath. | Each | 11
Microprocessor controlled equipment. Electronic control.
stainless steel, Digital Display.
Circulating
Perforated shelf with cover
Capacity 14-20 liters, temperature 250C-850C.approx.
Temperature : ambient+ 5 -85°C.approx.
Sensitivity, ± 0.2o C
Uniformity ± 0.1o C
Complete with cover.
Power supply 100 – 240 v 50 Hz
77 | Pulse oximeter | Each | 44
A portable and light weight NIBP/Spo2 monitor with
plethysmographic waveform display.
Features include:
 High resolution graphic LCD with backlit with a
plethysmographic display
Display of plethysmograph with numerics of %O2, Pulse Rate,
Systolic, Diastolic & Mean BP
 Easy menu driven operation
 Programmable audiovisual alarms
Trend – 24 hour SPO2 / 100 NIBP readings
78 | Cardio tachograph CTG | Each | 21
Continuous monitoring of FHR, Toco& Foetal movements with trace on LCD
FHR display with beat to beat monitoring with auto correlation technique using EMEBREDDED Technology for computation of the FHR in B.P.M.
Graph on LCD & Dot –matrix printer using plain paper
Annotation of Hospital & patient details
Memory of 8 – 10 hours /20 patients
Automatic TOCO zero
Tachycardia & bradycardia alarms
Multi crystal ultrasound Transducer & digitally controlled feather touch keys
Technical Data-
FHR Probe : Large 50 mm diameter transducer
Operating Frequency: 2MHz
Fetal Heart Rate: Measurements Range: 60-220 BPM
Uterine Contraction: External : Toco Transducer
Event Marker: Hand held, patient operated
Dimensions: 380 (W) Χ 100 (h) Χ 330 (d) mm
Printer: Type : 80 Column dot matrix printer
with 10 packs of paper
Trolley
Voltage: 230 V A.C+/- 10% 50 Hz
79 | Oxygen concentrator with double outlet | Each | 6
Size: : compact and portable
Weight : upto 25 kgs
Sound level : < 50 dba
Power consumption : 400 watts avg
Safety alarms : power failure
process failure
high and low pressure
low oxygen purity
Oxygen purity : 0-2 lpm-94%, 2-4 lpm-93%, 5 lpm- 90%
Outlet pressure : 7 psi
Oxygen output : variable – 0 to 5 lpm (1/2 litre increment)
Filters : cabinet, compressor intake, and bacteria
80 | Cheatle Forceps With Jar | pcs | 28
Forceps, sterilizer, Cheatle, 27 cm, curved
Martensitic stainless steel (quenched, magnetic steel) composition: From 0.16% to 0.25% carbon and From 12 to 14% chromium.
Length: 270 approx. mm +/- 5 mm