United Nations Population Fund

Address 131, 31 August 1989 street, Chisinau,

Republic of Moldova

Website: www.unfpa.org

07 November 2024

INVITATION TO BID

ITB No. UNFPA/MDA/24/004

SUPPLY AND INSTALLATION OF MEDICAL OXYGEN GENERATOR

IN REPUBLIC OF MOLDOVA

Dear Sir/Madam,

- The United Nations Population Fund (UNFPA), an international development agency, invites sealed bids for the supply of Medical oxygen generator (for ventilatory therapy (ICU ventilators, Anesthesia machine) for refugee response in the Republic of Moldova

- Bidding shall be conducted through ONE envelope. The technical bid containing the technical specifications and the financial bid containing price information shall be submitted together.

- The Bidder shall be required to quote for all items.

- To enable you to submit a bid, please read the following attached documents carefully:

| Section I: | Instructions to Bidders |
| --- | --- |
| Section II: | Technical Specifications and Schedule of Requirements |
| Section III: | UNFPA General Conditions of Contract |
| Section IV: | UNFPA Special Conditions for Contracts |
| Section V: | Bidding Forms |

- The bid shall reach UNFPA’s reception or the email inbox of tender.mda@unfpa.org no later than 05 December 2024, at 22:00 Chisinau GMT+2.

- The bid shall be opened on 09 December 2024, at 10:00am at 131, 31 August 1989 street, Chisinau, Republic of Moldova, MD-2012, bidders or their authorized representatives may attend the bid opening.  Kindly confirm by e-mail by 5 December 2024 whether your company shall be represented at the bid opening.

- Bids received after the stipulated date and time shall not be accepted under any circumstances. Bids delivered through courier and posted later than the due date shall not be registered and shall be returned unopened or shall be shredded. Bids submitted to any other email address than tender.mda@unfpa.org shall be rejected.

- Bidders shall acknowledge receipt of this Invitation to Bid according to the Bid Confirmation Form,  Section V, 1 of this solicitation document by email to Rajeev Rajasekhar, rajasekhar@unfpa.org  no later than 25 November 2024 and to indicate whether or not a bid shall be submitted.  The acknowledgement shall provide company name, telephone number, fax number and the name of a contact person. If you are declining to bid, please confirm this via e-mail to UNFPA and please state the reasons for UNFPA to improve its effectiveness in future invitations.

- Any questions relating to the attached documents shall be addressed in writing to the following UNFPA personnel:

- Rajeev Rajasekhar, rajasekhar@unfpa.org for questions relating to the bidding exercise.

Do not submit your bid to these contacts, or your bid will be disqualified.

- This letter is not to be construed in any way as an offer to contract with your firm.

- UNFPA strongly encourages all Bidders to register on the United Nations Global Marketplace (http://www.ungm.org). The UNGM is the procurement portal of the United Nations system. By registering on UNGM, vendors become part of the database that UN buyers use when searching for suppliers. Vendors can also access all UN tenders online and, by subscribing to the Bid Tender Service, vendors can be automatically notified via e-mail of all UN business opportunities that match the products and services for which they have registered. Instructions on how to subscribe to the Tender Alert Service can be found in the UNGM Interactive Guide for Suppliers http://www.ungm.org/Publications/UserManuals/Suppliers/UserManual_Supplier.pdf .

Yours sincerely,

Rajeev Rajasekhar

UNFPA

Republic of Moldova

UNITED NATIONS POPULATION FUND

INVITATION TO BID

ITB NO.: UNFPA/MDA/24/004

Bid document for the manufacture and/or supply of products and related services

07 October 2024

Table of Contents

SECTION I: Instructions to Bidders	6

A. Introduction	7

1. Scope	7

2. Eligible Bidders	7

Fraud and Corruption	7

B. Solicitation Documents	7

4 UNFPA Solicitation document	7

5 Clarifications of solicitation document	8

6 Amendments to UNFPA bid solicitation document	8

C. Preparation of Bids	8

7 Documents to be submitted with the bid	8

8 Bid Currency and Prices	9

9 Validity of Bid	9

D. Submission of Bids and Bid Opening	10

10 Partial Bids	10

11 Alternative Bids	10

12 Bids	10

13 Sealing and Marking of Bids (hard copies)	10

14 Electronic Submissions	11

15 Bid Submission Deadline/Late Bids	11

16 Storage of Bids	12

17 Bid Opening	12

E. Evaluation and Comparison of Bids	12

18. Confidentiality	12

19. Clarification of Bids	13

20. Responsiveness of bids	13

21. Nonconformities, Errors, and Omissions	13

22. Preliminary examination of Bids	13

23. Examination of Terms and Conditions and Technical Evaluation	14

24. Conversion to Single Currency	14

25. Evaluation of Bids	14

26. Comparison of Price Bids	14

27. Post-qualification of the Bidder	14

28. UNFPA’s Right to Accept Any Bid and to Reject Any or All Bids	15

29. UNFPA’s Right to Annul a Bidding Process	16

F. Award of Contract	16

30. Award Criteria	16

31. Right to Vary Requirements at Time of Award	16

32. Signing of the contract	16

33. Publication of Contract Award	16

SECTION II: Technical Specifications and Schedule of Requirements	17

SECTION III: UNFPA General Conditions of Contract	30

SECTION IV: UNFPA Special Conditions for Contracts	31

SECTION V: Bidding Forms	31

1. Bid Confirmation Form	34

2. Bid Submission Form	35

3. Bidders Identification Form	36

4. Product Item Overview Form	38

5. Price Schedule Form	39

# SECTION I: Instructions to Bidders

# Introduction

## Scope

  - The goods and related services to be procured are

- Medical Oxygen Generator and Installation of the unit along with Training of Bio engineers at the hospital. The selected supplier will be responsible to connect the unit to the existing oxygen network, including all the connectors and piping necessary for the   installation and commissioning of the unit.

- The supplier shall provide in the bidding document the details about the site preparation for the installation of the Oxygen Generator, including the drawing.

- The supplier shall exclude the piping cost for the entire hospital, which is already in place, only connectors necessary for connecting to the existing network shall be included.

- The spare parts for next two years shall be included.

## Eligible Bidders

  - All Bidders found to have a conflict of interest shall be disqualified. Bidders may be considered to have a conflict of interest if they are or have been associated in the past, with a firm or any of its affiliates that have been engaged by UNFPA to provide consulting services under these bidding documents.

  - Bidders shall not be eligible to submit a bid if at the time of bid submission:

- The Bidder is listed as suspended on United Nations Global Marketplace (http://www.ungm.org) as a result of having committed fraudulent activities,

- The Bidder’s name is mentioned in the UN 1267 list issued by the Security Council resolution 1267 that establishes a sanctions regime to cover individuals and entities associated with Al-Qaida and/or the Taliban;

- The Bidder is debarred by the World Bank Group.

## Fraud and Corruption

  - UNFPA’s policy regarding fraud and corruption is available at http://www.unfpa.org/about-procurement#FraudCorruption  and applies fully to this Invitation to Bid. The submission of any offer implies that the Bidder is aware of this policy.

# Solicitation Documents

## UNFPA Solicitation document

  - Bidders are expected to examine all instructions, forms, specifications, terms and conditions contained within this UNFPA solicitation document. Failure to comply with these documents shall be at the Bidder’s risk and may affect the evaluation of the bids, or may result in the rejection of the bid.

  - Bidders are cautioned to read the specifications carefully (see Section II Technical Specifications and Schedule of Requirements), as there may be special requirements. The technical specifications presented herein are not to be construed as defining a particular manufacturer’s product. Bidders are encouraged to advise UNFPA if they disagree.

  - The specifications are the minimum requirements for the products and related services. Products and services offered must meet or exceed all requirements herein. The products shall conform in strength, quality and workmanship to the accepted standards of the relevant industry. Modifications of or additions to basic standard products of less size or capability to meet these requirements will not be acceptable.

## Clarifications of solicitation document

  - A prospective Bidder requiring any clarification on the bid solicitation documents may notify UNFPA in writing within 10 days from the date of issue of the bid. UNFPA shall respond in writing to any request for clarification received and circulate its response (including an explanation of the query but without identifying the source of enquiry) to all prospective Bidders who have received the bid solicitation documents. A copy of UNFPA’s answer shall also be posted on the UN Global Marketplace, http://www.ungm.org/

## Amendments to UNFPA bid solicitation document

  - At any time prior to the deadline for submission of bids, UNFPA may for any reason, whether at its own initiative or in response to a clarification requested by a prospective Bidder, modify the bidding documents by amendment.

  - All prospective Bidders that have received the bidding documents shall be notified in writing of all the amendments to the bidding documents. In order to give prospective Bidders reasonable time to take the amendments into account in preparing their bids UNFPA may, at its discretion, extend the deadline for the submission of bids.

# Preparation of Bids

## Documents to be submitted with the bid

  - Documents Establishing the Eligibility of the Bidder

To establish their eligibility, Bidders shall:

- Complete the Bid Submission Form, Section V, 2.

- Complete Bidders Identification Form, Section V, 3.

  - Documents Establishing the Qualifications of the Bidder

To establish its qualifications, the Bidder shall submit to UNFPA’s satisfaction the following documents:

- Evidence that the Bidder is established as a company and legally incorporated in the country where it resides; e.g. through provision of certification of incorporation or other documentary evidence (this is not required for companies already registered in national, regional or international Stock Exchanges);

- Copy of valid manufacturing license from the country of manufacturing and/or a copy of company registration in the country of operation demonstrating that is duly authorized to supply these goods to the country of destination.

- If a Bidder is not operating in the Moldova or Romania, they must be represented by an Agent within these countries who can fulfill the supplier’s obligations for maintenance, training, repairs, and spare parts as outlined in Section After-sales service, Technical Specifications, and Schedule of Requirements.

- Written confirmation from the Bidder that the Bidder is neither suspended by the United Nations system nor debarred by the World Bank group.

- The availability of spare parts and after-sales services for the Oxygen Generator.

- Post qualification documentation outlined in Instructions to Bidders, Sub-Clause 27

Failure to furnish all the information required for submission shall be at the Bidder’s risk as it may then be determined that the bid does not substantially respond to the UNFPA bid document in every respect.  This may result in a rejection of the bid.

  - Documents Establishing the Eligibility and Conformity of the Goods and Related Services

Bidders shall submit:

- Documentary evidence that the goods conform to the Technical Specifications and standards specified in Section II Technical Specifications and Schedule of Requirements.

- Completed Product Item Overview Form, Section V, 4.

- Product catalogues containing pictures of the product(s)

- Manufacturer’s technical product specifications or datasheets

- Copies of current certifications as stated in the Technical Specifications and Schedule of Requirements Section II

- The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the goods during 2 years following commencement of the use of the goods by UNFPA. Bidders must complete and submit with their bid the Excel table containing the individual item details, as per Form in Section V.4. Bidding Forms.

- Manufacturer’s warranty letter indicating the time of the warranty duration.

- A copy of user, service, spare parts manual and installation, if not provided with the submission of the documents, a commitment letter to deliver such documents will be provided along with the Oxygen Generator.

- Details about the site preparation for the installation of the Oxygen Generator, including the drawings.

- Training details and duration of the training.

## Bid Currency and Prices

  - All prices shall be quoted in US Dollars (USD).

Bidders are requested to quote the following based on INCOTERMS 2010:

- Price of goods FCA Point of departure

- Freight cost CPT, Republic of Moldova, Ministry of Interior Hospital, Chisinau

  - Where installation, commissioning, training or other similar services are required to be performed by the Bidder, the Bidder shall include an itemized list of the prices for those services.

## Validity of Bid

  - The prices of the bid shall be valid for 90 days after the closing date of bid submission as specified by UNFPA. A bid valid for a shorter period shall be rejected by UNFPA on the grounds that it is non-responsive.

  - In exceptional circumstances, UNFPA may solicit the Bidder's consent for an extension of the period of validity under exceptional circumstances. The request and the responses shall be made in writing.

# Submission of Bids and Bid Opening

## Partial Bids

  - Partial bids are not allowed under this tender.

## Alternative Bids

  - Alternative bids will not be accepted. In the event of a supplier submitting more than one bid, the following shall apply:

- All bids marked alternative bids will be rejected and only the base bid will be evaluated.

- All bids will be rejected if no indication is provided as to which bids are alternative bids.

## Bids

  - Bids shall be submitted in one envelope or transmitted in an email to a secure email address designated by UNFPA tender.mda@unfpa.org.

  - Bids shall be prepared in accordance with Section II: Schedule of Requirements and Technical Specifications and shall include the requested documentation as per Instructions to Bidders Clause 7, and in in accordance with the Price Schedule Form in Section V, 5 of the bid forms.

  - Bids shall be signed by the Bidder or a person or persons duly authorized to bind the Bidder to the contract. A bid shall contain no interlineations, erasures, or overwriting except as necessary to correct errors made by the Bidder. In that case such corrections shall be initialled by the person or persons signing the bid.

## Sealing and Marking of Bids (hard copies)

  - When submitting bids in hard copies the Bidder shall prepare one set of sealed bids containing the technical and price components.

  - The envelope shall also indicate the name and address of the Bidder to enable the bid to be returned unopened in case it is declared “late.”

  - If the outer envelope is not sealed and marked as required, UNFPA shall assume no responsibility for the bid’s misplacement or premature opening.

  - The outer envelope must be clearly marked with the following:

UNITED NATIONS POPULATION FUND (UNFPA)

131, 31 August 1989 street,

Chisinau, Republic of Moldova, MD-2012

Invitation to Bid No. UNFPA/MDA/24/004

Attention: Rajeev Rajasekhar – Supply Chain Specialist

ONLY TO BE OPENED BY AUTHORISED UNFPA PERSONNEL

## Electronic Submissions

  - Bids may be submitted electronically. Please note the following guidelines for electronic submissions:

  - Bidders shall make clear reference to the specific bid in the subject field as instructed, otherwise bids may be rejected. Clearly specify the following text in the subject line: ITB No. UNFPA/MDA/24/004, Bidder’s Name.

  - The bid shall be submitted to tender.mda@unfpa.org .  Bids received at the tender.mda@unfpa.org  mailbox are kept undisclosed and shall not be opened before the scheduled opening date. Sending to any other email address will violate confidentiality and invalidate the bid.

  - E-mail submission shall not exceed 10 MB, including the size of the cover email. It is recommended that all the bidding documents are consolidated into as few attachments as possible which shall be in commonly used file formats. If the bid consists of large electronic files, it is recommended to send these files separately before the deadline indicating the order of emails (email 1, email 2, etc.) after the bid reference number and the Bidder’s name in the subject line of each email.

  - It shall be the Bidder’s responsibility to ensure that bids sent by e-mail are received by the deadline. All Bidders shall receive an auto-reply acknowledging the receipt of their email. Bidders shall not receive responses to questions sent to tender.mda@unfpa.org  since it is a secure mailbox.

  - In order to avoid last minute internet congestion it is recommended to send your bid as early as possible before the deadline.

## Bid Submission Deadline/Late Bids

  - Bids must be delivered to the office on or before the date and time specified in the introductory letter of this solicitation document.  If any doubt exists as to the time zone in which the bid should be submitted please refer to www.timeanddate.com/worldclock, or contact the bid focal point.

  - UNFPA may, under special and exceptional circumstances, extend the bid submission deadline and such changes shall be notified in UNGM before the expiration of the original period.

  - Any bid received by UNFPA after the bid submission deadline shall be rejected and returned unopened to the Bidder. UNFPA shall not be legally responsible for bids that arrived late due to the Bidder’s problems with transmission of bid submissions via email and/or with the courier company.

## Storage of Bids

  - Bids received prior to the deadline of submission and the time of opening shall be securely kept unopened until the specified bid opening date stated in the UNFPA’s solicitation document. No responsibility shall be attached to UNFPA for prematurely opening an improperly addressed and/or identified bid.

## Bid Opening

  - UNFPA shall conduct the bid opening in public at the following address, date and time.

Street Address:

UNFPA

131, 31 August 1989 street,

Chisinau, Republic of Moldova, MD-2012

Date:  09 December 2024

Time: 10am, Chisinau +2GMT

  - Bids received electronically by the required deadline will be printed and a copy of the bids will be put in a sealed envelope that will be opened at the time and date specified in the bid document. Only the last received bid will be opened if multiple bids are sent by a same Bidder.

  - The bids shall be opened publicly at the time and place specified in the ITB and an immediate record made thereof.

  - Only those who have submitted bids or their authorized agent or representative may attend the bid opening.

  - The report shall be available for viewing by Bidders for a period of thirty days from the date of the opening. No information that is not included in the bid opening report can be given to Bidders.

  - No bid shall be rejected at bid opening, except for late bids, which shall be returned unopened to the Bidder.

# Evaluation and Comparison of Bids

## Confidentiality

  - Information relating to the examination, evaluation, comparison, and post-qualification of bids, and recommendation of contract award shall not be disclosed to Bidders or any other persons not officially concerned with such process until the contract award is published.

  - Any effort by a Bidder to influence UNFPA in the examination, evaluation, comparison, and post-qualification of the bids or contract award decisions may result in the rejection of its bid.

## Clarification of Bids

  - To assist in the examination, evaluation and comparison of bids, UNFPA may ask Bidders for clarification of their bids. The request for clarification and the response shall be in writing by UNFPA and no change in price or substance of the bid shall be sought, offered or permitted.

## Responsiveness of bids

  - UNFPA’s determination of a bid’s responsiveness is to be based on the contents of the bid itself.

  - A substantially responsive bid is one that conforms to all the terms, conditions, and specifications of the bidding documents without material deviation, reservation, or omission. A material deviation, reservation, or omission is one that:

- affects in any substantial way the scope, quality, or performance of the goods and related services specified in the contract; or

- limits in any substantial way, inconsistent with the bidding documents, UNFPA’s rights or the Bidder’s obligations under the contract; or

- if rectified would unfairly affect the competitive position of other Bidders presenting substantially responsive bids.

## Nonconformities, Errors, and Omissions

  - Provided that a bid is substantially responsive:

- UNFPA may waive any non-conformities or omissions in the bid that do not constitute a material deviation.

- UNFPA may request that the Bidder submit the necessary information or documentation within a reasonable period of time to rectify non material non conformities or omissions in the bid related to documentation requirements.  Such omission shall not be related to any aspect of the price of the bid.  Failure of the Bidder to comply with the request may result in the rejection of its bid.

- UNFPA shall correct arithmetical errors on the following basis:

- If there is a discrepancy between the unit price and the line item total that is obtained by multiplying the unit price by the quantity, the unit price shall prevail and the line item total shall be corrected, unless in the opinion of UNFPA there is an obvious misplacement of the decimal point in the unit price. In that case the line item total as quoted shall govern and the unit price shall be corrected;

- if there is a discrepancy between words and figures, the amount in words shall prevail;

- if there is an error in a total corresponding to the addition or subtraction of subtotals, the subtotals shall prevail and the total shall be corrected; and

## Preliminary examination of Bids

  - UNFPA shall examine the bids to determine whether they are complete, that all documents and technical documentation requested as per Instructions to Bidders Clause 7 have been provided and to determine the completeness of each document submitted. UNFPA will also examine whether any computational errors have been made, whether the documents are properly signed, and whether the bids are generally in order.

## Examination of Terms and Conditions and Technical Evaluation

  - UNFPA shall examine the bid to confirm that it does not contain any material deviations, reservation, or omission related to the conditions and requirements specified in the Section II Technical Specifications and Schedule of Requirements, Section III UNFPA General Conditions of Contract and Section IV UNFPA Special Conditions for Contracts.

  - If after the examination of the terms and conditions and the technical evaluation UNFPA determines that the bid is not substantially responsive in accordance with Instructions to Bidders Clause 21, the bid shall be rejected.

## Conversion to Single Currency

  - To facilitate evaluation and comparison, UNFPA will convert all bid prices expressed in the amounts in various currencies in which the bid prices are payable to US dollars at the official UN exchange rate on the last day for submission of bids.

## Evaluation of Bids

  - UNFPA shall evaluate each bid that has been determined, up to this stage of the evaluation, to be substantially responsive.

## Comparison of Price Bids

  - UNFPA shall compare all substantially responsive bids to determine the lowest priced substantially responsive bid

  - Bid comparison will be made on the total cost, delivered to final destination. UNFPA reserves the right to compare freight prices of Bidders with rates of reputable freight forwarders and to consider such rates for the purpose of bid evaluation. In the event that Bidder’s freight prices are found to be less competitive than the rates offered by freight forwarders, UNFPA may issue a contract on FCA basis to the Vendor instead of CPT, and issue a separate contract for freight to a freight forwarder if deemed in the best financial interest of UNFPA.

## Post-qualification of the Bidder

  - UNFPA shall determine to its satisfaction whether the Bidder with the lowest priced, substantially responsive bid is qualified to perform the contract satisfactorily.

  - The determination shall be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted in the bid.

  - To evaluate a Bid, UNFPA shall consider the following:

- Copy of last year audited company Balance and Financial Statements

- Copy of valid manufacturing license from the country of manufacturing and/or a copy of company registration in the country of operation demonstrating that is duly authorized to supply these goods to the country of destination

- Financial Capability:

  - Liquidity ratio: Current ratio (Current Assets/ Current liabilities) > 1.

- Experience and Technical Capacity:

  - Details of experience and past performance of the Bidder on equipments offered and on those of similar nature within the past five years

  - Registration details of the company Experience to undertake the contract

  - List of similar contracts executed for other clients, including contract details (i.e. Bidders with previous successful supply history to UN system or valid LTA with UN entity).

  - Evidence that  the  Bidder  possesses  experience  in the  geographical  area required by the bid.

  - At least three years of experience in performing similar contracts.

  - Company’s managerial capability:

  - Details of company’s managerial structure.

  - The Bidder shall disclose instances of previous past performance that may have resulted in adverse actions taken against the Bidder and the manufacturers whose products are being offered by the Bidder, in the last five years. Such adverse actions may be treated as unsatisfactory performance history while deciding the award of contract. If no instance of previous past performance has resulted into adverse actions, this must be clearly indicated in the Bidder’s bid.

For non manufacturer Bidders:

- Legally enforceable authorization from the manufacturer assuring full guarantee and warranty obligations as per the tender conditions for the goods offered; and

- The Bidder, as authorized by the manufacturers, has supplied and provided after sales service for similar goods to the extent of at least 20 percent of the quantities indicated in the tender requirements in any one of the last three years, and the goods must be in satisfactory operation.

- Financial  Experience  and  Technical  Capacity  requirements  of  the  manufacturer similar to those mentioned above.

  - Notwithstanding anything stated above, UNFPA reserves the right to assess the Bidder’s capabilities and capacity to execute the contract satisfactorily before deciding on award.

  - Even though the Bidders may meet the above qualifying criteria, they can be subject to disqualification if they have made misleading or false representations in the forms, statements and attachments submitted in proof of the qualification requirements, and/or record of poor performance such as, not properly completing contracts, inordinate delays in completion, litigation history, financial failures, etc.

## UNFPA’s Right to Accept Any Bid and to Reject Any or All Bids

  - A bid that is rejected by UNFPA may not be made responsive by the Bidder by correction of the non-conformity. A responsive bid is defined as one which conforms to all the terms and conditions of the UNFPA’s bid solicitation documents without material deviations. UNFPA shall determine the responsiveness of each bid against the UNFPA solicitation documents.

  - UNFPA reserves the right to reject any bid if a Bidder has previously failed to perform properly or complete on time in accordance with contracts or the Bidder who in UNFPA’s perspective is not in a position to perform the contract.

  - The Bidders waive all rights to appeal against the decision made by UNFPA.

## UNFPA’s Right to Annul a Bidding Process

  - UNFPA reserves the right to annul the bidding process and reject all bids at any time prior to award of purchase order, without thereby incurring any liability to the affected Bidder(s) or any obligation to provide information on the grounds for UNFPA’s action.

# Award of Contract

## Award Criteria

  - In the event of a contract award, UNFPA shall award the to the lowest priced Bidder(s) whose bid has been determined to be substantially responsive with the bidding documents.

  - If required, the Bidder shall permit UNFPA representatives access to their facilities at any reasonable time to inspect the premises that shall be used for the production, testing and packaging of the products. The Bidder shall also provide reasonable assistance to the representatives for such inspection, including copies of any test results or quality control reports as may be necessary. UNFPA may inspect the manufacturing facilities of the lowest evaluated responsive Bidder to assess his capability to successfully perform the contract as per the terms and conditions specified in the ITB.

## Right to Vary Requirements at Time of Award

  - UNFPA reserves the right at the time of award of contract to increase or decrease by up to 20% the quantity of goods specified in this bid without any change in unit price or other terms and conditions.

## Signing of the contract

  - Prior to the expiration of the period of bid validity, UNFPA shall send the successful Bidder the Purchase Orders for the specific procurement requirements (items and quantities listed in this solicitation document), which constitute the notification of award of the specific procurement requirements quantified in this solicitation document. The successful Bidder shall accept the Purchase Order within 7 days of the receipt of the Purchase Order. After receipt of  the Purchase Order, the successful Bidder(s) shall deliver the commodities in accordance with the quantity, quality and delivery schedule outlined in its bid in conjunction with UNFPA terms and conditions.

## Publication of Contract Award

  - UNFPA shall publish the contract award on United Nations Global Marketplace http://www.ungm.org, with the information of the awarded Bidder company name, contract amount or LTA and the date of the contract.

Suppliers perceiving that they have been unjustly treated in connection with the solicitation or award of a contract may lodge a complaint directly with the UNFPA Head of Office at nersesyan@unfpa.org . The UNFPA Head of Office will then make an assessment of the complaint and provide a reply to the supplier within a week. If the supplier is not satisfied with the reply provided by the UNFPA Head of Office, the supplier may escalate the complaint to the Chief, Supply Chain Management Unit at procurement@unfpa.org, who will reply to the supplier within a week and advise the Supplier on further recourse if required.

# SECTION II: Technical Specifications and Schedule of Requirements

- Technical Specifications

| Name | Medical oxygen generator : QTY-1 |  |  |
| --- | --- | --- | --- |
| Type of the system | Medical oxygen generator is a device which generate pure oxygen for medical use, for ventilatory therapy (ICU ventilators, Anesthesia machine) |  |  |
| General requirements | Oxygen output pressure |  | 4 - 7 bar adjustable |
|  | System start up time taken to reach required oxygen purity |  | 5 - 10 min |
|  | Possibility to connect back up power supply |  | Yes |
|  | Possibility to connect at backup oxygen supply |  | Yes |
|  | Operational conditions |  | Temperature -30 - + 45 °C, |
|  | Maximum allowable generated noise |  | inside of the container 80 dB, at one meter from container 70 dB |
|  | Complaints with international standards | CE 93/42 or 2017/745 | Yes, European certificate for medical devices. |
|  |  | Free Sales Certificate | Yes |
|  |  | ISO 13485 | Yes, for medical devices |
|  |  | ISO 9001 | Yes |
|  |  | ISO 7396-1 | Yes, for medical gas pipeline systems |
|  |  | ISO 8573-1, -2, -4 | Yes, standard for gas contamination |
|  |  | ISO 10083 | Yes, for oxygen concentrators |
|  |  | ISO 5011 | Yes, Standard for air cleaning equipment |
|  |  | ISO 21969 | Yes, Standard for flexible connections |
|  |  | CE 2014/68/EU certificate | Yes |
|  |  | ISO 15001 | Yes, Oxygen compatibility materials |
|  |  | IEC 60601-1 | Yes, for electrical safety of medical equipment |
|  |  | ISO 14001 | Yes, for environmental management systems |
| System components |  | Description | Requirement |
| Compressed air system | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Possibility to have an standby or redundand compressor for continuous operation | Yes |
|  |  | Working pressure | 7.5 bar, adjustable |
|  |  | Flow | Minimum 8 m3/min. |
|  |  | Power supply | 380 V / 50 Hz |
|  |  | With oil or oil-free | according to ISO 8573 standard |
|  |  | Electronic oil temperature management system | Yes |
|  |  | Insulation class/engine protection degree | Minimum IP 55 |
|  |  | Electric Motor Efficiency | IE4 according to IEC 60034-30 |
|  |  | Easy start-up via 1:1 direct coupling | Yes |
|  |  | Radial Cooling Fan | Yes |
|  |  | Noise level | According to ISO 2151 |
|  |  | Laminated mineral wool lined panels for sound insulation. | Yes |
|  |  | Operating temperature, with values ranging from | +5°C to +40°C |
|  |  | Rotation control (protected against phase reversal) | Yes |
|  |  | Electro-pneumatic Equivalent / Analog Inlet Valve | Yes |
|  |  | Compressor elements - motor overheating safety, alarm at 105°C, shutdown at 110°C | Yes |
|  |  | Automatic restart system after restarting power in the event of a power failure | Yes |
|  |  | Electronic control system with display and menu at least in Romanian and English | Yes |
|  |  | The control panel with the ability to show the operating hours until the replacement of: Filters, Rumors (including their lubrication), valve grouts, oil. | Yes |
|  |  | Internal self-diagnosis and error display system | Yes |
|  |  | Remote monitoring via computer or smartphone or tablet. Integrated software with the necessary accessories for connection. Monitoring is carried out through TCP, IP or similar protocols. | Yes |
|  |  | Indicators for: pressure and temperature | Yes |
|  |  | Counter for: the total number of hours of shutdown and the total number of hours in the charging phase | Yes |
|  |  | The compressor hoses must be made of lightly non-corrosive metal (without rubber). | Yes |
|  |  | Ball valve at compressor outlet | Yes |
|  |  | Air Dryer |  |
|  |  | Integrated dew point sensor with measure possibility - 20° + 20° C | Yes |
|  |  | Flow | Minimum 13 m3/min |
|  |  | Power supply | 220 V / 50 Hz |
|  |  | Maximum working pressure | 3 - 16 bar |
|  |  | Real pressure Dew Point Temperature generated | 3°-5°, alarm if is greated |
|  |  | Supported air inlet temperature | Minim 60 °C |
| Compressed air filtration system | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Coarse filter with self-draining for large particles |  |
|  |  | Flow: filtration capacity at 7 bar pressure | Minimum 630 m3/hour |
|  |  | Working pressure | Minim 16 bar |
|  |  | Solid particle retention | 5 μm |
|  |  | Clogging indicator gauge | Yes |
|  |  | Coarse filter with self-draining for small particles |  |
|  |  | Flow: filtration capacity at 7 bar pressure | Minimum 630 m3/hour |
|  |  | Working pressure | Minim 16 bar |
|  |  | Retention of solid particles, up to | 1 μm |
|  |  | Clogging (clogging) indicator gauge | Yes |
|  |  | Coal Tower |  |
|  |  | Flow: filtration capacity at 7 bar pressure | Minim 490 m3/hour |
|  |  | Working pressure | Minim 16 bar |
|  |  | Maximum amount of residual oil | Maximum 0.003 mg/m3 |
|  |  | High-efficiency filter |  |
|  |  | Flow: filtration capacity at 7 bar pressure | Minimum 630 m3/hour |
|  |  | Working pressure | Minim 16 bar |
|  |  | Filtration efficiency | 99.99% |
|  |  | Solid particle retention | bigger than 0.01 μm |
|  |  | Dust filter with manual drain |  |
|  |  | Solid particle retention | bigger than 0.01 μm |
|  |  | Working pressure | Minim 16 bar |
|  |  | Flow: filtration capacity at 7 bar pressure | Minimum 630 m3/hour |
| Compressed air tank | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Galvanized tank | Yes |
|  |  | Capacity | Minimum 1500 L |
|  |  | Maximum working pressure supported by the vessel | Minimum 11 bar |
|  |  | Safety valve and pressure gauge | Yes |
|  |  | Valve for automatic discharge of condensate from the air tank | Yes |
| Medical oxygen production station | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Oxygen production technology | Pressure Swing Adsorption (PSA) |
|  |  | Oxygen purity | 93%±2 |
|  |  | Oxygen flow at a concentration of 93% | Minimum 30 m3/hour |
|  |  | Power supply | 220 V / 50 Hz |
|  |  | Compressed air inlet pressure | Minimum 7.5 bar |
|  |  | The installation must not affect the environment and the ozone layer | Yes |
|  |  | Operating temperature, with values ranging from | +10°C to +40°C |
|  |  | Central Control Panel | Yes |
|  |  | Continuous monitoring with +/- 1% accuracy of compressed air inlet pressure in the oxygen generator | Yes |
|  |  | Continuous monitoring with integrated pressure gauges on the Oxygen generator for Zeolite, Oxygen and Air Tanks | Yes |
|  |  | System with valves to block the flow of oxygen in case of a drop in the concentration below the level of 90%, after returning to normal, the valves open automatically for the distribution of oxygen to the hospital. | Yes |
|  |  | Continuous monitoring of the purity of the oxygen produced by the generator with included analyzer | Yes |
|  |  | Monitoring the air temperature at the entrance to the generator so as not to exceed the temperature recommended by the manufacturer, and in case of exceeding - to generate an alarm about it | Yes |
|  |  | Tank Pressure Monitoring System | Yes |
|  | Control Panel / User Interface | Control panel / user interface, fully automated PLC, touch screen | Yes |
|  |  | Oxygen Generator Control Panel | Yes |
|  |  | Digital display language: | At least English and Romanian |
|  |  | Oxygen purity [%] | Yes |
|  |  | Oxygen production trend [Nm3/hour] | Yes |
|  |  | Outlet pressure | Yes |
|  |  | System status, including current maintenance needs | Yes |
|  |  | Cumulative hours of operation | Yes |
|  |  | Calculation of the system efficiency, how much air is converted in usable oxygen | Yes |
|  |  | Audible and visual alarms for: |  |
|  |  | High temperature | Yes |
|  |  | Low/High Pressure | Yes |
|  |  | Low oxygen concentration (<90%) | Yes |
|  |  | Power outage, system failure | Yes |
|  |  | Remote monitoring and operation via computer or smartphone or tablet. Integrated software with the necessary accessories for connection. Monitoring and operation is carried out through TCP/IP or similar protocols. | Yes |
|  |  | Permanent monitoring of instantaneous oxygen flow in the network (liters/minute or m3/hour) | Yes |
|  |  | Permanent monitoring of oxygen pressure in the hospital network | Yes |
|  |  | Permanent monitoring of generator uptime | Yes |
|  |  | Internal storage of data and operating parameters (oxygen purity, flow rates, alarms, system performance) | Yes, for at least 6 months |
|  |  | Noise attenuator to reduce noise levels, mounted on the discharge side of the oxygen generator | Yes |
|  |  | Connecting a Visual alarm in the Operator's room, to observe the shutdown of the oxygen generator | Yes |
|  |  | Internal storage of alarms with the date and time produced | Yes |
|  | Others | All oxygen generator pipes must be made of medical grade pipe, made of copper or stainless steel min AISI 304 | Yes |
| Oxygen Tank | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Capacity | Minimum 1500 L |
|  |  | Pressure reducer at the outlet of the oxygen tank that will operate with values between | 3 - 6 bar adjustable |
|  |  | Maximum working pressure supported by the vessel | Minimum 11 bar |
|  |  | Pressure monitoring at the outlet | Yes |
| Antibacterial Medical Oxygen Filtration System | Quantity | 1 pc. |  |
|  | Overview of functional requirements | Oxygen filtration flow rate | Minimum 60 m3/hour |
|  |  | Maximum working pressure | Minium 16 bar |
|  |  | Solid particle retention filter | Yes |
|  |  | Antibacterial/sterile filter | Yes |
|  |  | Clogging (clogging) indicator gauge | Yes |
| Oxygen Pipe Connection | Quantity | 1 pc. |  |
|  |  | Connection box | Yes |
|  |  | Safety valve on the connecting pipe for releasing oxygen into the outside air in case of emergency. | Yes |
|  |  | Oxygen pressure monitoring and operational alarm system connected as a priority | Yes |
| Prefabricated container | Quantity | 1 pc. |  |
|  | Container, features | Container system or similar | Yes |
|  |  | Transportable | Yes |
|  |  | System plug and play | Yes |
|  |  | The total area, height and volume of the container shall be sufficient for access by technical personnel for monitoring and technical maintenance purposes. | Yes |
|  |  | Resistant to mechanical shock (handling or movement) | Yes |
|  |  | The container or construction must be treated and painted on the surface to prevent corrosion | Yes |
|  |  | Automatic system for closing/opening the external air intake grille. | Yes |
|  |  | The floor should be covered with a washable surface to maintain cleanliness | Yes |
|  |  | Drainage system for the evacuation of condensate from the compressor and dryer. | Yes |
|  | Environmental conditions | Air conditioning to maintain indoor temperature and humidity strictly in accordance with the recommendations provided by the manufacturer of the oxygen generator offered | Yes |
|  |  | Freestanding air intake, thermally conditioned by the compressor | Yes |
|  |  | Waterproofed container | Yes |
|  |  | Resistant and airtight to dust, sand and fog corrosion | Yes |
|  |  | Interior lighting system for the operator | Yes |
|  |  | Emergency interior lighting system. | Yes |
|  |  | Outdoor lighting system in the entrance area. | Yes |
|  |  | Container Oxygen Concentration Monitoring System | Yes |
|  |  | At least 4 220V electrical outlets must be installed for corrective maintenance and preventive maintenance works. | Yes |
|  |  | Smoke detectors with alarm | Yes |
| Oxygen Plant Installation Requirements | Delivery & Shipping | Transportation of oxygen station equipment according to the cumulative recommendations of the manufacturer(s) of all basic components (oxygen generator, compressor, air dryer, other components) | Yes |
|  |  | All equipment required for unloading and placement must be included. | Yes |
|  | Commissioning | All pipes, connectors, and accessories necessary to connect to the hospital's main distribution pipe must be included. | Yes |
|  |  | The connection to the existing oxygen network must be carried out without risk to the hospital, as well as in the place necessary for it. | Yes |
|  |  | All the necessary electrical components for equipment interconnection. | Yes |
|  |  | Commissioning by representative company team | Yes |
|  |  | Test after commissioning | Yes, with certified tester for oxygen purity and dew point, testing of the flow rates, safety tests |
| After-sales service | User training and maintenance | Warranty period | Minimum 36 months |
|  |  | Local representative in Republic of Moldova or Romania | Yes |
|  |  | Training of a bioengineer/responsible person in the process of technical use of the system as well as carrying out preventive and corrective maintenance tasks. | Yes |
|  |  | Technical support from the supplier for the process of using the control system and carrying out preventive and corrective maintenance works | Yes |
|  |  | Documentation (included, minimum in English): Certificate of quality, calibration and inspection | Yes, Original certificates at delivery |
|  |  | Access keys, PIN codes, passwords and technical manuals and user manuals must be delivered at least in english | Yes |
|  |  | Remote support for maintenance | Yes |
|  |  | Presentation of the training certificate from the manufacturer of an engineer of the company that will perform the installation and maintenance work of the device | Yes |
|  | Lifespan – Wear Warranty | Support warranty and the presence of spare parts from the manufacturer for a period of 10 years | Yes |

  - Schedule of Requirements

| - List of Goods and Delivery Schedule |  |  |  |  |
| --- | --- | --- | --- | --- |
| Line Item | Description of Goods | Quantity | Unit of measure | Delivery Schedule from date of Contract |
| 1 | Medical oxygen generator | 1 | EA | 90Days |

| - Consignee Address and Consignee-wise Quantity Distribution |  |  |  |  |
| --- | --- | --- | --- | --- |
| Line Item | Consignee Address | Contact person | Quantity | Unit of measure |
| 1 | United Nations Population Fund 131, 31 August 1989 street, Chisinau, Republic of Moldova, MD-2012 | Will Provide while awarding the contract | 1 | EA |

| - List of Related Services and Completion Schedule |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| No. | Description of Service | Quantity | Physical Unit | Place where Services shall be performed | Final Completion Date(s) of Services |
| 1 | Installation of Medical oxygen generator | 1 | EA | Ministry of Interior Hospital, Chisinau |  |

Commissioning and Testing:

The supplier is responsible for testing the Medical Oxygen Generator in compliance with the minimum conditions outlined below;

Confirm that the generator produces oxygen at the specified purity level ≥93% purity for medical applications and that flow rates meet the design requirements.

- Ensure the system is calibrated according to manufacturer specifications, including adjustments for pressure, flow rates, and alarm thresholds.

- Verify that all safety mechanisms, including pressure relief valves, oxygen purity alarms, and emergency shutdowns, are functioning properly.

- Assess the generator’s performance under expected operating conditions, such as ambient temperature, humidity, and power stability.

- Provide detailed commissioning documentation in English and Russian and offer training to facility staff on system operation, maintenance, and emergency protocols.

Training:

- Train bioengineers on starting, stopping, and monitoring the oxygen generator, including reading control panels and interpreting key performance indicators.

- Provide guidance on regular maintenance tasks, such as filter replacements, system cleaning, and checking for leaks, to ensure ongoing reliability.

- Instruct on safety measures, including handling oxygen safely, recognizing alarm signals, and responding to emergency shutdowns or other critical situations.

- Train on basic troubleshooting techniques for common issues, such as low oxygen purity or pressure drops, and the steps to take if more complex problems arise.

- Teach the importance of maintaining accurate operational logs and reporting maintenance or operational issues according to facility protocols.

- Training must be held in Romanian or Russian.

Support Requirement During Warranty Period:

- Carry out routine preventative maintenance every 6 months or as recommended by the manufacturer, including inspection, cleaning, lubrication, filter replacements, and calibration. Provide detailed maintenance records and notify the facility of upcoming maintenance activities in advance.

- Ensure minimal downtime by guaranteeing a response time of no more than 24 hours for critical issues and 48 hours for non-critical issues. For any issue impacting system functionality, the supplier should commit to resolving within 72 hours to ensure the system remains operational.

- Provide 24/7 technical support by phone or online for immediate troubleshooting assistance. On-site support should be made available within the stipulated response time if remote troubleshooting fails to resolve the issue.

- Supply and replace any defective parts at no additional cost under the warranty, with an emphasis on quick procurement and installation. This includes compressors, filters, valves, sensors, or any part affecting output performance and purity standards.

- Conduct a full annual system check, including output verification and calibration to ensure compliance with required oxygen purity and flow rates. Provide a detailed report on performance and any adjustments or parts replacements made during the check.

# SECTION III: UNFPA General Conditions of Contract

The General Conditions of Contract can be found at:

http://www.unfpa.org/resources/unfpa-general-conditions-contract

# SECTION IV: UNFPA Special Conditions for Contracts

| WARRANTY | The warranty period shall be 36 months. Details on Warranty Services required are included in Section II: Technical Specifications and Schedule of Requirements. |
| --- | --- |
| GOODS AND SERVICES DEFINED | Goods are hereinafter deemed to include, without limitation, equipment, spare parts, commodities, raw materials, components, customized and standard software as required, intermediate products and products which the Supplier is required to supply under the Purchase Order. Services are to include design, installation and commissioning, training services, technical assistance and warranty services as required to supply in the Purchase Order. |
| AFTER SALES SERVICES | Bidder must be maintaining the unit for 3 years. Ensuring the availability of spare parts for the product’s lifecycle 10 years, often with a priority supply arrangement to reduce wait times for critical parts. Offering 24/7 access to technical assistance, either remote (phone or online) or on-site, to resolve any operational issues that arise and minimize downtime |
| TRANSPORTATION AND FREIGHT | Responsibility for transportation of the Goods shall be as specified in the INCOTERMS. All non-containerized Goods must be shipped below deck |
| SHIPPING AND PAYMENT INSTRUCTIONS | Access the following link for shipping and payment instructions: Shipping Instructions |

# SECTION V: Bidding Forms

The following checklist is provided as a courtesy to Bidders. Please use this checklist while preparing the bid to ensure that your bid contains all required information. This checklist is for the Bidder’s internal reference and does not need to be submitted with the bid.

| ACTIVITY | LOCATION | YES / NO/ NOT APPLICABLE | REMARKS |
| --- | --- | --- | --- |
| Have you noted the bid closing deadline? | Cover letter, #5 |  |  |
| Have you read and understood all of the Instructions to Bidders in Section I of the bidding documents? | Section I |  |  |
| Have you reviewed and agreed to the UNFPA General Conditions of Contract? | Section III |  |  |
| Have you reviewed and agreed to the UNFPA Special Conditions for Contracts? | Section IV |  |  |
| Have you completed the Bid Confirmation Form? | Section V, 1 |  |  |
| Have you completed the Bid Submission Form? | Section V, 2 |  |  |
| Have you completed the Bidder’s Identification Form? | Section V, 3 |  |  |
| Have you completed the Product Item Overview Form? | Section V, 4 |  |  |
| Have you completed and signed the Price Schedule Form? | Section V, 5 |  |  |
| Have you reviewed all of the relevant contract form(s)? | Section VI |  |  |
| Have you provided evidence that your firm is established as a company and legally incorporated in the country where it resides? | Section I, Sub-Clause 7.2, a |  |  |
| Have you prepared a copy of your valid manufacturing license from the country of manufacturing? | Section I, Sub-Clause 7.2, b. |  |  |
| Have you provided written confirmation that your company is neither suspended by the United Nations system nor debarred by the World Bank Group? | Section I, Sub-Clause 2.4 |  |  |
| Have you prepared documentary evidence that the goods conform to the technical specifications and standards specified in Section II Technical Specifications and Schedule of Requirements? | Section I, Sub-Clause 7.3, a. |  |  |
| Have you prepared product catalogues containing pictures of the product(s)? | Section I, Sub-Clause 7.3, c. |  |  |
| Have you prepared the manufacturer’s technical product specifications or data sheets? | Section I, Sub-Clause 7.3, d. |  |  |
| Have you provided any copies of current certificates as stated in the Technical Specifications and Schedule of Requirements, in Section II? | Section I, Sub-Clause 7.3, f. |  |  |
| Have you provided a copy of the valid authorization letter issued by the manufacturer for each product, if you are not the manufacturer? | Section I, Sub-Clause 7.3, g. |  |  |
| Have you furnished a list of full particulars, regarding the available sources and current prices of space parts, special tools, etc., necessary for the proper and continuing functions of the goods within the Product Item Overview Form, Section V, 5? | Section I, Sub-Clause7.3, h. |  |  |
| Have you sealed and marked the bids according to Instructions to Bidders Clause 13 (hard copy bids) or Clause 14 (electronic bids)? | Section I, Sub-Clause 13 & 14 |  |  |
| If submitted electronically, is the file size of the bid less than 10MB? (If the file size is above 10MB, refer to Instructions to Bidders Sub-Clause 14.4) | Section I, Sub-Clause 14.4 |  |  |
| Have you prepared a copy of the previous year’s audited company Balance and Financial Statements? | Section I, Sub-Clause 27.3 |  |  |
| For non-manufacturer Bidders: Have you provided a legally enforceable authorization from the manufacturer, assuring full guarantee and warranty obligations as per the tender conditions for the goods offered? | Section I, Sub-Clause 27.3, a. |  |  |
| Have you provided evidence that you, as authorized by the manufacturers, have supplied and provided after sales service for similar goods indicated in the tender requirements in the last three years, and that the goods are in satisfactory operation? | Section I, Sub-Clause 27.3, b. |  |  |

# 1. Bid Confirmation Form

[Complete this page and return it prior to bid opening]

|  |  | Date: |
| --- | --- | --- |
| To: | UNFPA Republic of Moldova | mail: tender.mda@unfpa.org |
| From: | [Company name] |  |
|  | [Contact person] |  |
|  | [Telephone] |  |
|  | [Email address] |  |
|  | [Postal address] |  |
| Subject: | ITB No.: UNFPA/MDA/24/004 |  |

YES, we intend to submit an bid.

NO, we are unable to submit a bid in response to the above mentioned Invitation to Bid due to the following reason(s):

( )	The requested products and services are not within our range of supply

( )	We are unable to submit a competitive bid for the requested products at the moment

( )	The requested products are not available at the moment

( )	We cannot meet the requested specifications

( )	We cannot offer the requested type of packing

( )	We can only offer FCA prices

( )	The information provided for quotation purposes is insufficient

( )	Your ITB is too complicated

( )	Insufficient time is allowed to prepare a quotation

( )	We cannot meet the delivery requirements

( )	We cannot adhere to your terms and conditions (please specify: payment terms, request for performance security, etc)

( )	We do not export

( )	Our production capacity is currently full

( )	We are closed during the holiday season

( )	We had to give priority to other clients’ requests

( )	We do not sell directly, but through distributors

( )	We have no after-sales service available in the recipient country

( )	The person handling bid is away from the office

( )	Other (please specify)

Please confirm one of the following two options:

( )	We would like to receive future ITBs for this type of goods

( )	We don’t want to receive ITBs for this type of goods

If UNFPA has questions to the Bidder concerning this NO BID, UNFPA should contact Mr./Ms._________________, phone/email ________________, who will be able to assist.

# 2. Bid Submission Form

[The Bidder shall fill in this form in accordance with the instructions indicated. No alterations to its format shall be permitted and no substitutions shall be accepted.]

Date: [insert date (as day, month and year) of Bid Submission]

ITB No.: UNFPA/MDA/24/004

To: UNFPA

Dear Sir / Madam,

We the Undersigned have examined and have no reservations to the Bidding Documents No. UNFPA/MDA/24/004 and amendments We hereby offer to supply, in conformity with the Bidding Documents and in accordance with the Delivery Schedules specified in the Schedule of Requirements, the following goods and related services ___________________________________________________ which are subject to UNFPA General Conditions of Contract and other terms and conditions specified in the document.

We agree to abide by this bid for a period of 90 days from the date fixed for opening of bids in the Invitation to Bid, and it shall remain binding upon us and may be accepted at any time before the expiration of that period.

We have no conflict of interest in accordance with Instructions to Bidders Sub-Clause 2.1;

Our firm, its affiliates or subsidiaries—including any subcontractors or suppliers for any part of the contract—have not been declared ineligible by UNFPA, in accordance with Instructions to Bidders Sub-Clause 2.2;

We understand that you are not bound to accept the lowest evaluated bid or any other bid that you may receive.

Dated on .............day of ......................................[year].

| Signature: | ……………………………………………………………… [insert signature of person whose name and capacity are shown] |
| --- | --- |
| In the capacity of: | ……………………………………………………………… [insert legal capacity of person signing the Bid Submission Form] |
| Name: | ……………………………………………………………… [insert complete name of person signing the Bid Submission Form] |
| Company: | ……………………………………………………………… [insert name of company] |

# 3. Bidders Identification Form

Bid No. UNFPA/MDA/24/004

- Organization

| Company/Institution Name |  |
| --- | --- |
| Address, City, Country |  |
| Telephone/FAX |  |
| Website |  |
| Date of establishment |  |
| Legal Representative: Name/Surname/Position |  |
| Legal structure: natural person/Co.Ltd, NGO/institution/other (please specify) |  |
| Organizational Type: Manufacturer, Wholesaler, Trader, Service provider, etc. |  |
| Areas of expertise of the organization |  |
| Current Licenses, if any, and permits (with dates, numbers and expiration dates) |  |
| Years supplying to UN organizations |  |
| Years supplying to UNFPA |  |
| Number of years in the industry and experience with similar projects |  |
| Record of successful installations of similar oxygen generators, especially in healthcare settings |  |
| Production Capacity |  |
| Subsidiaries in the region (please indicate names of subsidiaries and addresses, if relevant to the bid) |  |
| Commercial Representatives in the country: Name/Address/Phone (for international companies only) |  |

- Quality Assurance Certification

| International Quality Management System (QMS) |  |
| --- | --- |
| List of other ISO certificates or equivalent certificates |  |
| Presence and characteristics of in-house quality control laboratory (if relevant to bid) |  |

- Expertise of Staff

| Total number of staff |  |
| --- | --- |
| Number of staff involved in similar supply contracts |  |
| Names and qualifications of technical staff, available for on-site support in the destination country (CV, copy of Certifications) |  |

- Client Reference List

Please provide references of main client details from similar projects.

| Name of company | Contact person | Telephone | E-mail |
| --- | --- | --- | --- |
| 1. |  |  |  |
| 2. |  |  |  |
| 3. |  |  |  |

- Contact details of persons that UNFPA may contact for requests for clarification during bid evaluation

| Name/Surname |  |
| --- | --- |
| Telephone Number (direct) |  |
| Email address (direct) |  |

P.S.: This person must be available during the next two weeks following receipt of bid

- Product Item Overview Form

Important. Please note that bidders should provide their own technical specifications rather than copying and pasting the specifications provided by UNFPA. This helps ensure that your proposal accurately reflects the unique features and standards of your products

| Item No. | Description and minimum /mandatory specifications | Description of items offered and Bidder’s statements on deviations (To be completed by the Bidder) | Compliant? (Y/N) (To be completed by UNFPA during evaluation) |
| --- | --- | --- | --- |
| 1 | Medical oxygen generator is a device which generate pure oxygen for medical use, for ventilatory therapy (ICU ventilators, Anesthesia machine) As per the requirement specified in the SECTION II: Technical Specifications and Schedule of Requirements |  |  |

# 5. Price Schedule Form

| BIDDER’S TOTAL PRICES (Price & Currency to be entered by Bidder): |  |
| --- | --- |
| TOTAL FIRM FCA PRICE |  |
| TOTAL FIRM CPT PRICE |  |
| TOTAL PRICE FOR SERVICES (if applicable) |  |
| FREIGHT COST PER 20/40 FT CONTAINER (if applicable) |  |

| BIDDER’S PRICES FOR GOODS (Price & Currency to be entered by Bidder): |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| ITEM/LOT | DESCRIPTION OF THE GOODS | QTY (a) | CURRENCY: |  |  |  |
|  |  |  | UNIT PRICE FCA (b) | UNIT PRICE CPT (c) | TOTAL PRICE FCA (a)x(b) | TOTAL PRICE CPT (a)x(c) |
| 1. | Medical oxygen generator | 1 |  |  |  |  |
| 2. | Spare parts for 2 years(list all the items and cost per items) |  |  |  |  |  |
| 3. |  |  |  |  |  |  |

| BIDDER’S PRICES FOR SERVICES (Price & Currency to be entered by Bidder): |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| ITEM/LOT | DESCRIPTION OF THE SERVICES | COUNTRY OF ORIGIN | QUANTITY AND PHYSICAL UNIT (a) | UNIT PRICE (b) | TOTAL PRICE PER SERVICE (a)x(b) |
| 1. | Installation, Commissioning and Testing |  |  |  |  |
| 2. | Training |  |  |  |  |

| #### BIDDER’S DELIVERY DATA |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Country of origin of offered products: | Item 1 |  |  |  |  |
| FCA point(s) of delivery for offered products: | Item 1 |  |  |  |  |
| Delivery time (FCA from date of order): | Item 1 |  |  |  |  |
| Shipment dimensions of offered products (including package): |  | Gross weight | Total volume | Containers (if applicable): |  |
|  |  |  |  | Number | Size |
|  | Item 1 |  |  |  |  |
|  | Item 2 |  |  |  |  |
|  | Item 3 |  |  |  |  |
|  | Total |  |  |  |  |

| BIDDER'S SIGNATURE AND CONFIRMATION OF THE ITB |  |
| --- | --- |
| PROVIDED THAT A PURCHASE ORDER IS ISSUED BY UNFPA WITHIN THE REQUIRED BID VALIDITY PERIOD, THE UNDERSIGNED HEREBY COMMITS, SUBJECT TO THE TERMS OF SUCH PURCHASE ORDER, TO FURNISH ANY OR ALL ITEMS AT THE PRICES OFFERED AND TO DELIVER SAME TO THE DESIGNATED POINT(S) WITHIN THE DELIVERY TIME STATED ABOVE. |  |
| Exact name and address of company COMPANY NAME ADDRESS PHONE NO. FAX NO. EMAIL ADDRESS OF CONTACT PERSON OTHER EMAIL ADDRESSES | AUTHORIZED SIGNATURE DATE NAME OF AUTHORIZED SIGNATORY (TYPE OR PRINT) FUNCTIONAL TITLE OF SIGNATORY WEB SITE |

# SECTION VI: Contract Forms

- xxx