## Pricing Sheet
S/N | Item | Qty | Unit Price FCA | Total Cost FCA | Total Freight Cost to Hargeisa Group Hospital (shall include CRATING & EXPORT PACKING) | Total Insurance | Total DAP to Hargeisa Group Hospital | Volume Discounts (if applicable) | Delivery Time from confirmed order (indicate if split shipments are applicable) | indicate total weight (kg) & volume (m3) | mode of transport | Brand / Model Number (Mandatory that you provide the required information in this column) | Country of Origin of Product (Mandatory that you provide the required information in this column) | Warranty period, if different from requirements stipulated in the technical specifications. Warranty begins from date of acceptance of goods by FMoH through their signature on UNOPS' Receiving & Inspection Reports (RIRs)
|  |  | USD | USD | USD | USD | USD | USD
IMPORTANT: Pricing Sheets submitted in any other format will not be accepted by UNOPS
1 | Electrosurgical Unit Analyzer: | 1
2 | Vital signs Simulator: | 1
3 | Defibrillator analyzers: | 1
4 | Electrical Safety Analyzer | 1
5 | X-Ray Test Tools: | 1
6 | Gas Flow Analyzer: | 1
Grand Total Cost for Equipment
Vendor to undertake installation of above equipment . Alll-inclusive lump sum cost for installation of equipment in ____________. This amount should include for transportation, accommodation meals etc.
| Please indicate the time it would take to complete installlation of the equipment in | We estimate it would take ……………………….. (indicate period) to complete installation
Vendor to undertake inspection, testing and commissioning of equipment in ___________ . All-inclusive lump sum cost for inspection, testing and commissioning of installed equipment.
| Please indicate the time it would take to complete inspection, testing and commissioning of the equipment in | We estimate it would take ……………………….. (indicate period) to complete inspection, testing and commissioning of installed equipment
Vendor to undertake training of equipment in ___________ . All-inclusive lump sum cost for training. This amount should include for transportation, accommodation meals etc. for the trainers
| Vendors are required to provide user training in _____________, of all equipment supplied | Please provide your full training plan as a supporting document
Local Warranty Support (please provide name / address / contact details of the warranty centre in ____________. Warranty begins from date of UNOPS clients' signature on UNOPS' Receiving & Inspection Report (RIR) as their acceptance of goods supplied. |  | Local Warranty Centre Contact Details:
Acknowledgement
We confirm that all specifications meet the UNOPS requirements and our quotations are valid for a period of 90 (ninety) days from date of bid closing.
Company Name:
Name of Company Representative
Email of Company Representative:
Telephone ( Landline):
Telephone (Mobile):
Signature of Company Representative:
Acknowledgement
IMPORTANT: In order for your company to remain as a registered vendor, you must return this acknowledgement even if you do not intend to submit a proposal at this time.
If you do not intend to submit a proposal to the UNOPS, please indicate the reason:
Mark "x"
We do not have the capacity to submit a proposal at this time.
We cannot meet the technical requirement for this RFQ.
We do not think we can make a competitive offer at this time.
Other reasons: (Please specify):
Name and Signature of Company's Representative
Date:
Company Stamp

## Technical Sheet
| SPECIFICATIONS: All goods and materials to be incorporated in the Goods must meet international electricity standards, be new, unused, merchantable quality and of the most recent or current models, and that they incorporate all recent improvements in design and materials, unless provided for otherwise in the technical specifications herein or resultant contract(s). | The products offered are in accordance with required specifications and technical requirements. (Yes / No). Any Deviation must be listed below.
| Bid Validity: All offers must be valid for a period of 90 (ninety) days from the date of bid closing. | Yes | No- mandatory to indicate variation in specifications.
EQUIVALENT PRODUCTS ARE ACCEPTABLE Standards for workmanship, process, material, and equipment, as well as references to brand names or specified by UNOPS in the ITB, are intended to be descriptive only and not restrictive. The Bidder may offer other standards of quality, brand names, and/or catalogue numbers, provided that it demonstrates, to UNOPS’s satisfaction, that the substitutions ensure substantial equivalence or are superior to those specified in the ITB.
1 . INTEGRATING SPHERE FOR LUMEN MEASUREMENT OF LUMINARIES
1 - Electrosurgical Unit Analyzer: | Electrosurgical Unit Analyzer:
General Description:
The unit is intended to measures ESU output, high frequency leakage, verification tests on the return electrode contact quality monitors (RECQM Test). The Testing tool has to be able of Instantaneous measurement or sample over time. The Tool has to have oscilloscope output for Waveform visualization
Specific required Features:
· MODES OF OPERATIONS: Continuous Operation, Single Operation, Power Distribution, RF Leakage current, RECQM, Manual remote
Automatic power distribution measurement, including power, current, peak-to-peak voltage closed load only), and crest factor
Oscilloscope output
High-frequency leakage measurements with accuracy of ± 5 % of reading
128 internal user-selectable test loads from 10 Ω to 5200 Ω
Foot-switch output for triggering the ESU under test
Automated test protocols
Large display 240 x 64 pixel matrix
RS-232 and printer interface
Data Input/Output: Parallel printer port and bidirectional RS232
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· User Manual, operation Manual,
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model (EQUIVALENT PRODUCTS ARE ACCEPTABLE) : FLUKE QA-ES Electrosurgery Analyzer
http://www.flukebiomedical.com/biomedical/usen/Electrosurgery-Analyzers/QA-ES-Electrosurgery-Analyzer.htm?PID=56636
2 - Vital signs Simulator: | Vital signs Simulator:
General Description:
Comprehensive Patient Simulator to test vital signs monitoring equipment. The Equipment has to test ECG (including fetal ECG/IUP and arrhythmias), respiration, temperature, IBP/cardiac catheterization, cardiac output and NIBP. The Simulator shall have the following additional features: Wireless PC communication, customized presets, auto-sequences, barcode scanning, onboard memory, direct data capture, printing functions and single-step adjustability.
Specific Features
· All-in-one patient Simulator;
· ECG simulator, fetal simulator, arrhythmia simulator, respiration simulator, temperature simulator, IBP simulator, cardiac output simulator, cardiac catheterization simulator, NIBP simulator, SpO2 tester, and is the premier SpO2 simulator to test Rainbow multi-wavelength waveforms
· Custom SpO2 r-curve
· Static pressure linearity testing
· Repeatable NIBP simulation for dynamic pressure repeatability testing
· Physiologically synchronized pulses across all parameters
· Barcode scanning and printing functionality
· Onboard, customizable patient pre-sets and auto sequences for fast/easy testing
· Multi-language user interface (English)
· Integrated, easily-replaceable long-life battery
· PC-interface
· Wireless communication for remote PC control of test device, as well as data transfer and automated regulatory reporting
· User Manual, operation Manual in English Language.
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model (EQUIVALENT PRODUCTS ARE ACCEPTABLE): FLUKE ProSim 8 Vital Signs Patient Simulator
http://www.flukebiomedical.com/biomedical/usen/patient-simulators/prosim-8-vital-signs-simulator.htm?PID=72624
3 - Defibrillator analyzers: | Defibrillator analyzers:
General Description
The Defibrillator analyzer is intended to be used for testing life-support cardiac-resuscitation equipment, including mono- and bi-phasic defibrillators and AEDs. The Defibrillator Analyzer has to be capable to test the global spectrum of established pulse shapes, compatible with AED with accuracy readings ± 1% of reading +0.1J. The Defibrillator Analyzer should be capable to test between 6 and 10 independent ECG outputs with 12 lead combinations for standardized clinical signals.
Specific Features
· IEC 60601-2-4 Compliance
· Lown, Edmark, trapezoidal, biphasic and pulsed-biphasic defibrillation technology compatibility
· AED compatibility
· 12-lead ECG simulation
· Internal pacer brand selections
· Flexible heart-rate settings (1BPM step)
· Waveform capture, store and replay
· Accuracy ± 1% of reading + 0.1J
· Long lasting, rechargeable battery
· User Manual, operation Manual in English Language.
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model (EQUIVALENT PRODUCTS ARE ACCEPTABLE): Impulse 6000D/7000DP Defibrillator Tester
Website Link:
http://www.flukebiomedical.com/Biomedical/usen/Defibrillator-Analyzers/Impulse-7000DP-defibrillator-tester-analyzer.htm?PID=56326
4 - Electrical Safety Analyzer | Electrical Safety Analyzer
General Description
The Electrical Safety Analyzer that performs electrical safety testing on medical equipment both in the field and in facilities especially designed for to healthcare technology management professionals. The device shall perform all primary electrical safety tests, including line (mains) voltage, ground wire (protective earth) resistance, insulation resistance, device current, and lead (patient) leakage tests. It also offers ECG simulation and point–to-point voltage, leakage, and resistance tests. Versatile to any global electrical safety standard of choice, the Device has to be capable to test according to ANSI/AAMI ES-1, NFPA-99, IEC62353, VDE751, IEC60601-1 (2nd and 3rd editions) and AS/NZS 355 standards.
Specific Features
On-board automation with automated test sequences for easy compliance to key global electrical safety standards (ANSI/AAMI ES-1 (NFPA-99), IEC62353 (VDE751), IEC60601-1 2nd and 3rd editions, and AS/NZS 3551)
Portable, ergonomic design with an integrated handle and tilt stand
Human-factors-designed user interface for streamlined testing
ECG waveform tests and dual-lead measurements to combine the functionality of a simulator, multimeter and safety analyzer in a single test tool.
20 A at 120 V current capability
Five applied parts jacks and easy ECG snap connection; optional expander box for up to 12-lead ECG testing
Easy data entry through barcode, external keyboard or on-board keypad
Wireless communication plus removable memory card for fast and convenient data storage and exchange
Replaceable of mains fuses
Custom language selections: English,
Large, easy-to-read display with adjustable contrast
Rigorously tested for rugged field applications, with CE and CSA
Two-year extended warranty
· User Manual, operation Manual in English Language.
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model(EQUIVALENT PRODUCTS ARE ACCEPTABLE): ESA620 Electrical Safety Analyzer
Website Link:
http://www.flukebiomedical.com/biomedical/usen/Electrical-Safety-Analyzers/ESA615-Electrical-Safety-Analyzer.htm?PID=76829
5 - X-Ray Test Tools: | X-Ray Test Tools:
General Description
X-Ray Test Tools for assuring quality and safety of diagnostic x-ray imaging systems. It has to possible to select all-in-one-exposure solid-state detector, dosimeter, ion chambers, optional mA/mAs invasive shunt or non-invasive clamp device, and choice of handheld display or laptop interface (both completely wireless). With long battery life ensures uninterrupted operation all day.
Specific Features
Choice of all-in-one detector, dosimeter, integral ma/mAs and handheld display or user’s own laptop interface.
100 % wireless operation
Compact handheld design
Prompt Sets up in seconds
Easy displays of all values in one shot (all-in-one detector)
Simple user interface
Sturdy construction for long-term reliability
40 kHz kV sampling rate
Two-year extended warranty
· User Manual, operation Manual in English Language.
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model (EQUIVALENT PRODUCTS ARE ACCEPTABLE): TNT 12000 X-Ray Radiation Dosimeter Test Tools
Website link:
http://www.flukebiomedical.com/biomedical/usen/x-ray-test-tools/TNT-12000-x-ray-radiation-dosimeter-test-tools.htm?PID=56783
6 - Gas Flow Analyzer: | Gas Flow Analyzer:
General Description
Gas Flow Analyzer for general-purpose gas-flow analyzer, compact and portable for evaluating performance of a wide variety of medical gas-flow/pressure devices and measures 16 ventilator parameters. Front panel with user-friendly interface with capability to displays waveforms on the tool's screen. Capable to store test results to be retrieved and exported to PC. The base unit measures high- and low-flow ranges, volume, pressure, and oxygen concentration.
Specific Features
Bidirectional flow (high- and low-flow ranges), volume, vacuum, pressure, and oxygen-concentration measurements
16 ventilator parameter measurements
Trending and statistical analysis of all measured values
Onboard graphical display
Portable and compact
RS232 for computer control
Memory for storing results
VT for Windows PC software
Optional sensor assembly for temperature and humidity measurements
Two-year extended warranty
· User Manual, operation Manual in English Language.
Power: 115/230 VAC; 48 to 66 Hz, 35 VA, Shuko Type
· The equipment has to come complete with all accessories for a full and complete utilization.
Reference Brand and Model (EQUIVALENT PRODUCTS ARE ACCEPTABLE): VT MOBILE Gas Flow Analyzer
Website link:
http://www.flukebiomedical.com/biomedical/usen/Gas-Flow-Analyzers/VT-MOBILE-Gas-Flow-Analyzer.htm?PID=56858