United Nations Population Fund, UNFPA

International Zone (IZ) – UNAMI compound – D2 building

Iraq

Website: iraq.unfpa.org/

Date: 26 July 2017

Request for Quotation No. UNFPA/IRQ/RFQ/17/010

Dear Sir/Madam,

We hereby solicit your quotation for the supply of Medical supplies as Specified in Annex1 of this RFQ.

If you are interested in submitting a quotation for these items, kindly fill in the attached Quotation Submission Form – Annex 2, and send by email to: bid.iraq@unfpa.org or the address indicated below not later than 30 July 2017, and 21:00 (GMT+3).

Name of Contact Person in UNFPA: 	Mr. Habeeb Abdullah

Street address:	Iraq – Baghdad – International Zone (IZ) – UNAMI compound – D2 building – UNFPA office

Email: bid.iraq@unfpa.org

Quotations submitted by email must be free from any form of virus or corrupted contents, or the quotations shall be rejected.

It shall remain your responsibility to ensure that your quotation will reach the address above on or before the deadline. Quotations that are received by UNFPA after the deadline indicated above, for whatever reason, shall not be considered for evaluation.  If you are submitting your quotation by email, kindly ensure that it is signed and in the .pdf format, and free from any virus or corrupted files.

Full acceptance of the UNFPA General Terms and Conditions is mandatory. They can be located on this webpage at: http://www.unfpa.org/resources/unfpa-general-conditions-contract . Non-acceptance of the terms of the General Terms and Conditions (GTC) shall be grounds for disqualification from this procurement process.

Note: Current UNFPA supplier policies apply to this solicitation and can be found at: http://www.unfpa.org/suppliers.

Please take note of the following requirements and conditions pertaining to the supply of the abovementioned goods:

| Goods | All the goods to be supplied under the contract may have their origin in any country should be international standard |
| --- | --- |
| Delivery Terms [INCOTERMS 2010] | CPT/CFR (Baghdad, Iraq) |
| Customs clearance, if needed, shall be done by: | UNFPA Country Office Iraq |
| Insurance | Cargo Insurance is covered by UNFPA global insurance policy. Please do not include insurance costs in your quotation. |
| Delivery Location | Erbil, Iraq |
| Delivery Schedule | Required |
| Mode of Transport | LAND |
| Currency of Quotation | United States Dollars |
| Documents to be submitted | - Duly Accomplished Form as provided in Annex 2, and in accordance with the list of requirements in Annex 1; - Your previous experience of supplying UN agencies (which agency, when, where, what kind of product) - Your company profile, date of establishment, number of staff, legal registration, current licenses - List of other clients - Catalogs, including images for all items specified in Annex 1 - Manufacturing dates for medical supplies - Your company address, and telephone and name of contact person - Quality Certificates (ISO, etc.); - Certificate of Exclusive Distributorship in the country (if applicable, and if Supplier is not the manufacturer); - Written Self-Declaration of not being included in the UN Security Council 1267/1989 list, UN Procurement Division List or other UN Ineligibility List; |
| Period of Validity of Quotes starting the Submission Date | 60 days |
| Delivery time | 10 days upon issuing PO |
| Partial Quotes | Not Permitted |
| Partial Delivery | Permitted |
| Payment Terms | 100% after delivery |
| Evaluation Criteria | - Full compliance to requirements and lowest price - Full acceptance of the UNFPA General Terms and Conditions - Fully compliance with UNFPA Technical requirements for MEDICAL DEVICES |
| UNFPA will award to: | One suppliers |
| Type of Contract to be Signed | Purchase Order |
| Conditions for Release of Payment | Written Acceptance of Goods of beneficiary based on full compliance with RFQ requirements |
| Annexes to this RFQ | - Specifications of the Goods Required (Annex 1) - Quotation Form (Annex 2) Annex 2 need to be duly completed and submitted in order to qualify for the procurement process. |
| Special Condition | - Offers without documents specified will not be considered - Vendors should strictly follow delivery time |

Annex 1

Technical Specifications

| Items to be Supplied | Quantity | Specifications of Goods |
| --- | --- | --- |
| Portable ultrasound | 10 | Portable high-performance ultrasound system, main unit. Full digital portable model Compact and lightweight, easy to carry via handle Size Size WDH, 31.5* 35.3* 25.3cm, Weight 6.4kg Monitor with screen: easy operation handling Scanning method: electronic convex sector probe, electronic linear, also support endo-rectal probes Imaging display modes: , B, B+B, B+M, B+M/M, M, 4B Imaging technologies: dynamic frequency imaging, multi-stage focusing and aperture control. Image gray-scale: 256 levels Depth range selection: convex sector image and linear image, 16 steps adjustable Image orientation: support for lateral and vertical inversion (in B mode). 128-frame CINE loop . 90-frame image storage |
| Suction machine 1.5 - 2 Liters capacity, electric, with infant and adult tubes | 20(1,000 adult tubes and 1,000 infant tube) | For use in minor surgical procedures, E.N.T., resuscitation and for aspiration of secretions, pus, bloods etc. Mounted on a stable, portable stand with castors/wheels and handle Castors / wheels to have fully 360 degree swivel, minimum size 75mm. Sturdy shock resistant casing, easy to carry and easy clean. Vacuum Extractor for intermittent and continuous suction. Reciprocating, oil free pump, twin piston type (metallic). Adjustable suction from 0-675mm/Hg (800 mbar) Suction flow maximum, approximately 17L/minute Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 100%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 100%. Suction pipe length: approx. 1.5 m. Explosion proof foot switch for control of vacuum (vacuum 9 bar) Supplied with graduated plastic jar auto-clavable of 3 litters with overflow safety device and antibacterial filter Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Power consumption, approximately 100W Material: Graduated jar: polysulphone Connection/suction tube: silicone (medical grade Supplied with: Sterilizable/disposable silicone tubes set for molar aspiration & missed abortion sizes 6, 7, 8, 9 and 10 mm , length approximately 2.5m Connectors for fitting the silicone tubes. Instruction for user: manual with instructions for assembly, use and maintenance, and list accessories and spare parts, in 3 languages (English, French and Spanish). Accessories/Spare parts/Consumables: 1 x Set of silicone tubes 1 x Set of connectors fitting the silicone tubes 1 x Spare jar with cover, gasket and overflow valve 1 x Set of spare antibacterial filters 1 x Set of spare fuses |
| Autoclave 24 liters | 20 | Steam sterilizer, pressure type, capacity 24 litres. Product Description: Steam sterilizer/autoclave for sterilizing medical materials such as dressings and surgical instruments under superheated steam. Metal vessel with high-pressure seal suitable for sterilization under superheated steam. Maximum pressure: 21 PSI / 1.5 bar. Maximum temperature: 259ºF / 126ºC. Metal to metal seal (no rubber gaskets). Safety clamping locks: retaining bayonet clamp and (6) Bakelite wing nuts to prevent displacement of cover while sterilizer is under pressure. Inner Chamber dimensions: Chamber, approximately: diameter 315mm x height 290mm. Aluminium container, approximately: diameter 285mm x height 250mm. Sterilization capacity, approximately: 24 litres. Removable cover equipped with: Bakelite handle. Dial type geared steam gauge graduated in kg / cm², PSI and degrees Fahrenheit, and with colour-coding showing sterilizing zone (green) and caution zone (red). Control valve and flexible metal exhaust tube. Excess pressure relief valve and over-pressure rubber plug. Supplied with: 1 x spare over pressure rubber plug Instructions for use and cleaning/maintenance and with diagrams for assembly/disassembly; in 3 languages (English, French and Spanish). Listing of parts, accessories and spare parts. |

| O2 cylinders with regulators | 30 | Standard Flow Rate: 40 l/min Maximum Flow Rate: 150 l/min Single Stage Pressure Regulator with Flowmeter 0-15 l/per min. Pre-Set Pressure Relief Valve, calibrated at 10 bar Vertical/Horizontal Positioning available 6cm long adaptor between flowmeter & regulator Tubing Nipple Every unit rigorously tested prior to dispatch Batch Number stamped on regulator for identification purposes Comply fully with the requirements of HTM 2022 |
| --- | --- | --- |
| Sonic aid | 30 | Heart Rate Display (LCD) Automatic Display Freezing After remove from the body Heart Rate Range 50-240 bpm Accuracy ± 2 or 2% Ultrasound Frequency 2MHz Ultrasound Intensity <10mW/cm2 Out put Headset with Volume Control Double Audio Output sockets for Headset Power Supply: 1.5V Battery x 2 Sensitivity From 10 - 12 weeks PC Interface cable |
| Portable light source chargeable | 30 | A light used during examinations or treatment of the patient. It is usually used during examination or treatment during minor events Designed to be easily moved from one location to another. Arm: 105 +/- cm articulated, spring loaded arm, arm with on/off switch and incorporated electronically transformer. Integral rechargeable battery for operation without mains electricity. Mobile stand with swivel castors. Lamp emits natural white light: colour temperature approx. 4000 K. Halogen bulb approx. 12V/20W Light intensity: approx 20.000 Lux at 40 cm. Reflector adjustable for positioning. Free cord: length approx 3 m. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. The country of destination would be stated in the UNFPA’s PO. Supplied with: 1 x spare bulb a 1x spare fuse. 1x Instructions for use: Manual with instructions for use, cleaning, diagrams for assembly/disassembly (if applicable) and maintenance , storage and handling in English, French, and Spanish) Accessories/ spare parts/consumables 1 spare fuse 1 spare bulb |
| Infant Weighing Scale | 30 | Max. 12 Kgs Digital scale |
| Adult Weight Scale | 30 | Max. 250 Kgs Digital Scale |
| ICU bed | 06 | Patient Bed, Critical Care, electric with following features: Patient Controls: - Type: Clearly visible with symbols - Functions: Positioning, and nurse call Automatic Knee Break (Contour): Yes Retractable: Yes Maximum Patient Weight, kg: ≥225 Adjustable Height ≥80 cm Side rail: Yes Casters: Diameter 12.5 cm, swivel, conductive, ≥2 locking Central Brake System: Yes Removable Headboard: Yes IV Pole: ≥1 on each side Power: 230 VAC Items for the critical care patient bed, electric: a- Pressure-reducing mattress for full body pressure redistribution. The mattress should meet the flammability standards. Mattress-padding waterproof, washable plastic and flame retardant. An anti-microbial agent to be contained in the surface of the mattress to assist in prohibiting growth of bacteria and fungus. Blankets for each bed. Bed sheets for each bed Pillows with cover for each bed. |
| A Bedside Patient Monitor | 04 | Patient Monitor Monitors parameters: ECG, NIBP, SpO2, respiration and temperature ECG: - Number of leads: 5 - HR range, bpm: 30-250 - Alarms: Yes - Arrhythmia detect: Yes - ST analysis: Yes NIBP: Cuff size: adult and child Pulse Oximetry- SpO2: Yes - Probe types: adult and pediatric - Reusable probe: Yes Respiration: Waveform displayed: Yes Temperature: number of inputs: 1 Alarms: visual and audible Display: - Flat panel color LCD ≥ 10.1" - Traces: 4 Line Power: 230 VAC Built-in rechargeable battery with operating time of at least ≥60 minutes. |
| Ventilator ICU | 06 | Ventilator, Intensive Care, with medical air compressor. Microprocessor controlled unit for adult, pediatric and infant. Mounted on mobile cart (caster antistatic wheels) with brakes. The unit shall incorporate ≥11" color LCD touch screen, TFT with configurable display screen. Line power: 230 VAC and rechargeable internal backup battery with operating time 1 hour |
| Mobile examination light | 30 | A light used during examinations or treatment of the patient. It is usually used during examination or treatment during minor events Designed to be easily moved from one location to another. Arm: 105 +/- cm articulated, spring loaded arm, arm with on/off switch and incorporated electronically transformer. Integral rechargeable battery for operation without mains electricity. Mobile stand with swivel castors. Lamp emits natural white light: colour temperature approx. 4000 K. |
| HB Meter | 30 | Haemoglobin photometer handheld w/accessories General Description: Hemoglobinometer, portable, with printer Product description Hand-held device easy transportation and set-up. Provides direct absorbance reading of haemoglobin from inserted cuvette with whole blood. Measuring range, approx: 0-26 g/dL. Accuracy: Correlation with ICSH reference method (HiCN): 0.99 Wavelength: 506 nm (iso-bestic point) and 880 nm (turbidity compensation). Display informs: haemoglobin reading, reading errors, systems errors, battery status. Read-out in: g/dL. Interfaces: RS 232 or USB to printer or computer. Uses dedicated single-use micro-cuvette(closed system). Cuvette allows collecting blood from patient’s skin by capillary action. Sample size in cuvette approx: 8 to 10 uL (capillary, venous or arterial whole blood). Automatic zero setting between measurements. Reading time, approx: 15 sec. Power supply: 110 to 220V, 50 and 60 Hz via universal adaptor or 4 AA / R6 1.5 V batteries. Suitable for use in tropical countries with temperature up to 40 degree celsius. Portable Factory calibrated and built-in self-test (when device is switched-on). Supplied as set containing: 1x Hand-held haemoglobin measuring device. 1 x Set of 200 micro-cuvette. 1 x Box of 200 safety lancets (sterile single-use, auto-disable, incision 2.2mm) 1 x Set of 5 cleaners. 1 x Set of 4 AA / R6 1.5 V batteries (separately packed). 1 x CD, user training and troubleshooting (English, French, Spanish). 1 x Storage and transportation box (hard-case) |
| Incubator | 25 | 1. Servo Control Mode: Air temperature, skin temperature; 2. Instant Parameter Displaying: Set temp, air temp, skin temp, internal humidity, heat power rate 3. Audible & Visual Alarms: Power failure, sensor failure, higher & lower temp, air& skin temp override, fan failure, system failure; 4. >37° C Preset Mode: Power failure reset memory, alarm silence, keypad lock, self-test function; 5. Hood with dual slope: 6 hand-ports; Left & right cabinets, top-shelf; Display parameters: - Patient temperature - Air Temperature - Visual and audible alarms for: i. Patient and air high/low temperature alarm. ii. Air circulation / probe / system / power failure alarm. iii. Humidity alarm. - Temporal alarm silencer. - Heater power indicator |
| Electrosurgical unit, w/accessories with Diathermy machine (bipolar) | 01 | Product Description: Electro surgical generator for use in general surgery for the cutting of soft tissues and coagulation Microprocessor controlled generator provides mono-polar and bi-polar output High frequency Suitable for monopolar cutting and coagulation Suitable for bipolar electro surgery Monopolar maximum power up to 200W (cutting), up to 100W (coagulation) Bipolar maximum power up to 45 W (coagulation) Modes: cutting (pure, blend and haemostasis) and coagulating (soft, force, spray and bi-polar) Front panel allows control of: power cutting, power coagulation, on/off Double function foot switch (mono and bi-polar), with yellow pedal cutting and blue pedal coagulation Hand switch mode when button-activated probes are connected Return circuit sensing monitors and deactivates generator in case patient plate fails Display reports: output power, system errors and electrode failure Automatic limitation of output voltage Reduced emission of interferences and limitation of HF leakage current Alarm in case of HF overpower Visual and audible activation indicators required Visual and audible cable disconnection alarm required Protected against defibrillation shocks Autoclavable handle and electrodes that can withstands more than 250 sterilization cycles at 134ºC. Power supply: The supplier has to supply the right voltage, plug and socket type that fit the country it is delivering the goods to. Supplied with: 1 x Foot switch, two pedals, yellow and blue, with connecting cable 2 x Patient plate, reusable, with connecting cable (adult & child) 2 x Mono-polar electrode handle, reusable, foot switch controlled, with connecting cable 2 x Mono-polar electrode handle, reusable, finger switch controlled, with connecting cable 1 x Set different mono-polar reusable electrodes (needle, blade, ball and loop) 2 x Bi-polar forceps, reusable, foot switch controlled, with connecting cable (short, straight, tip-angled) |

Annex 2

Quotation Form

Name of Bidder:

Date of Bid:

Request for Quotation No:

Currency of Bid price:

Delivery time (10 days upon issuing PO):

Expiration of Validity of Quotation (The quotation shall be

valid for a period of at least 60 days after the Closing date.):

Price Schedule:

| Item No. | FOB/FCA price/unit specify port of shipment | ## Quantity | Transportation cost to port of Destination (specify mode of trp.) | Total CPT/CFR (Destination port) | Delivery schedule (months) |
| --- | --- | --- | --- | --- | --- |

Vendor’s Comments:

I hereby certify that this company, which I am duly authorized to sign for, accepts the terms and conditions of UNFPA and we will abide by this quotation until it expires.

Name and title					Date and Place